## PS
Pressure Support
Lung Model
|  | Elastance (cm H2O/L) | 25
|  | Compliance (L/cm H2O) | 0.04 |  | 40
|  | Resistance (cm H2O/L/s) | 10 |  | 10
|  | Time Constant (s) | 0.4
|  | Tidal Volume (mL) | 355.2399
Effort Model
|  | Delay Before Effort (s) | 0.12 |  | 12
|  | Duration of Effort (s) | 0.2 |  | 4
|  | Freq (hz) | 1.25
|  | Period (s) | 0.4
|  | max Pmus (cm H2O) | 5 | 49 | 50
Ventilator Settings
|  | Inspiratory Pressure Target (cm H2O) | 10 |  | 10
|  | Peak Inspiratory Pressure (cm H2O) | 15
|  | Mean Inspiratory Pressure (cm H2O) | 15
|  | Peak Inspiratory Flow (L/min) | 69.1034
|  | Inspirtory Time (s) | 0.57
|  | Flow Trigger Threshold (L/min) | 4 |  | 4
|  | Flow Cycle Threshold (%) | 10 |  | 10
|  | Flow Cycle Threshold (L/min) | 6.9103
|  | PEEP (cm H2O) | 5 |  | 5
|  | Pressure Rise Time (s) | 0 |  | 0
|  | Peak Inspiratory Flow (L/min) | 69.1034
|  | End Inspiratory Flow (L/min) | 7.0627
|  | Inspirtory Time (s) | 0.57
|  | Peak Expiratory Flow (L/min) | -53.286
1245
Pk Flow
69.1034
time Pk Flow
[10 empty] | Vmus equiv to constant mean Pmus | [5 empty] | 0.21
[7 empty] | DP | P rise-t |  | 0.080532 |  | Initial |  |  |  | Pk exp Flow |  |  |  | Mean
[7 empty] | 10 | 0.001 |  |  |  | 0.2 |  |  | Insp | -53.286 |  |  |  | Pressure
[9 empty] | Vent | Muscle | Lung | Lung | Tidal | Vent | Time | Vent | Max |  | Relative | (cm H2O)
[6 empty] | Trigger |  | Relative | Pressure | Pressure | Pressure | Volume | Volume | Flow | 0.57 | Flow | Flow | Cycle | Time (Pvent) | 15
[6 empty] | Signal | Time | Time (musc) | (cm H2O) | (cm H2O) | (cm H2O) | (L) | (mL) | (L/s) |  | (L/min) | (L/min) | Signal | (s)
[6 empty] | 0 | 0 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.01 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.02 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.03 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.04 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.05 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.06 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.07 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.08 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.09 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.1 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.11 | 0 | 5 | 0 | 5 | 0.2 | 0 | 0 | 0 | 0 | 0 | 1 | 0
[6 empty] | 0 | 0.12 | 0.01 | 5 | 0.3923 | 5 | 0.2 | 0 | 0.03923 | 0 | 2.3538 | 2.3538 | 1 | 0
[6 empty] | 0 | 0.13 | 0.02 | 5 | 0.78217 | 5.0098 | 0.20039 | 0.3923 | 0.077236 | 0 | 4.6342 | 4.6342 | 1 | 0
[6 empty] | 1 | 0.14 | 0.03 | 14.9995 | 1.1672 | 5.0291 | 0.20116 | 1.1647 | 1.1138 | 0.01 | 66.8259 | 66.8259 | 1 | 0.01 | 14.9995
[6 empty] | 1 | 0.15 | 0.04 | 15 | 1.5451 | 5.3076 | 0.2123 | 12.3023 | 1.1238 | 0.02 | 67.4252 | 67.4252 | 1 | 0.02 | 15
[6 empty] | 1 | 0.16 | 0.05 | 15 | 1.9134 | 5.5885 | 0.22354 | 23.5398 | 1.1325 | 0.03 | 67.9495 | 67.9495 | 1 | 0.03 | 15
[6 empty] | 1 | 0.17 | 0.06 | 15 | 2.27 | 5.8716 | 0.23486 | 34.8648 | 1.1398 | 0.04 | 68.39 | 68.39 | 1 | 0.04 | 15
[6 empty] | 1 | 0.18 | 0.07 | 15 | 2.6125 | 6.1566 | 0.24626 | 46.2631 | 1.1456 | 0.05 | 68.7355 | 68.7355 | 1 | 0.05 | 15
[6 empty] | 1 | 0.19 | 0.08 | 15 | 2.9389 | 6.443 | 0.25772 | 57.719 | 1.1496 | 0.06 | 68.9757 | 68.9757 | 1 | 0.06 | 15
[6 empty] | 1 | 0.2 | 0.09 | 15 | 3.2472 | 6.7304 | 0.26921 | 69.215 | 1.1517 | 0.07 | 69.1012 | 69.1012 | 1 | 0.07 | 15
[6 empty] | 1 | 0.21 | 0.1 | 15 | 3.5355 | 7.0183 | 0.28073 | 80.7318 | 1.1517 | 0.08 | 69.1034 | 69.1034 | 1 | 0.08 | 15
[6 empty] | 1 | 0.22 | 0.11 | 15 | 3.802 | 7.3062 | 0.29225 | 92.2491 | 1.1496 | 0.09 | 68.9748 | 69.1034 | 1 | 0.09 | 15
[6 empty] | 1 | 0.23 | 0.12 | 15 | 4.0451 | 7.5936 | 0.30374 | 103.7449 | 1.1451 | 0.1 | 68.7088 | 69.1034 | 1 | 0.1 | 15
[6 empty] | 1 | 0.24 | 0.13 | 15 | 4.2632 | 7.8799 | 0.3152 | 115.1963 | 1.1383 | 0.11 | 68.2998 | 69.1034 | 1 | 0.11 | 15
[6 empty] | 1 | 0.25 | 0.14 | 15 | 4.455 | 8.1645 | 0.32658 | 126.5796 | 1.1291 | 0.12 | 67.7433 | 69.1034 | 1 | 0.12 | 15
[6 empty] | 1 | 0.26 | 0.15 | 15 | 4.6194 | 8.4468 | 0.33787 | 137.8702 | 1.1173 | 0.13 | 67.0359 | 69.1034 | 1 | 0.13 | 15
[6 empty] | 1 | 0.27 | 0.16 | 15 | 4.7553 | 8.7261 | 0.34904 | 149.0428 | 1.1029 | 0.14 | 66.1753 | 69.1034 | 1 | 0.14 | 15
[6 empty] | 1 | 0.28 | 0.17 | 15 | 4.8618 | 9.0018 | 0.36007 | 160.072 | 1.086 | 0.15 | 65.1603 | 69.1034 | 1 | 0.15 | 15
[6 empty] | 1 | 0.29 | 0.18 | 15 | 4.9384 | 9.2733 | 0.37093 | 170.9321 | 1.0665 | 0.16 | 63.9908 | 69.1034 | 1 | 0.16 | 15
[6 empty] | 1 | 0.3 | 0.19 | 15 | 4.9846 | 9.5399 | 0.3816 | 181.5972 | 1.0445 | 0.17 | 62.6679 | 69.1034 | 1 | 0.17 | 15
[6 empty] | 1 | 0.31 | 0.2 | 15 | 5 | 9.801 | 0.39204 | 192.0419 | 1.0199 | 0.18 | 61.1937 | 69.1034 | 1 | 0.18 | 15
[6 empty] | 1 | 0.32 | 0.21 | 15 | 4.9846 | 10.056 | 0.40224 | 202.2408 | 0.99286 | 0.19 | 59.5714 | 69.1034 | 1 | 0.19 | 15
[6 empty] | 1 | 0.33 | 0.22 | 15 | 4.9384 | 10.3042 | 0.41217 | 212.1694 | 0.96342 | 0.2 | 57.8052 | 69.1034 | 1 | 0.2 | 15
[6 empty] | 1 | 0.34 | 0.23 | 15 | 4.8618 | 10.5451 | 0.4218 | 221.8036 | 0.93168 | 0.21 | 55.9006 | 69.1034 | 1 | 0.21 | 15
[6 empty] | 1 | 0.35 | 0.24 | 15 | 4.7553 | 10.778 | 0.43112 | 231.1204 | 0.89773 | 0.22 | 53.8636 | 69.1034 | 1 | 0.22 | 15
[6 empty] | 1 | 0.36 | 0.25 | 15 | 4.6194 | 11.0024 | 0.4401 | 240.0976 | 0.8617 | 0.23 | 51.7017 | 69.1034 | 1 | 0.23 | 15
[6 empty] | 1 | 0.37 | 0.26 | 15 | 4.455 | 11.2179 | 0.44871 | 248.7146 | 0.82372 | 0.24 | 49.423 | 69.1034 | 1 | 0.24 | 15
[6 empty] | 1 | 0.38 | 0.27 | 15 | 4.2632 | 11.4238 | 0.45695 | 256.9518 | 0.78394 | 0.25 | 47.0364 | 69.1034 | 1 | 0.25 | 15
[6 empty] | 1 | 0.39 | 0.28 | 15 | 4.0451 | 11.6198 | 0.46479 | 264.7912 | 0.74253 | 0.26 | 44.5518 | 69.1034 | 1 | 0.26 | 15
[6 empty] | 1 | 0.4 | 0.29 | 15 | 3.802 | 11.8054 | 0.47222 | 272.2165 | 0.69966 | 0.27 | 41.9797 | 69.1034 | 1 | 0.27 | 15
[6 empty] | 1 | 0.41 | 0.3 | 15 | 3.5355 | 11.9803 | 0.47921 | 279.2131 | 0.65552 | 0.28 | 39.3312 | 69.1034 | 1 | 0.28 | 15
[6 empty] | 1 | 0.42 | 0.31 | 15 | 3.2472 | 12.1442 | 0.48577 | 285.7683 | 0.6103 | 0.29 | 36.6182 | 69.1034 | 1 | 0.29 | 15
[6 empty] | 1 | 0.43 | 0.32 | 15 | 2.9389 | 12.2968 | 0.49187 | 291.8713 | 0.56421 | 0.3 | 33.8529 | 69.1034 | 1 | 0.3 | 15
[6 empty] | 1 | 0.44 | 0.33 | 15 | 2.6125 | 12.4378 | 0.49751 | 297.5135 | 0.51747 | 0.31 | 31.0479 | 69.1034 | 1 | 0.31 | 15
[6 empty] | 1 | 0.45 | 0.34 | 15 | 2.27 | 12.5672 | 0.50269 | 302.6881 | 0.47027 | 0.32 | 28.2165 | 69.1034 | 1 | 0.32 | 15
[6 empty] | 1 | 0.46 | 0.35 | 15 | 1.9134 | 12.6848 | 0.50739 | 307.3909 | 0.42286 | 0.33 | 25.3719 | 69.1034 | 1 | 0.33 | 15
[6 empty] | 1 | 0.47 | 0.36 | 15 | 1.5451 | 12.7905 | 0.51162 | 311.6195 | 0.37546 | 0.34 | 22.5276 | 69.1034 | 1 | 0.34 | 15
[6 empty] | 1 | 0.48 | 0.37 | 15 | 1.1672 | 12.8844 | 0.51537 | 315.3741 | 0.32829 | 0.35 | 19.6972 | 69.1034 | 1 | 0.35 | 15
[6 empty] | 1 | 0.49 | 0.38 | 15 | 0.78217 | 12.9664 | 0.51866 | 318.657 | 0.28157 | 0.36 | 16.8945 | 69.1034 | 1 | 0.36 | 15
[6 empty] | 1 | 0.5 | 0.39 | 15 | 0.3923 | 13.0368 | 0.52147 | 321.4727 | 0.23555 | 0.37 | 14.1329 | 69.1034 | 1 | 0.37 | 15
[6 empty] | 1 | 0.51 | 0.4 | 15 | -6.049e-15 | 13.0957 | 0.52383 | 323.8282 | 0.19043 | 0.38 | 11.4258 | 69.1034 | 1 | 0.38 | 15
[6 empty] | 1 | 0.52 | 0.41 | 15 | 0 | 13.1433 | 0.52573 | 325.7325 | 0.18567 | 0.39 | 11.1401 | 69.1034 | 1 | 0.39 | 15
[6 empty] | 1 | 0.53 | 0.42 | 15 | 0 | 13.1897 | 0.52759 | 327.5892 | 0.18103 | 0.4 | 10.8616 | 69.1034 | 1 | 0.4 | 15
[6 empty] | 1 | 0.54 | 0.43 | 15 | 0 | 13.235 | 0.5294 | 329.3995 | 0.1765 | 0.41 | 10.5901 | 69.1034 | 1 | 0.41 | 15
[6 empty] | 1 | 0.55 | 0.44 | 15 | 0 | 13.2791 | 0.53116 | 331.1645 | 0.17209 | 0.42 | 10.3253 | 69.1034 | 1 | 0.42 | 15
[6 empty] | 1 | 0.56 | 0.45 | 15 | 0 | 13.3221 | 0.53289 | 332.8854 | 0.16779 | 0.43 | 10.0672 | 69.1034 | 1 | 0.43 | 15
[6 empty] | 1 | 0.57 | 0.46 | 15 | 0 | 13.3641 | 0.53456 | 334.5632 | 0.16359 | 0.44 | 9.8155 | 69.1034 | 1 | 0.44 | 15
[6 empty] | 1 | 0.58 | 0.47 | 15 | 0 | 13.405 | 0.5362 | 336.1992 | 0.1595 | 0.45 | 9.5701 | 69.1034 | 1 | 0.45 | 15
[6 empty] | 1 | 0.59 | 0.48 | 15 | 0 | 13.4449 | 0.53779 | 337.7942 | 0.15551 | 0.46 | 9.3309 | 69.1034 | 1 | 0.46 | 15
[6 empty] | 1 | 0.6 | 0.49 | 15 | 0 | 13.4837 | 0.53935 | 339.3493 | 0.15163 | 0.47 | 9.0976 | 69.1034 | 1 | 0.47 | 15
[6 empty] | 1 | 0.61 | 0.5 | 15 | 0 | 13.5216 | 0.54087 | 340.8656 | 0.14784 | 0.48 | 8.8702 | 69.1034 | 1 | 0.48 | 15
[6 empty] | 1 | 0.62 | 0.51 | 15 | 0 | 13.5586 | 0.54234 | 342.3439 | 0.14414 | 0.49 | 8.6484 | 69.1034 | 1 | 0.49 | 15
[6 empty] | 1 | 0.63 | 0.52 | 15 | 0 | 13.5946 | 0.54379 | 343.7853 | 0.14054 | 0.5 | 8.4322 | 69.1034 | 1 | 0.5 | 15
[6 empty] | 1 | 0.64 | 0.53 | 15 | 0 | 13.6298 | 0.54519 | 345.1907 | 0.13702 | 0.51 | 8.2214 | 69.1034 | 1 | 0.51 | 15
[6 empty] | 1 | 0.65 | 0.54 | 15 | 0 | 13.664 | 0.54656 | 346.5609 | 0.1336 | 0.52 | 8.0159 | 69.1034 | 1 | 0.52 | 15
[6 empty] | 1 | 0.66 | 0.55 | 15 | 0 | 13.6974 | 0.5479 | 347.8969 | 0.13026 | 0.53 | 7.8155 | 69.1034 | 1 | 0.53 | 15
[6 empty] | 1 | 0.67 | 0.56 | 15 | 0 | 13.73 | 0.5492 | 349.1995 | 0.127 | 0.54 | 7.6201 | 69.1034 | 1 | 0.54 | 15
[6 empty] | 1 | 0.68 | 0.57 | 15 | 0 | 13.7617 | 0.55047 | 350.4695 | 0.12383 | 0.55 | 7.4296 | 69.1034 | 1 | 0.55 | 15
[6 empty] | 1 | 0.69 | 0.58 | 15 | 0 | 13.7927 | 0.55171 | 351.7078 | 0.12073 | 0.56 | 7.2438 | 69.1034 | 1 | 0.56 | 15
[6 empty] | 1 | 0.7 | 0.59 | 15 | 0 | 13.8229 | 0.55292 | 352.9151 | 0.11771 | 0.57 | 7.0627 | 69.1034 | 1 | 0.57 | 15
[6 empty] | 1 | 0.71 | 0.6 | 15 | 0 | 13.8523 | 0.55409 | 354.0922 | 0.11477 | 0 | 6.8862 | 69.1034 | 0 | 0.58
[6 empty] | 1 | 0.72 | 0.61 | 5 | 0 | 13.881 | 0.55524 | 355.2399 | -0.8881 | 0 | -53.286 | 69.1034 | 0 | 0.59
[6 empty] | 1 | 0.73 | 0.62 | 5 | 0 | 13.659 | 0.54636 | 346.3589 | -0.8659 | 0 | -51.9538 | 69.1034 | 0 | 0.6
[6 empty] | 1 | 0.74 | 0.63 | 5 | 0 | 13.4425 | 0.5377 | 337.6999 | -0.84425 | 0 | -50.655 | 69.1034 | 0 | 0.61
[6 empty] | 1 | 0.75 | 0.64 | 5 | 0 | 13.2314 | 0.52926 | 329.2574 | -0.82314 | 0 | -49.3886 | 69.1034 | 0 | 0.62
[6 empty] | 1 | 0.76 | 0.65 | 5 | 0 | 13.0256 | 0.52103 | 321.026 | -0.80256 | 0 | -48.1539 | 69.1034 | 0 | 0.63
[6 empty] | 1 | 0.77 | 0.66 | 5 | 0 | 12.825 | 0.513 | 313.0003 | -0.7825 | 0 | -46.9501 | 69.1034 | 0 | 0.64
[6 empty] | 1 | 0.78 | 0.67 | 5 | 0 | 12.6294 | 0.50518 | 305.1753 | -0.76294 | 0 | -45.7763 | 69.1034 | 0 | 0.65
[6 empty] | 1 | 0.79 | 0.68 | 5 | 0 | 12.4386 | 0.49755 | 297.546 | -0.74386 | 0 | -44.6319 | 69.1034 | 0 | 0.66
[6 empty] | 1 | 0.8 | 0.69 | 5 | 0 | 12.2527 | 0.49011 | 290.1073 | -0.72527 | 0 | -43.5161 | 69.1034 | 0 | 0.67
[6 empty] | 1 | 0.81 | 0.7 | 5 | 0 | 12.0714 | 0.48285 | 282.8546 | -0.70714 | 0 | -42.4282 | 69.1034 | 0 | 0.68
[6 empty] | 1 | 0.82 | 0.71 | 5 | 0 | 11.8946 | 0.47578 | 275.7833 | -0.68946 | 0 | -41.3675 | 69.1034 | 0 | 0.69
[6 empty] | 1 | 0.83 | 0.72 | 5 | 0 | 11.7222 | 0.46889 | 268.8887 | -0.67222 | 0 | -40.3333 | 69.1034 | 0 | 0.7
[6 empty] | 1 | 0.84 | 0.73 | 5 | 0 | 11.5542 | 0.46217 | 262.1665 | -0.65542 | 0 | -39.325 | 69.1034 | 0 | 0.71
[6 empty] | 1 | 0.85 | 0.74 | 5 | 0 | 11.3903 | 0.45561 | 255.6123 | -0.63903 | 0 | -38.3418 | 69.1034 | 0 | 0.72
[6 empty] | 1 | 0.86 | 0.75 | 5 | 0 | 11.2305 | 0.44922 | 249.222 | -0.62305 | 0 | -37.3833 | 69.1034 | 0 | 0.73
[6 empty] | 1 | 0.87 | 0.76 | 5 | 0 | 11.0748 | 0.44299 | 242.9914 | -0.60748 | 0 | -36.4487 | 69.1034 | 0 | 0.74
[6 empty] | 1 | 0.88 | 0.77 | 5 | 0 | 10.9229 | 0.43692 | 236.9167 | -0.59229 | 0 | -35.5375 | 69.1034 | 0 | 0.75
[6 empty] | 1 | 0.89 | 0.78 | 5 | 0 | 10.7748 | 0.43099 | 230.9937 | -0.57748 | 0 | -34.6491 | 69.1034 | 0 | 0.76
[6 empty] | 1 | 0.9 | 0.79 | 5 | 0 | 10.6305 | 0.42522 | 225.2189 | -0.56305 | 0 | -33.7828 | 69.1034 | 0 | 0.77
[6 empty] | 1 | 0.91 | 0.8 | 5 | 0 | 10.4897 | 0.41959 | 219.5884 | -0.54897 | 0 | -32.9383 | 69.1034 | 0 | 0.78
[6 empty] | 1 | 0.92 | 0.81 | 5 | 0 | 10.3525 | 0.4141 | 214.0987 | -0.53525 | 0 | -32.1148 | 69.1034 | 0 | 0.79
[6 empty] | 1 | 0.93 | 0.82 | 5 | 0 | 10.2187 | 0.40875 | 208.7462 | -0.52187 | 0 | -31.3119 | 69.1034 | 0 | 0.8
[6 empty] | 1 | 0.94 | 0.83 | 5 | 0 | 10.0882 | 0.40353 | 203.5276 | -0.50882 | 0 | -30.5291 | 69.1034 | 0 | 0.81
[6 empty] | 1 | 0.95 | 0.84 | 5 | 0 | 9.961 | 0.39844 | 198.4394 | -0.4961 | 0 | -29.7659 | 69.1034 | 0 | 0.82
[6 empty] | 1 | 0.96 | 0.85 | 5 | 0 | 9.837 | 0.39348 | 193.4784 | -0.4837 | 0 | -29.0218 | 69.1034 | 0 | 0.83
[6 empty] | 1 | 0.97 | 0.86 | 5 | 0 | 9.716 | 0.38864 | 188.6415 | -0.4716 | 0 | -28.2962 | 69.1034 | 0 | 0.84
[6 empty] | 1 | 0.98 | 0.87 | 5 | 0 | 9.5981 | 0.38393 | 183.9254 | -0.45981 | 0 | -27.5888 | 69.1034 | 0 | 0.85
[6 empty] | 1 | 0.99 | 0.88 | 5 | 0 | 9.4832 | 0.37933 | 179.3273 | -0.44832 | 0 | -26.8991 | 69.1034 | 0 | 0.86
[6 empty] | 1 | 1 | 0.89 | 5 | 0 | 9.3711 | 0.37484 | 174.8441 | -0.43711 | 0 | -26.2266 | 69.1034 | 0 | 0.87
[6 empty] | 1 | 1.01 | 0.9 | 5 | 0 | 9.2618 | 0.37047 | 170.473 | -0.42618 | 0 | -25.5709 | 69.1034 | 0 | 0.88
[6 empty] | 1 | 1.02 | 0.91 | 5 | 0 | 9.1553 | 0.36621 | 166.2112 | -0.41553 | 0 | -24.9317 | 69.1034 | 0 | 0.89
[6 empty] | 1 | 1.03 | 0.92 | 5 | 0 | 9.0514 | 0.36206 | 162.0559 | -0.40514 | 0 | -24.3084 | 69.1034 | 0 | 0.9
[6 empty] | 1 | 1.04 | 0.93 | 5 | 0 | 8.9501 | 0.358 | 158.0045 | -0.39501 | 0 | -23.7007 | 69.1034 | 0 | 0.91
[6 empty] | 1 | 1.05 | 0.94 | 5 | 0 | 8.8514 | 0.35405 | 154.0544 | -0.38514 | 0 | -23.1082 | 69.1034 | 0 | 0.92
[6 empty] | 1 | 1.06 | 0.95 | 5 | 0 | 8.7551 | 0.3502 | 150.203 | -0.37551 | 0 | -22.5305 | 69.1034 | 0 | 0.93
[6 empty] | 1 | 1.07 | 0.96 | 5 | 0 | 8.6612 | 0.34645 | 146.4479 | -0.36612 | 0 | -21.9672 | 69.1034 | 0 | 0.94
[6 empty] | 1 | 1.08 | 0.97 | 5 | 0 | 8.5697 | 0.34279 | 142.7867 | -0.35697 | 0 | -21.418 | 69.1034 | 0 | 0.95
[6 empty] | 1 | 1.09 | 0.98 | 5 | 0 | 8.4804 | 0.33922 | 139.2171 | -0.34804 | 0 | -20.8826 | 69.1034 | 0 | 0.96
[6 empty] | 1 | 1.1 | 0.99 | 5 | 0 | 8.3934 | 0.33574 | 135.7367 | -0.33934 | 0 | -20.3605 | 69.1034 | 0 | 0.97
[6 empty] | 1 | 1.11 | 1 | 5 | 0 | 8.3086 | 0.33234 | 132.3432 | -0.33086 | 0 | -19.8515 | 69.1034 | 0 | 0.98
[6 empty] | 1 | 1.12 | 1.01 | 5 | 0 | 8.2259 | 0.32903 | 129.0347 | -0.32259 | 0 | -19.3552 | 69.1034 | 0 | 0.99
[6 empty] | 1 | 1.13 | 1.02 | 5 | 0 | 8.1452 | 0.32581 | 125.8088 | -0.31452 | 0 | -18.8713 | 69.1034 | 0 | 1
[6 empty] | 1 | 1.14 | 1.03 | 5 | 0 | 8.0666 | 0.32266 | 122.6636 | -0.30666 | 0 | -18.3995 | 69.1034 | 0 | 1.01
[6 empty] | 1 | 1.15 | 1.04 | 5 | 0 | 7.9899 | 0.3196 | 119.597 | -0.29899 | 0 | -17.9395 | 69.1034 | 0 | 1.02
[6 empty] | 1 | 1.16 | 1.05 | 5 | 0 | 7.9152 | 0.31661 | 116.6071 | -0.29152 | 0 | -17.4911 | 69.1034 | 0 | 1.03
[6 empty] | 1 | 1.17 | 1.06 | 5 | 0 | 7.8423 | 0.31369 | 113.6919 | -0.28423 | 0 | -17.0538 | 69.1034 | 0 | 1.04
[6 empty] | 1 | 1.18 | 1.07 | 5 | 0 | 7.7712 | 0.31085 | 110.8496 | -0.27712 | 0 | -16.6274 | 69.1034 | 0 | 1.05
[6 empty] | 1 | 1.19 | 1.08 | 5 | 0 | 7.702 | 0.30808 | 108.0783 | -0.2702 | 0 | -16.2118 | 69.1034 | 0 | 1.06
[6 empty] | 1 | 1.2 | 1.09 | 5 | 0 | 7.6344 | 0.30538 | 105.3764 | -0.26344 | 0 | -15.8065 | 69.1034 | 0 | 1.07
[6 empty] | 1 | 1.21 | 1.1 | 5 | 0 | 7.5685 | 0.30274 | 102.742 | -0.25685 | 0 | -15.4113 | 69.1034 | 0 | 1.08
[6 empty] | 1 | 1.22 | 1.11 | 5 | 0 | 7.5043 | 0.30017 | 100.1734 | -0.25043 | 0 | -15.026 | 69.1034 | 0 | 1.09
[6 empty] | 1 | 1.23 | 1.12 | 5 | 0 | 7.4417 | 0.29767 | 97.6691 | -0.24417 | 0 | -14.6504 | 69.1034 | 0 | 1.1
[6 empty] | 1 | 1.24 | 1.13 | 5 | 0 | 7.3807 | 0.29523 | 95.2274 | -0.23807 | 0 | -14.2841 | 69.1034 | 0 | 1.11
[6 empty] | 1 | 1.25 | 1.14 | 5 | 0 | 7.3212 | 0.29285 | 92.8467 | -0.23212 | 0 | -13.927 | 69.1034 | 0 | 1.12
[6 empty] | 1 | 1.26 | 1.15 | 5 | 0 | 7.2631 | 0.29053 | 90.5255 | -0.22631 | 0 | -13.5788 | 69.1034 | 0 | 1.13
[6 empty] | 1 | 1.27 | 1.16 | 5 | 0 | 7.2066 | 0.28826 | 88.2624 | -0.22066 | 0 | -13.2394 | 69.1034 | 0 | 1.14
[6 empty] | 1 | 1.28 | 1.17 | 5 | 0 | 7.1514 | 0.28606 | 86.0558 | -0.21514 | 0 | -12.9084 | 69.1034 | 0 | 1.15
[6 empty] | 1 | 1.29 | 1.18 | 5 | 0 | 7.0976 | 0.2839 | 83.9044 | -0.20976 | 0 | -12.5857 | 69.1034 | 0 | 1.16
[6 empty] | 1 | 1.3 | 1.19 | 5 | 0 | 7.0452 | 0.28181 | 81.8068 | -0.20452 | 0 | -12.271 | 69.1034 | 0 | 1.17
[6 empty] | 1 | 1.31 | 1.2 | 5 | 0 | 6.994 | 0.27976 | 79.7616 | -0.1994 | 0 | -11.9642 | 69.1034 | 0 | 1.18
[6 empty] | 1 | 1.32 | 1.21 | 5 | 0 | 6.9442 | 0.27777 | 77.7676 | -0.19442 | 0 | -11.6651 | 69.1034 | 0 | 1.19
[6 empty] | 1 | 1.33 | 1.22 | 5 | 0 | 6.8956 | 0.27582 | 75.8234 | -0.18956 | 0 | -11.3735 | 69.1034 | 0 | 1.2
[6 empty] | 1 | 1.34 | 1.23 | 5 | 0 | 6.8482 | 0.27393 | 73.9278 | -0.18482 | 0 | -11.0892 | 69.1034 | 0 | 1.21
[6 empty] | 1 | 1.35 | 1.24 | 5 | 0 | 6.802 | 0.27208 | 72.0796 | -0.1802 | 0 | -10.8119 | 69.1034 | 0 | 1.22
[6 empty] | 1 | 1.36 | 1.25 | 5 | 0 | 6.7569 | 0.27028 | 70.2776 | -0.17569 | 0 | -10.5416 | 69.1034 | 0 | 1.23
[6 empty] | 1 | 1.37 | 1.26 | 5 | 0 | 6.713 | 0.26852 | 68.5207 | -0.1713 | 0 | -10.2781 | 69.1034 | 0 | 1.24
[6 empty] | 1 | 1.38 | 1.27 | 5 | 0 | 6.6702 | 0.26681 | 66.8077 | -0.16702 | 0 | -10.0212 | 69.1034 | 0 | 1.25
[6 empty] | 1 | 1.39 | 1.28 | 5 | 0 | 6.6284 | 0.26514 | 65.1375 | -0.16284 | 0 | -9.7706 | 69.1034 | 0 | 1.26
[6 empty] | 1 | 1.4 | 1.29 | 5 | 0 | 6.5877 | 0.26351 | 63.509 | -0.15877 | 0 | -9.5264 | 69.1034 | 0 | 1.27
[6 empty] | 1 | 1.41 | 1.3 | 5 | 0 | 6.548 | 0.26192 | 61.9213 | -0.1548 | 0 | -9.2882 | 69.1034 | 0 | 1.28
[6 empty] | 1 | 1.42 | 1.31 | 5 | 0 | 6.5093 | 0.26037 | 60.3733 | -0.15093 | 0 | -9.056 | 69.1034 | 0 | 1.29
[6 empty] | 1 | 1.43 | 1.32 | 5 | 0 | 6.4716 | 0.25886 | 58.864 | -0.14716 | 0 | -8.8296 | 69.1034 | 0 | 1.3
[6 empty] | 1 | 1.44 | 1.33 | 5 | 0 | 6.4348 | 0.25739 | 57.3924 | -0.14348 | 0 | -8.6089 | 69.1034 | 0 | 1.31
[6 empty] | 1 | 1.45 | 1.34 | 5 | 0 | 6.3989 | 0.25596 | 55.9575 | -0.13989 | 0 | -8.3936 | 69.1034 | 0 | 1.32
[6 empty] | 1 | 1.46 | 1.35 | 5 | 0 | 6.364 | 0.25456 | 54.5586 | -0.1364 | 0 | -8.1838 | 69.1034 | 0 | 1.33
[6 empty] | 1 | 1.47 | 1.36 | 5 | 0 | 6.3299 | 0.25319 | 53.1946 | -0.13299 | 0 | -7.9792 | 69.1034 | 0 | 1.34
[6 empty] | 1 | 1.48 | 1.37 | 5 | 0 | 6.2966 | 0.25186 | 51.8648 | -0.12966 | 0 | -7.7797 | 69.1034 | 0 | 1.35
[6 empty] | 1 | 1.49 | 1.38 | 5 | 0 | 6.2642 | 0.25057 | 50.5682 | -0.12642 | 0 | -7.5852 | 69.1034 | 0 | 1.36
[6 empty] | 1 | 1.5 | 1.39 | 5 | 0 | 6.2326 | 0.2493 | 49.304 | -0.12326 | 0 | -7.3956 | 69.1034 | 0 | 1.37
[6 empty] | 1 | 1.51 | 1.4 | 5 | 0 | 6.2018 | 0.24807 | 48.0714 | -0.12018 | 0 | -7.2107 | 69.1034 | 0 | 1.38
[6 empty] | 1 | 1.52 | 1.41 | 5 | 0 | 6.1717 | 0.24687 | 46.8696 | -0.11717 | 0 | -7.0304 | 69.1034 | 0 | 1.39
[6 empty] | 1 | 1.53 | 1.42 | 5 | 0 | 6.1424 | 0.2457 | 45.6978 | -0.11424 | 0 | -6.8547 | 69.1034 | 0 | 1.4
[6 empty] | 1 | 1.54 | 1.43 | 5 | 0 | 6.1139 | 0.24456 | 44.5554 | -0.11139 | 0 | -6.6833 | 69.1034 | 0 | 1.41
[6 empty] | 1 | 1.55 | 1.44 | 5 | 0 | 6.086 | 0.24344 | 43.4415 | -0.1086 | 0 | -6.5162 | 69.1034 | 0 | 1.42
[6 empty] | 1 | 1.56 | 1.45 | 5 | 0 | 6.0589 | 0.24236 | 42.3555 | -0.10589 | 0 | -6.3533 | 69.1034 | 0 | 1.43
[6 empty] | 1 | 1.57 | 1.46 | 5 | 0 | 6.0324 | 0.2413 | 41.2966 | -0.10324 | 0 | -6.1945 | 69.1034 | 0 | 1.44
[6 empty] | 1 | 1.58 | 1.47 | 5 | 0 | 6.0066 | 0.24026 | 40.2642 | -0.10066 | 0 | -6.0396 | 69.1034 | 0 | 1.45
[6 empty] | 1 | 1.59 | 1.48 | 5 | 0 | 5.9814 | 0.23926 | 39.2576 | -0.098144 | 0 | -5.8886 | 69.1034 | 0 | 1.46
[6 empty] | 1 | 1.6 | 1.49 | 5 | 0 | 5.9569 | 0.23828 | 38.2761 | -0.09569 | 0 | -5.7414 | 69.1034 | 0 | 1.47
[6 empty] | 1 | 1.61 | 1.5 | 5 | 0 | 5.933 | 0.23732 | 37.3192 | -0.093298 | 0 | -5.5979 | 69.1034 | 0 | 1.48
[6 empty] | 1 | 1.62 | 1.51 | 5 | 0 | 5.9097 | 0.23639 | 36.3862 | -0.090966 | 0 | -5.4579 | 69.1034 | 0 | 1.49
[6 empty] | 1 | 1.63 | 1.52 | 5 | 0 | 5.8869 | 0.23548 | 35.4766 | -0.088691 | 0 | -5.3215 | 69.1034 | 0 | 1.5
[6 empty] | 1 | 1.64 | 1.53 | 5 | 0 | 5.8647 | 0.23459 | 34.5897 | -0.086474 | 0 | -5.1885 | 69.1034 | 0 | 1.51
[6 empty] | 1 | 1.65 | 1.54 | 5 | 0 | 5.8431 | 0.23372 | 33.7249 | -0.084312 | 0 | -5.0587 | 69.1034 | 0 | 1.52
[6 empty] | 1 | 1.66 | 1.55 | 5 | 0 | 5.822 | 0.23288 | 32.8818 | -0.082205 | 0 | -4.9323 | 69.1034 | 0 | 1.53
[6 empty] | 1 | 1.67 | 1.56 | 5 | 0 | 5.8015 | 0.23206 | 32.0598 | -0.080149 | 0 | -4.809 | 69.1034 | 0 | 1.54
[6 empty] | 1 | 1.68 | 1.57 | 5 | 0 | 5.7815 | 0.23126 | 31.2583 | -0.078146 | 0 | -4.6887 | 69.1034 | 0 | 1.55
[6 empty] | 1 | 1.69 | 1.58 | 5 | 0 | 5.7619 | 0.23048 | 30.4768 | -0.076192 | 0 | -4.5715 | 69.1034 | 0 | 1.56
[6 empty] | 1 | 1.7 | 1.59 | 5 | 0 | 5.7429 | 0.22971 | 29.7149 | -0.074287 | 0 | -4.4572 | 69.1034 | 0 | 1.57
[6 empty] | 1 | 1.71 | 1.6 | 5 | 0 | 5.7243 | 0.22897 | 28.972 | -0.07243 | 0 | -4.3458 | 69.1034 | 0 | 1.58
[6 empty] | 1 | 1.72 | 1.61 | 5 | 0 | 5.7062 | 0.22825 | 28.2477 | -0.070619 | 0 | -4.2372 | 69.1034 | 0 | 1.59
[6 empty] | 1 | 1.73 | 1.62 | 5 | 0 | 5.6885 | 0.22754 | 27.5415 | -0.068854 | 0 | -4.1312 | 69.1034 | 0 | 1.6
[6 empty] | 1 | 1.74 | 1.63 | 5 | 0 | 5.6713 | 0.22685 | 26.853 | -0.067132 | 0 | -4.0279 | 69.1034 | 0 | 1.61
[6 empty] | 1 | 1.75 | 1.64 | 5 | 0 | 5.6545 | 0.22618 | 26.1817 | -0.065454 | 0 | -3.9272 | 69.1034 | 0 | 1.62
[6 empty] | 1 | 1.76 | 1.65 | 5 | 0 | 5.6382 | 0.22553 | 25.5271 | -0.063818 | 0 | -3.8291 | 69.1034 | 0 | 1.63
[6 empty] | 1 | 1.77 | 1.66 | 5 | 0 | 5.6222 | 0.22489 | 24.8889 | -0.062222 | 0 | -3.7333 | 69.1034 | 0 | 1.64
[6 empty] | 1 | 1.78 | 1.67 | 5 | 0 | 5.6067 | 0.22427 | 24.2667 | -0.060667 | 0 | -3.64 | 69.1034 | 0 | 1.65
[6 empty] | 1 | 1.79 | 1.68 | 5 | 0 | 5.5915 | 0.22366 | 23.66 | -0.05915 | 0 | -3.549 | 69.1034 | 0 | 1.66
[6 empty] | 1 | 1.8 | 1.69 | 5 | 0 | 5.5767 | 0.22307 | 23.0685 | -0.057671 | 0 | -3.4603 | 69.1034 | 0 | 1.67
[6 empty] | 1 | 1.81 | 1.7 | 5 | 0 | 5.5623 | 0.22249 | 22.4918 | -0.05623 | 0 | -3.3738 | 69.1034 | 0 | 1.68
[6 empty] | 1 | 1.82 | 1.71 | 5 | 0 | 5.5482 | 0.22193 | 21.9295 | -0.054824 | 0 | -3.2894 | 69.1034 | 0 | 1.69
[6 empty] | 1 | 1.83 | 1.72 | 5 | 0 | 5.5345 | 0.22138 | 21.3813 | -0.053453 | 0 | -3.2072 | 69.1034 | 0 | 1.7
[6 empty] | 1 | 1.84 | 1.73 | 5 | 0 | 5.5212 | 0.22085 | 20.8468 | -0.052117 | 0 | -3.127 | 69.1034 | 0 | 1.71
[6 empty] | 1 | 1.85 | 1.74 | 5 | 0 | 5.5081 | 0.22033 | 20.3256 | -0.050814 | 0 | -3.0488 | 69.1034 | 0 | 1.72
[6 empty] | 1 | 1.86 | 1.75 | 5 | 0 | 5.4954 | 0.21982 | 19.8175 | -0.049544 | 0 | -2.9726 | 69.1034 | 0 | 1.73
[6 empty] | 1 | 1.87 | 1.76 | 5 | 0 | 5.4831 | 0.21932 | 19.322 | -0.048305 | 0 | -2.8983 | 69.1034 | 0 | 1.74
[6 empty] | 1 | 1.88 | 1.77 | 5 | 0 | 5.471 | 0.21884 | 18.839 | -0.047097 | 0 | -2.8258 | 69.1034 | 0 | 1.75
[6 empty] | 1 | 1.89 | 1.78 | 5 | 0 | 5.4592 | 0.21837 | 18.368 | -0.04592 | 0 | -2.7552 | 69.1034 | 0 | 1.76
[6 empty] | 1 | 1.9 | 1.79 | 5 | 0 | 5.4477 | 0.21791 | 17.9088 | -0.044772 | 0 | -2.6863 | 69.1034 | 0 | 1.77
[6 empty] | 1 | 1.91 | 1.8 | 5 | 0 | 5.4365 | 0.21746 | 17.4611 | -0.043653 | 0 | -2.6192 | 69.1034 | 0 | 1.78
[6 empty] | 1 | 1.92 | 1.81 | 5 | 0 | 5.4256 | 0.21702 | 17.0245 | -0.042561 | 0 | -2.5537 | 69.1034 | 0 | 1.79
[6 empty] | 1 | 1.93 | 1.82 | 5 | 0 | 5.415 | 0.2166 | 16.5989 | -0.041497 | 0 | -2.4898 | 69.1034 | 0 | 1.8
[6 empty] | 1 | 1.94 | 1.83 | 5 | 0 | 5.4046 | 0.21618 | 16.184 | -0.04046 | 0 | -2.4276 | 69.1034 | 0 | 1.81
[6 empty] | 1 | 1.95 | 1.84 | 5 | 0 | 5.3945 | 0.21578 | 15.7794 | -0.039448 | 0 | -2.3669 | 69.1034 | 0 | 1.82
[6 empty] | 1 | 1.96 | 1.85 | 5 | 0 | 5.3846 | 0.21538 | 15.3849 | -0.038462 | 0 | -2.3077 | 69.1034 | 0 | 1.83
[6 empty] | 1 | 1.97 | 1.86 | 5 | 0 | 5.375 | 0.215 | 15.0003 | -0.037501 | 0 | -2.25 | 69.1034 | 0 | 1.84
[6 empty] | 1 | 1.98 | 1.87 | 5 | 0 | 5.3656 | 0.21463 | 14.6253 | -0.036563 | 0 | -2.1938 | 69.1034 | 0 | 1.85
[6 empty] | 1 | 1.99 | 1.88 | 5 | 0 | 5.3565 | 0.21426 | 14.2596 | -0.035649 | 0 | -2.1389 | 69.1034 | 0 | 1.86
[6 empty] | 1 | 2 | 1.89 | 5 | 0 | 5.3476 | 0.2139 | 13.9031 | -0.034758 | 0 | -2.0855 | 69.1034 | 0 | 1.87
[6 empty] | 1 | 2.01 | 1.9 | 5 | 0 | 5.3389 | 0.21356 | 13.5556 | -0.033889 | 0 | -2.0333 | 69.1034 | 0 | 1.88
[6 empty] | 1 | 2.02 | 1.91 | 5 | 0 | 5.3304 | 0.21322 | 13.2167 | -0.033042 | 0 | -1.9825 | 69.1034 | 0 | 1.89
[6 empty] | 1 | 2.03 | 1.92 | 5 | 0 | 5.3222 | 0.21289 | 12.8862 | -0.032216 | 0 | -1.9329 | 69.1034 | 0 | 1.9
[6 empty] | 1 | 2.04 | 1.93 | 5 | 0 | 5.3141 | 0.21256 | 12.5641 | -0.03141 | 0 | -1.8846 | 69.1034 | 0 | 1.91
[6 empty] | 1 | 2.05 | 1.94 | 5 | 0 | 5.3062 | 0.21225 | 12.25 | -0.030625 | 0 | -1.8375 | 69.1034 | 0 | 1.92
[6 empty] | 1 | 2.06 | 1.95 | 5 | 0 | 5.2986 | 0.21194 | 11.9437 | -0.029859 | 0 | -1.7916 | 69.1034 | 0 | 1.93
[6 empty] | 1 | 2.07 | 1.96 | 5 | 0 | 5.2911 | 0.21165 | 11.6451 | -0.029113 | 0 | -1.7468 | 69.1034 | 0 | 1.94
[6 empty] | 1 | 2.08 | 1.97 | 5 | 0 | 5.2839 | 0.21135 | 11.354 | -0.028385 | 0 | -1.7031 | 69.1034 | 0 | 1.95
[6 empty] | 1 | 2.09 | 1.98 | 5 | 0 | 5.2768 | 0.21107 | 11.0702 | -0.027675 | 0 | -1.6605 | 69.1034 | 0 | 1.96
[6 empty] | 1 | 2.1 | 1.99 | 5 | 0 | 5.2698 | 0.21079 | 10.7934 | -0.026984 | 0 | -1.619 | 69.1034 | 0 | 1.97
[6 empty] | 1 | 2.11 | 2 | 5 | 0 | 5.2631 | 0.21052 | 10.5236 | -0.026309 | 0 | -1.5785 | 69.1034 | 0 | 1.98
[6 empty] | 1 | 2.12 | 2.01 | 5 | 0 | 5.2565 | 0.21026 | 10.2605 | -0.025651 | 0 | -1.5391 | 69.1034 | 0 | 1.99
[6 empty] | 1 | 2.13 | 2.02 | 5 | 0 | 5.2501 | 0.21 | 10.004 | -0.02501 | 0 | -1.5006 | 69.1034 | 0 | 2
[6 empty] | 1 | 2.14 | 2.03 | 5 | 0 | 5.2438 | 0.20975 | 9.7539 | -0.024385 | 0 | -1.4631 | 69.1034 | 0 | 2.01
[6 empty] | 1 | 2.15 | 2.04 | 5 | 0 | 5.2378 | 0.20951 | 9.51 | -0.023775 | 0 | -1.4265 | 69.1034 | 0 | 2.02
[6 empty] | 1 | 2.16 | 2.05 | 5 | 0 | 5.2318 | 0.20927 | 9.2723 | -0.023181 | 0 | -1.3908 | 69.1034 | 0 | 2.03
[6 empty] | 1 | 2.17 | 2.06 | 5 | 0 | 5.226 | 0.20904 | 9.0405 | -0.022601 | 0 | -1.3561 | 69.1034 | 0 | 2.04
[6 empty] | 1 | 2.18 | 2.07 | 5 | 0 | 5.2204 | 0.20881 | 8.8145 | -0.022036 | 0 | -1.3222 | 69.1034 | 0 | 2.05
[6 empty] | 1 | 2.19 | 2.08 | 5 | 0 | 5.2149 | 0.20859 | 8.5941 | -0.021485 | 0 | -1.2891 | 69.1034 | 0 | 2.06
[6 empty] | 1 | 2.2 | 2.09 | 5 | 0 | 5.2095 | 0.20838 | 8.3792 | -0.020948 | 0 | -1.2569 | 69.1034 | 0 | 2.07
[6 empty] | 1 | 2.21 | 2.1 | 5 | 0 | 5.2042 | 0.20817 | 8.1698 | -0.020424 | 0 | -1.2255 | 69.1034 | 0 | 2.08
[6 empty] | 1 | 2.22 | 2.11 | 5 | 0 | 5.1991 | 0.20797 | 7.9655 | -0.019914 | 0 | -1.1948 | 69.1034 | 0 | 2.09
[6 empty] | 1 | 2.23 | 2.12 | 5 | 0 | 5.1942 | 0.20777 | 7.7664 | -0.019416 | 0 | -1.165 | 69.1034 | 0 | 2.1
[6 empty] | 1 | 2.24 | 2.13 | 5 | 0 | 5.1893 | 0.20757 | 7.5722 | -0.018931 | 0 | -1.1358 | 69.1034 | 0 | 2.11
[6 empty] | 1 | 2.25 | 2.14 | 5 | 0 | 5.1846 | 0.20738 | 7.3829 | -0.018457 | 0 | -1.1074 | 69.1034 | 0 | 2.12
[6 empty] | 1 | 2.26 | 2.15 | 5 | 0 | 5.18 | 0.2072 | 7.1983 | -0.017996 | 0 | -1.0798 | 69.1034 | 0 | 2.13
[6 empty] | 1 | 2.27 | 2.16 | 5 | 0 | 5.1755 | 0.20702 | 7.0184 | -0.017546 | 0 | -1.0528 | 69.1034 | 0 | 2.14
[6 empty] | 1 | 2.28 | 2.17 | 5 | 0 | 5.1711 | 0.20684 | 6.8429 | -0.017107 | 0 | -1.0264 | 69.1034 | 0 | 2.15
[6 empty] | 1 | 2.29 | 2.18 | 5 | 0 | 5.1668 | 0.20667 | 6.6719 | -0.01668 | 0 | -1.0008 | 69.1034 | 0 | 2.16
[6 empty] | 1 | 2.3 | 2.19 | 5 | 0 | 5.1626 | 0.20651 | 6.5051 | -0.016263 | 0 | -0.97576 | 69.1034 | 0 | 2.17
[6 empty] | 1 | 2.31 | 2.2 | 5 | 0 | 5.1586 | 0.20634 | 6.3424 | -0.015856 | 0 | -0.95136 | 69.1034 | 0 | 2.18
[6 empty] | 1 | 2.32 | 2.21 | 5 | 0 | 5.1546 | 0.20618 | 6.1839 | -0.01546 | 0 | -0.92758 | 69.1034 | 0 | 2.19
[6 empty] | 1 | 2.33 | 2.22 | 5 | 0 | 5.1507 | 0.20603 | 6.0293 | -0.015073 | 0 | -0.90439 | 69.1034 | 0 | 2.2
[6 empty] | 1 | 2.34 | 2.23 | 5 | 0 | 5.147 | 0.20588 | 5.8785 | -0.014696 | 0 | -0.88178 | 69.1034 | 0 | 2.21
[6 empty] | 1 | 2.35 | 2.24 | 5 | 0 | 5.1433 | 0.20573 | 5.7316 | -0.014329 | 0 | -0.85974 | 69.1034 | 0 | 2.22
[6 empty] | 1 | 2.36 | 2.25 | 5 | 0 | 5.1397 | 0.20559 | 5.5883 | -0.013971 | 0 | -0.83824 | 69.1034 | 0 | 2.23
[6 empty] | 1 | 2.37 | 2.26 | 5 | 0 | 5.1362 | 0.20545 | 5.4486 | -0.013621 | 0 | -0.81729 | 69.1034 | 0 | 2.24
[6 empty] | 1 | 2.38 | 2.27 | 5 | 0 | 5.1328 | 0.20531 | 5.3124 | -0.013281 | 0 | -0.79685 | 69.1034 | 0 | 2.25
[6 empty] | 1 | 2.39 | 2.28 | 5 | 0 | 5.1295 | 0.20518 | 5.1796 | -0.012949 | 0 | -0.77693 | 69.1034 | 0 | 2.26
[6 empty] | 1 | 2.4 | 2.29 | 5 | 0 | 5.1263 | 0.20505 | 5.0501 | -0.012625 | 0 | -0.75751 | 69.1034 | 0 | 2.27
[6 empty] | 1 | 2.41 | 2.3 | 5 | 0 | 5.1231 | 0.20492 | 4.9238 | -0.01231 | 0 | -0.73857 | 69.1034 | 0 | 2.28
[6 empty] | 1 | 2.42 | 2.31 | 5 | 0 | 5.12 | 0.2048 | 4.8007 | -0.012002 | 0 | -0.72011 | 69.1034 | 0 | 2.29
[6 empty] | 1 | 2.43 | 2.32 | 5 | 0 | 5.117 | 0.20468 | 4.6807 | -0.011702 | 0 | -0.70211 | 69.1034 | 0 | 2.3
[6 empty] | 1 | 2.44 | 2.33 | 5 | 0 | 5.1141 | 0.20456 | 4.5637 | -0.011409 | 0 | -0.68455 | 69.1034 | 0 | 2.31
[6 empty] | 1 | 2.45 | 2.34 | 5 | 0 | 5.1112 | 0.20445 | 4.4496 | -0.011124 | 0 | -0.66744 | 69.1034 | 0 | 2.32
[6 empty] | 1 | 2.46 | 2.35 | 5 | 0 | 5.1085 | 0.20434 | 4.3384 | -0.010846 | 0 | -0.65075 | 69.1034 | 0 | 2.33
[6 empty] | 1 | 2.47 | 2.36 | 5 | 0 | 5.1057 | 0.20423 | 4.2299 | -0.010575 | 0 | -0.63448 | 69.1034 | 0 | 2.34
[6 empty] | 1 | 2.48 | 2.37 | 5 | 0 | 5.1031 | 0.20412 | 4.1241 | -0.01031 | 0 | -0.61862 | 69.1034 | 0 | 2.35
[6 empty] | 1 | 2.49 | 2.38 | 5 | 0 | 5.1005 | 0.20402 | 4.021 | -0.010053 | 0 | -0.60316 | 69.1034 | 0 | 2.36
[6 empty] | 1 | 2.5 | 2.39 | 5 | 0 | 5.098 | 0.20392 | 3.9205 | -0.0098013 | 0 | -0.58808 | 69.1034 | 0 | 2.37
[6 empty] | 1 | 2.51 | 2.4 | 5 | 0 | 5.0956 | 0.20382 | 3.8225 | -0.0095563 | 0 | -0.57338 | 69.1034 | 0 | 2.38
[6 empty] | 1 | 2.52 | 2.41 | 5 | 0 | 5.0932 | 0.20373 | 3.7269 | -0.0093174 | 0 | -0.55904 | 69.1034 | 0 | 2.39
[6 empty] | 1 | 2.53 | 2.42 | 5 | 0 | 5.0908 | 0.20363 | 3.6338 | -0.0090844 | 0 | -0.54507 | 69.1034 | 0 | 2.4
[6 empty] | 1 | 2.54 | 2.43 | 5 | 0 | 5.0886 | 0.20354 | 3.5429 | -0.0088573 | 0 | -0.53144 | 69.1034 | 0 | 2.41
[6 empty] | 1 | 2.55 | 2.44 | 5 | 0 | 5.0864 | 0.20345 | 3.4544 | -0.0086359 | 0 | -0.51815 | 69.1034 | 0 | 2.42
[6 empty] | 1 | 2.56 | 2.45 | 5 | 0 | 5.0842 | 0.20337 | 3.368 | -0.00842 | 0 | -0.5052 | 69.1034 | 0 | 2.43
[6 empty] | 1 | 2.57 | 2.46 | 5 | 0 | 5.0821 | 0.20328 | 3.2838 | -0.0082095 | 0 | -0.49257 | 69.1034 | 0 | 2.44
[6 empty] | 1 | 2.58 | 2.47 | 5 | 0 | 5.08 | 0.2032 | 3.2017 | -0.0080042 | 0 | -0.48025 | 69.1034 | 0 | 2.45
[6 empty] | 1 | 2.59 | 2.48 | 5 | 0 | 5.078 | 0.20312 | 3.1217 | -0.0078041 | 0 | -0.46825 | 69.1034 | 0 | 2.46
[6 empty] | 1 | 2.6 | 2.49 | 5 | 0 | 5.0761 | 0.20304 | 3.0436 | -0.007609 | 0 | -0.45654 | 69.1034 | 0 | 2.47
[6 empty] | 1 | 2.61 | 2.5 | 5 | 0 | 5.0742 | 0.20297 | 2.9675 | -0.0074188 | 0 | -0.44513 | 69.1034 | 0 | 2.48
[6 empty] | 1 | 2.62 | 2.51 | 5 | 0 | 5.0723 | 0.20289 | 2.8933 | -0.0072333 | 0 | -0.434 | 69.1034 | 0 | 2.49
[6 empty] | 1 | 2.63 | 2.52 | 5 | 0 | 5.0705 | 0.20282 | 2.821 | -0.0070525 | 0 | -0.42315 | 69.1034 | 0 | 2.5
[6 empty] | 1 | 2.64 | 2.53 | 5 | 0 | 5.0688 | 0.20275 | 2.7505 | -0.0068762 | 0 | -0.41257 | 69.1034 | 0 | 2.51
[6 empty] | 1 | 2.65 | 2.54 | 5 | 0 | 5.067 | 0.20268 | 2.6817 | -0.0067043 | 0 | -0.40226 | 69.1034 | 0 | 2.52
[6 empty] | 1 | 2.66 | 2.55 | 5 | 0 | 5.0654 | 0.20261 | 2.6147 | -0.0065367 | 0 | -0.3922 | 69.1034 | 0 | 2.53
[6 empty] | 1 | 2.67 | 2.56 | 5 | 0 | 5.0637 | 0.20255 | 2.5493 | -0.0063733 | 0 | -0.3824 | 69.1034 | 0 | 2.54
[6 empty] | 1 | 2.68 | 2.57 | 5 | 0 | 5.0621 | 0.20249 | 2.4856 | -0.0062139 | 0 | -0.37284 | 69.1034 | 0 | 2.55
[6 empty] | 1 | 2.69 | 2.58 | 5 | 0 | 5.0606 | 0.20242 | 2.4234 | -0.0060586 | 0 | -0.36351 | 69.1034 | 0 | 2.56
[6 empty] | 1 | 2.7 | 2.59 | 5 | 0 | 5.0591 | 0.20236 | 2.3628 | -0.0059071 | 0 | -0.35443 | 69.1034 | 0 | 2.57
[6 empty] | 1 | 2.71 | 2.6 | 5 | 0 | 5.0576 | 0.2023 | 2.3038 | -0.0057594 | 0 | -0.34557 | 69.1034 | 0 | 2.58
[6 empty] | 1 | 2.72 | 2.61 | 5 | 0 | 5.0562 | 0.20225 | 2.2462 | -0.0056155 | 0 | -0.33693 | 69.1034 | 0 | 2.59
[6 empty] | 1 | 2.73 | 2.62 | 5 | 0 | 5.0548 | 0.20219 | 2.19 | -0.0054751 | 0 | -0.3285 | 69.1034 | 0 | 2.6
[6 empty] | 1 | 2.74 | 2.63 | 5 | 0 | 5.0534 | 0.20214 | 2.1353 | -0.0053382 | 0 | -0.32029 | 69.1034 | 0 | 2.61
[6 empty] | 1 | 2.75 | 2.64 | 5 | 0 | 5.052 | 0.20208 | 2.0819 | -0.0052047 | 0 | -0.31228 | 69.1034 | 0 | 2.62
[6 empty] | 1 | 2.76 | 2.65 | 5 | 0 | 5.0507 | 0.20203 | 2.0298 | -0.0050746 | 0 | -0.30448 | 69.1034 | 0 | 2.63
[6 empty] | 1 | 2.77 | 2.66 | 5 | 0 | 5.0495 | 0.20198 | 1.9791 | -0.0049478 | 0 | -0.29687 | 69.1034 | 0 | 2.64
[6 empty] | 1 | 2.78 | 2.67 | 5 | 0 | 5.0482 | 0.20193 | 1.9296 | -0.0048241 | 0 | -0.28944 | 69.1034 | 0 | 2.65
[6 empty] | 1 | 2.79 | 2.68 | 5 | 0 | 5.047 | 0.20188 | 1.8814 | -0.0047035 | 0 | -0.28221 | 69.1034 | 0 | 2.66
[6 empty] | 1 | 2.8 | 2.69 | 5 | 0 | 5.0459 | 0.20183 | 1.8343 | -0.0045859 | 0 | -0.27515 | 69.1034 | 0 | 2.67
[6 empty] | 1 | 2.81 | 2.7 | 5 | 0 | 5.0447 | 0.20179 | 1.7885 | -0.0044712 | 0 | -0.26827 | 69.1034 | 0 | 2.68
[6 empty] | 1 | 2.82 | 2.71 | 5 | 0 | 5.0436 | 0.20174 | 1.7438 | -0.0043594 | 0 | -0.26157 | 69.1034 | 0 | 2.69
[6 empty] | 1 | 2.83 | 2.72 | 5 | 0 | 5.0425 | 0.2017 | 1.7002 | -0.0042505 | 0 | -0.25503 | 69.1034 | 0 | 2.7
[6 empty] | 1 | 2.84 | 2.73 | 5 | 0 | 5.0414 | 0.20166 | 1.6577 | -0.0041442 | 0 | -0.24865 | 69.1034 | 0 | 2.71
[6 empty] | 1 | 2.85 | 2.74 | 5 | 0 | 5.0404 | 0.20162 | 1.6162 | -0.0040406 | 0 | -0.24244 | 69.1034 | 0 | 2.72
[6 empty] | 1 | 2.86 | 2.75 | 5 | 0 | 5.0394 | 0.20158 | 1.5758 | -0.0039396 | 0 | -0.23637 | 69.1034 | 0 | 2.73
[6 empty] | 1 | 2.87 | 2.76 | 5 | 0 | 5.0384 | 0.20154 | 1.5364 | -0.0038411 | 0 | -0.23047 | 69.1034 | 0 | 2.74
[6 empty] | 1 | 2.88 | 2.77 | 5 | 0 | 5.0375 | 0.2015 | 1.498 | -0.0037451 | 0 | -0.2247 | 69.1034 | 0 | 2.75
[6 empty] | 1 | 2.89 | 2.78 | 5 | 0 | 5.0365 | 0.20146 | 1.4606 | -0.0036514 | 0 | -0.21909 | 69.1034 | 0 | 2.76
[6 empty] | 1 | 2.9 | 2.79 | 5 | 0 | 5.0356 | 0.20142 | 1.4241 | -0.0035601 | 0 | -0.21361 | 69.1034 | 0 | 2.77
[6 empty] | 1 | 2.91 | 2.8 | 5 | 0 | 5.0347 | 0.20139 | 1.3885 | -0.0034711 | 0 | -0.20827 | 69.1034 | 0 | 2.78
[6 empty] | 1 | 2.92 | 2.81 | 5 | 0 | 5.0338 | 0.20135 | 1.3537 | -0.0033844 | 0 | -0.20306 | 69.1034 | 0 | 2.79
[6 empty] | 1 | 2.93 | 2.82 | 5 | 0 | 5.033 | 0.20132 | 1.3199 | -0.0032998 | 0 | -0.19799 | 69.1034 | 0 | 2.8
[6 empty] | 1 | 2.94 | 2.83 | 5 | 0 | 5.0322 | 0.20129 | 1.2869 | -0.0032173 | 0 | -0.19304 | 69.1034 | 0 | 2.81
[6 empty] | 1 | 2.95 | 2.84 | 5 | 0 | 5.0314 | 0.20125 | 1.2547 | -0.0031368 | 0 | -0.18821 | 69.1034 | 0 | 2.82
[6 empty] | 1 | 2.96 | 2.85 | 5 | 0 | 5.0306 | 0.20122 | 1.2234 | -0.0030584 | 0 | -0.1835 | 69.1034 | 0 | 2.83
[6 empty] | 1 | 2.97 | 2.86 | 5 | 0 | 5.0298 | 0.20119 | 1.1928 | -0.0029819 | 0 | -0.17892 | 69.1034 | 0 | 2.84
[6 empty] | 1 | 2.98 | 2.87 | 5 | 0 | 5.0291 | 0.20116 | 1.163 | -0.0029074 | 0 | -0.17444 | 69.1034 | 0 | 2.85
[6 empty] | 1 | 2.99 | 2.88 | 5 | 0 | 5.0283 | 0.20113 | 1.1339 | -0.0028347 | 0 | -0.17008 | 69.1034 | 0 | 2.86
[6 empty] | 1 | 3 | 2.89 | 5 | 0 | 5.0276 | 0.20111 | 1.1055 | -0.0027638 | 0 | -0.16583 | 69.1034 | 0 | 2.87
Menu

## Menu
SIVA
Simulator Interface for Ventilation Analysis
V11.1.21
VC-CMV
PC-CMV
PC-CSV
Resources
Equations
Models
Exercises
Classify Modes
Vent Mode Map
Compare Vents
Glossary
About
References

## Menu (PC)
Options for Pressure Control
Inspiratory Pressure Target relative to body surface pressure
Inspiratory Pressure Target
relative to PEEP

## Menu (Ex)
Options for Exercises
Volume Control
Pressure Control

## VC (sq)
Volume Control
Lung Model
|  | Elastance (cm H2O/L) | 20
|  | Compliance (L/cm H2O) | 0.05 |  | 50
|  | Resistance (cm H2O/L/s) | 10 |  | 10
|  | Time Constant (s) | 0.5
|  | Tidal Lung Pressure (cm H2O) | 8.4
Effort Model
|  | Delay Before Effort (s) | 0 |  | 0
|  | Duration of Effort (s) | 0.1 |  | 2
|  | Freq (hz) | 2.5
|  | Period (s) | 0.2
|  | max Pmus (cm H2O) | 0 | 49 | 0
Ventilator Settings
Flow Waveform
|  | Flow (square wave, L/min) | 14 |  | 14
|  | Peak Flow (L/s) | 0.23333
|  | Tidal Volume (mL) | 420 |  | 420
|  | Tidal Volume (L) | 0.42
|  | Frequency (breaths/min) | 17 |  | 17
|  | PEEP (cm H2O) | 5 |  | 5
|  | autoPEEP (cm H2O) | 0.27291
|  | Pause Time (s) | 0 |  | 0
|  | Peak Airway Pressure (cm H2O) | 16.0062
|  | Total Cycle Time (s) | 3.5294
|  | Inspirtory Time (s) | 1.8
|  | Expiratory Time (s) | 1.7294
|  | I:E | 1:1
|  | Duty Cycle (%) | 51
|  | Plateau Pressure (cm H2O) | 13.6729
|  | Tidal (Driving) Pressure (cm H2O) | 8.4
|  | Peak Inspiratory Flow (L/min) | 14
|  | Peak Expiratory Flow (L/min) | -52.3174
|  | Mean Inspiratory Pressure (cm H2O) | 11.8062
|  | Mean Airway Pressure (cm H2O) | 8.4899
|  | Inspiratory Flow Time (s) | 1.8
Passive |  | Ventilator Elastic Work (J) | 0.3903
|  | Ventilator Resistive Work (J) | 0.096138
|  | Total Work (J) | 0.48644
Active |  | Ventilator Work (J) | 0.48914
|  | Patient Work (J) | 0
|  | Total Work (J) | 0.48914
|  | Work Shift Index (%) | 0
|  | Ventilator Inspiratory Power (W) | 0.13782
|  | Ventilator Inspiratory Power (J/min) | 8.2695
[20 empty] | Peak | Elastance
[20 empty] | 16.0062 | 0.02 | …………………….Pressure Volume Curve Data……………….
[18 empty] | Muscle | Muscle | total | [5 empty] | Seg |  | Seg |  | Seg |  | Seg |  | Muscle
[8 empty] | Pressure | Lung Volume | Volume | Flow | Airway | Mean | Pat. | Vent | Absolute | Relative | Pressure | volume | Pressure |  | Pstat |  | Pdyn |  | L1 |  | L2 |  | U1 |  | U2 |  | flow
[6 empty] | BreathTime | Time | (cm H2O) | (mL) | (mL) | (L/m) | Press | Press | Work [J] | Work [J] | Time | Time | (cm H2O) | mL | (cm H2O) |  | X | Y | X | Y | x | y | x | y | x | y | x | y | (L/s) | Time
[6 empty] | 0 | 0 | 7.6062 | 263.6454 | 0 | 14 | 7.6062 | 7.6062 | 0 | 0.0017411 | 0 | 0 | 0 | 0 | 7.6062 |  | 5 | 0 | 7.6062 | 0 | 0 | 0 | 5 | 0 | 0 | 420 | 13.4 | 420 | 0 | 0
[6 empty] | 0.01 | 0.01 | 7.6529 | 265.9787 | 2.3333 | 14 | 7.6529 | 7.6296 | 0 | 0.0034928 | 0.01 | 0.01 | 0 | 0 | 7.6529 |  | 13.4 | 420 | 16.0062 | 420 | 7.6062 | 0 | 7.6062 | 0 | 13.4 | 420 | 16.0062 | 420 | 0 | 0.01
[6 empty] | 0.02 | 0.02 | 7.6996 | 268.3121 | 4.6667 | 14 | 7.6996 | 7.6529 | 0 | 0.0052553 | 0.02 | 0.02 | 0 | 0 | 7.6996 | [13 empty] | 0 | 0.02
[6 empty] | 0.03 | 0.03 | 7.7462 | 270.6454 | 7 | 14 | 7.7462 | 7.6762 | 0 | 0.0070284 | 0.03 | 0.03 | 0 | 0 | 7.7462 | [13 empty] | 0 | 0.03
[6 empty] | 0.04 | 0.04 | 7.7929 | 272.9787 | 9.3333 | 14 | 7.7929 | 7.6996 | 0 | 0.0088122 | 0.04 | 0.04 | 0 | 0 | 7.7929 | [13 empty] | 0 | 0.04
[6 empty] | 0.05 | 0.05 | 7.8396 | 275.3121 | 11.6667 | 14 | 7.8396 | 7.7229 | 0 | 0.010607 | 0.05 | 0.05 | 0 | 0 | 7.8396 | [13 empty] | 0 | 0.05
[6 empty] | 0.06 | 0.06 | 7.8862 | 277.6454 | 14 | 14 | 7.8862 | 7.7462 | 0 | 0.012412 | 0.06 | 0.06 | 0 | 0 | 7.8862 | [13 empty] | 0 | 0.06
[6 empty] | 0.07 | 0.07 | 7.9329 | 279.9787 | 16.3333 | 14 | 7.9329 | 7.7696 | 0 | 0.014228 | 0.07 | 0.07 | 0 | 0 | 7.9329 | [13 empty] | 0 | 0.07
[6 empty] | 0.08 | 0.08 | 7.9796 | 282.3121 | 18.6667 | 14 | 7.9796 | 7.7929 | 0 | 0.016054 | 0.08 | 0.08 | 0 | 0 | 7.9796 | [13 empty] | 0 | 0.08
[6 empty] | 0.09 | 0.09 | 8.0262 | 284.6454 | 21 | 14 | 8.0262 | 7.8162 | 0 | 0.017891 | 0.09 | 0.09 | 0 | 0 | 8.0262 | [13 empty] | 0 | 0.09
[6 empty] | 0.1 | 0.1 | 8.0729 | 286.9787 | 23.3333 | 14 | 8.0729 | 7.8396 | 0 | 0.019739 | 0.1 | 0.1 | 0 | 0 | 8.0729 | [13 empty] | 0 | 0.1
[6 empty] | 0.11 | 0.11 | 8.1196 | 289.3121 | 25.6667 | 14 | 8.1196 | 7.8629 | 0 | 0.021598 | 0.11 | 0.11 | 0 | 0 | 8.1196 | [13 empty] | 0 | 0.11
[6 empty] | 0.12 | 0.12 | 8.1662 | 291.6454 | 28 | 14 | 8.1662 | 7.8862 | 0 | 0.023467 | 0.12 | 0.12 | 0 | 0 | 8.1662 | [13 empty] | 0 | 0.12
[6 empty] | 0.13 | 0.13 | 8.2129 | 293.9787 | 30.3333 | 14 | 8.2129 | 7.9096 | 0 | 0.025347 | 0.13 | 0.13 | 0 | 0 | 8.2129 | [13 empty] | 0 | 0.13
[6 empty] | 0.14 | 0.14 | 8.2596 | 296.3121 | 32.6667 | 14 | 8.2596 | 7.9329 | 0 | 0.027238 | 0.14 | 0.14 | 0 | 0 | 8.2596 | [13 empty] | 0 | 0.14
[6 empty] | 0.15 | 0.15 | 8.3062 | 298.6454 | 35 | 14 | 8.3062 | 7.9562 | 0 | 0.029139 | 0.15 | 0.15 | 0 | 0 | 8.3062 | [13 empty] | 0 | 0.15
[6 empty] | 0.16 | 0.16 | 8.3529 | 300.9787 | 37.3333 | 14 | 8.3529 | 7.9796 | 0 | 0.031051 | 0.16 | 0.16 | 0 | 0 | 8.3529 | [13 empty] | 0 | 0.16
[6 empty] | 0.17 | 0.17 | 8.3996 | 303.3121 | 39.6667 | 14 | 8.3996 | 8.0029 | 0 | 0.032974 | 0.17 | 0.17 | 0 | 0 | 8.3996 | [13 empty] | 0 | 0.17
[6 empty] | 0.18 | 0.18 | 8.4462 | 305.6454 | 42 | 14 | 8.4462 | 8.0262 | 0 | 0.034907 | 0.18 | 0.18 | 0 | 0 | 8.4462 | [13 empty] | 0 | 0.18
[6 empty] | 0.19 | 0.19 | 8.4929 | 307.9787 | 44.3333 | 14 | 8.4929 | 8.0496 | 0 | 0.036851 | 0.19 | 0.19 | 0 | 0 | 8.4929 | [13 empty] | 0 | 0.19
[6 empty] | 0.2 | 0.2 | 8.5396 | 310.3121 | 46.6667 | 14 | 8.5396 | 8.0729 | 0 | 0.038806 | 0.2 | 0.2 | 0 | 0 | 8.5396 | [13 empty] | 0 | 0.2
[6 empty] | 0.21 | 0.21 | 8.5862 | 312.6454 | 49 | 14 | 8.5862 | 8.0962 | 0 | 0.040771 | 0.21 | 0.21 | 0 | 0 | 8.5862 | [13 empty] | 0 | 0.21
[6 empty] | 0.22 | 0.22 | 8.6329 | 314.9787 | 51.3333 | 14 | 8.6329 | 8.1196 | 0 | 0.042747 | 0.22 | 0.22 | 0 | 0 | 8.6329 | [13 empty] | 0 | 0.22
[6 empty] | 0.23 | 0.23 | 8.6796 | 317.3121 | 53.6667 | 14 | 8.6796 | 8.1429 | 0 | 0.044734 | 0.23 | 0.23 | 0 | 0 | 8.6796 | [13 empty] | 0 | 0.23
[6 empty] | 0.24 | 0.24 | 8.7262 | 319.6454 | 56 | 14 | 8.7262 | 8.1662 | 0 | 0.046731 | 0.24 | 0.24 | 0 | 0 | 8.7262 | [13 empty] | 0 | 0.24
[6 empty] | 0.25 | 0.25 | 8.7729 | 321.9787 | 58.3333 | 14 | 8.7729 | 8.1896 | 0 | 0.048739 | 0.25 | 0.25 | 0 | 0 | 8.7729 | [13 empty] | 0 | 0.25
[6 empty] | 0.26 | 0.26 | 8.8196 | 324.3121 | 60.6667 | 14 | 8.8196 | 8.2129 | 0 | 0.050758 | 0.26 | 0.26 | 0 | 0 | 8.8196 | [13 empty] | 0 | 0.26
[6 empty] | 0.27 | 0.27 | 8.8662 | 326.6454 | 63 | 14 | 8.8662 | 8.2362 | 0 | 0.052788 | 0.27 | 0.27 | 0 | 0 | 8.8662 | [13 empty] | 0 | 0.27
[6 empty] | 0.28 | 0.28 | 8.9129 | 328.9787 | 65.3333 | 14 | 8.9129 | 8.2596 | 0 | 0.054828 | 0.28 | 0.28 | 0 | 0 | 8.9129 | [13 empty] | 0 | 0.28
[6 empty] | 0.29 | 0.29 | 8.9596 | 331.3121 | 67.6667 | 14 | 8.9596 | 8.2829 | 0 | 0.056879 | 0.29 | 0.29 | 0 | 0 | 8.9596 | [13 empty] | 0 | 0.29
[6 empty] | 0.3 | 0.3 | 9.0062 | 333.6454 | 70 | 14 | 9.0062 | 8.3062 | 0 | 0.05894 | 0.3 | 0.3 | 0 | 0 | 9.0062 | [13 empty] | 0 | 0.3
[6 empty] | 0.31 | 0.31 | 9.0529 | 335.9787 | 72.3333 | 14 | 9.0529 | 8.3296 | 0 | 0.061012 | 0.31 | 0.31 | 0 | 0 | 9.0529 | [13 empty] | 0 | 0.31
[6 empty] | 0.32 | 0.32 | 9.0996 | 338.3121 | 74.6667 | 14 | 9.0996 | 8.3529 | 0 | 0.063095 | 0.32 | 0.32 | 0 | 0 | 9.0996 | [13 empty] | 0 | 0.32
[6 empty] | 0.33 | 0.33 | 9.1462 | 340.6454 | 77 | 14 | 9.1462 | 8.3762 | 0 | 0.065189 | 0.33 | 0.33 | 0 | 0 | 9.1462 | [13 empty] | 0 | 0.33
[6 empty] | 0.34 | 0.34 | 9.1929 | 342.9787 | 79.3333 | 14 | 9.1929 | 8.3996 | 0 | 0.067293 | 0.34 | 0.34 | 0 | 0 | 9.1929 | [13 empty] | 0 | 0.34
[6 empty] | 0.35 | 0.35 | 9.2396 | 345.3121 | 81.6667 | 14 | 9.2396 | 8.4229 | 0 | 0.069408 | 0.35 | 0.35 | 0 | 0 | 9.2396 | [13 empty] | 0 | 0.35
[6 empty] | 0.36 | 0.36 | 9.2862 | 347.6454 | 84 | 14 | 9.2862 | 8.4462 | 0 | 0.071534 | 0.36 | 0.36 | 0 | 0 | 9.2862 | [13 empty] | 0 | 0.36
[6 empty] | 0.37 | 0.37 | 9.3329 | 349.9787 | 86.3333 | 14 | 9.3329 | 8.4696 | 0 | 0.07367 | 0.37 | 0.37 | 0 | 0 | 9.3329 | [13 empty] | 0 | 0.37
[6 empty] | 0.38 | 0.38 | 9.3796 | 352.3121 | 88.6667 | 14 | 9.3796 | 8.4929 | 0 | 0.075817 | 0.38 | 0.38 | 0 | 0 | 9.3796 | [13 empty] | 0 | 0.38
[6 empty] | 0.39 | 0.39 | 9.4262 | 354.6454 | 91 | 14 | 9.4262 | 8.5162 | 0 | 0.077975 | 0.39 | 0.39 | 0 | 0 | 9.4262 | [13 empty] | 0 | 0.39
[6 empty] | 0.4 | 0.4 | 9.4729 | 356.9787 | 93.3333 | 14 | 9.4729 | 8.5396 | 0 | 0.080143 | 0.4 | 0.4 | 0 | 0 | 9.4729 | [13 empty] | 0 | 0.4
[6 empty] | 0.41 | 0.41 | 9.5196 | 359.3121 | 95.6667 | 14 | 9.5196 | 8.5629 | 0 | 0.082322 | 0.41 | 0.41 | 0 | 0 | 9.5196 | [13 empty] | 0 | 0.41
[6 empty] | 0.42 | 0.42 | 9.5662 | 361.6454 | 98 | 14 | 9.5662 | 8.5862 | 0 | 0.084512 | 0.42 | 0.42 | 0 | 0 | 9.5662 | [13 empty] | 0 | 0.42
[6 empty] | 0.43 | 0.43 | 9.6129 | 363.9787 | 100.3333 | 14 | 9.6129 | 8.6096 | 0 | 0.086712 | 0.43 | 0.43 | 0 | 0 | 9.6129 | [13 empty] | 0 | 0.43
[6 empty] | 0.44 | 0.44 | 9.6596 | 366.3121 | 102.6667 | 14 | 9.6596 | 8.6329 | 0 | 0.088923 | 0.44 | 0.44 | 0 | 0 | 9.6596 | [13 empty] | 0 | 0.44
[6 empty] | 0.45 | 0.45 | 9.7062 | 368.6454 | 105 | 14 | 9.7062 | 8.6562 | 0 | 0.091145 | 0.45 | 0.45 | 0 | 0 | 9.7062 | [13 empty] | 0 | 0.45
[6 empty] | 0.46 | 0.46 | 9.7529 | 370.9787 | 107.3333 | 14 | 9.7529 | 8.6796 | 0 | 0.093377 | 0.46 | 0.46 | 0 | 0 | 9.7529 | [13 empty] | 0 | 0.46
[6 empty] | 0.47 | 0.47 | 9.7996 | 373.3121 | 109.6667 | 14 | 9.7996 | 8.7029 | 0 | 0.095621 | 0.47 | 0.47 | 0 | 0 | 9.7996 | [13 empty] | 0 | 0.47
[6 empty] | 0.48 | 0.48 | 9.8462 | 375.6454 | 112 | 14 | 9.8462 | 8.7262 | 0 | 0.097874 | 0.48 | 0.48 | 0 | 0 | 9.8462 | [13 empty] | 0 | 0.48
[6 empty] | 0.49 | 0.49 | 9.8929 | 377.9787 | 114.3333 | 14 | 9.8929 | 8.7496 | 0 | 0.10014 | 0.49 | 0.49 | 0 | 0 | 9.8929 | [13 empty] | 0 | 0.49
[6 empty] | 0.5 | 0.5 | 9.9396 | 380.3121 | 116.6667 | 14 | 9.9396 | 8.7729 | 0 | 0.10241 | 0.5 | 0.5 | 0 | 0 | 9.9396 | [13 empty] | 0 | 0.5
[6 empty] | 0.51 | 0.51 | 9.9862 | 382.6454 | 119 | 14 | 9.9862 | 8.7962 | 0 | 0.1047 | 0.51 | 0.51 | 0 | 0 | 9.9862 | [13 empty] | 0 | 0.51
[6 empty] | 0.52 | 0.52 | 10.0329 | 384.9787 | 121.3333 | 14 | 10.0329 | 8.8196 | 0 | 0.107 | 0.52 | 0.52 | 0 | 0 | 10.0329 | [13 empty] | 0 | 0.52
[6 empty] | 0.53 | 0.53 | 10.0796 | 387.3121 | 123.6667 | 14 | 10.0796 | 8.8429 | 0 | 0.1093 | 0.53 | 0.53 | 0 | 0 | 10.0796 | [13 empty] | 0 | 0.53
[6 empty] | 0.54 | 0.54 | 10.1262 | 389.6454 | 126 | 14 | 10.1262 | 8.8662 | 0 | 0.11162 | 0.54 | 0.54 | 0 | 0 | 10.1262 | [13 empty] | 0 | 0.54
[6 empty] | 0.55 | 0.55 | 10.1729 | 391.9787 | 128.3333 | 14 | 10.1729 | 8.8896 | 0 | 0.11395 | 0.55 | 0.55 | 0 | 0 | 10.1729 | [13 empty] | 0 | 0.55
[6 empty] | 0.56 | 0.56 | 10.2196 | 394.3121 | 130.6667 | 14 | 10.2196 | 8.9129 | 0 | 0.11629 | 0.56 | 0.56 | 0 | 0 | 10.2196 | [13 empty] | 0 | 0.56
[6 empty] | 0.57 | 0.57 | 10.2662 | 396.6454 | 133 | 14 | 10.2662 | 8.9362 | 0 | 0.11864 | 0.57 | 0.57 | 0 | 0 | 10.2662 | [13 empty] | 0 | 0.57
[6 empty] | 0.58 | 0.58 | 10.3129 | 398.9787 | 135.3333 | 14 | 10.3129 | 8.9596 | 0 | 0.121 | 0.58 | 0.58 | 0 | 0 | 10.3129 | [13 empty] | 0 | 0.58
[6 empty] | 0.59 | 0.59 | 10.3596 | 401.3121 | 137.6667 | 14 | 10.3596 | 8.9829 | 0 | 0.12337 | 0.59 | 0.59 | 0 | 0 | 10.3596 | [13 empty] | 0 | 0.59
[6 empty] | 0.6 | 0.6 | 10.4062 | 403.6454 | 140 | 14 | 10.4062 | 9.0062 | 0 | 0.12575 | 0.6 | 0.6 | 0 | 0 | 10.4062 | [13 empty] | 0 | 0.6
[6 empty] | 0.61 | 0.61 | 10.4529 | 405.9787 | 142.3333 | 14 | 10.4529 | 9.0296 | 0 | 0.12815 | 0.61 | 0.61 | 0 | 0 | 10.4529 | [13 empty] | 0 | 0.61
[6 empty] | 0.62 | 0.62 | 10.4996 | 408.3121 | 144.6667 | 14 | 10.4996 | 9.0529 | 0 | 0.13055 | 0.62 | 0.62 | 0 | 0 | 10.4996 | [13 empty] | 0 | 0.62
[6 empty] | 0.63 | 0.63 | 10.5462 | 410.6454 | 147 | 14 | 10.5462 | 9.0762 | 0 | 0.13296 | 0.63 | 0.63 | 0 | 0 | 10.5462 | [13 empty] | 0 | 0.63
[6 empty] | 0.64 | 0.64 | 10.5929 | 412.9787 | 149.3333 | 14 | 10.5929 | 9.0996 | 0 | 0.13539 | 0.64 | 0.64 | 0 | 0 | 10.5929 | [13 empty] | 0 | 0.64
[6 empty] | 0.65 | 0.65 | 10.6396 | 415.3121 | 151.6667 | 14 | 10.6396 | 9.1229 | 0 | 0.13782 | 0.65 | 0.65 | 0 | 0 | 10.6396 | [13 empty] | 0 | 0.65
[6 empty] | 0.66 | 0.66 | 10.6862 | 417.6454 | 154 | 14 | 10.6862 | 9.1462 | 0 | 0.14027 | 0.66 | 0.66 | 0 | 0 | 10.6862 | [13 empty] | 0 | 0.66
[6 empty] | 0.67 | 0.67 | 10.7329 | 419.9787 | 156.3333 | 14 | 10.7329 | 9.1696 | 0 | 0.14273 | 0.67 | 0.67 | 0 | 0 | 10.7329 | [13 empty] | 0 | 0.67
[6 empty] | 0.68 | 0.68 | 10.7796 | 422.3121 | 158.6667 | 14 | 10.7796 | 9.1929 | 0 | 0.14519 | 0.68 | 0.68 | 0 | 0 | 10.7796 | [13 empty] | 0 | 0.68
[6 empty] | 0.69 | 0.69 | 10.8262 | 424.6454 | 161 | 14 | 10.8262 | 9.2162 | 0 | 0.14767 | 0.69 | 0.69 | 0 | 0 | 10.8262 | [13 empty] | 0 | 0.69
[6 empty] | 0.7 | 0.7 | 10.8729 | 426.9787 | 163.3333 | 14 | 10.8729 | 9.2396 | 0 | 0.15016 | 0.7 | 0.7 | 0 | 0 | 10.8729 | [13 empty] | 0 | 0.7
[6 empty] | 0.71 | 0.71 | 10.9196 | 429.3121 | 165.6667 | 14 | 10.9196 | 9.2629 | 0 | 0.15266 | 0.71 | 0.71 | 0 | 0 | 10.9196 | [13 empty] | 0 | 0.71
[6 empty] | 0.72 | 0.72 | 10.9662 | 431.6454 | 168 | 14 | 10.9662 | 9.2862 | 0 | 0.15517 | 0.72 | 0.72 | 0 | 0 | 10.9662 | [13 empty] | 0 | 0.72
[6 empty] | 0.73 | 0.73 | 11.0129 | 433.9787 | 170.3333 | 14 | 11.0129 | 9.3096 | 0 | 0.15769 | 0.73 | 0.73 | 0 | 0 | 11.0129 | [13 empty] | 0 | 0.73
[6 empty] | 0.74 | 0.74 | 11.0596 | 436.3121 | 172.6667 | 14 | 11.0596 | 9.3329 | 0 | 0.16022 | 0.74 | 0.74 | 0 | 0 | 11.0596 | [13 empty] | 0 | 0.74
[6 empty] | 0.75 | 0.75 | 11.1062 | 438.6454 | 175 | 14 | 11.1062 | 9.3562 | 0 | 0.16276 | 0.75 | 0.75 | 0 | 0 | 11.1062 | [13 empty] | 0 | 0.75
[6 empty] | 0.76 | 0.76 | 11.1529 | 440.9787 | 177.3333 | 14 | 11.1529 | 9.3796 | 0 | 0.16532 | 0.76 | 0.76 | 0 | 0 | 11.1529 | [13 empty] | 0 | 0.76
[6 empty] | 0.77 | 0.77 | 11.1996 | 443.3121 | 179.6667 | 14 | 11.1996 | 9.4029 | 0 | 0.16788 | 0.77 | 0.77 | 0 | 0 | 11.1996 | [13 empty] | 0 | 0.77
[6 empty] | 0.78 | 0.78 | 11.2462 | 445.6454 | 182 | 14 | 11.2462 | 9.4262 | 0 | 0.17046 | 0.78 | 0.78 | 0 | 0 | 11.2462 | [13 empty] | 0 | 0.78
[6 empty] | 0.79 | 0.79 | 11.2929 | 447.9787 | 184.3333 | 14 | 11.2929 | 9.4496 | 0 | 0.17304 | 0.79 | 0.79 | 0 | 0 | 11.2929 | [13 empty] | 0 | 0.79
[6 empty] | 0.8 | 0.8 | 11.3396 | 450.3121 | 186.6667 | 14 | 11.3396 | 9.4729 | 0 | 0.17564 | 0.8 | 0.8 | 0 | 0 | 11.3396 | [13 empty] | 0 | 0.8
[6 empty] | 0.81 | 0.81 | 11.3862 | 452.6454 | 189 | 14 | 11.3862 | 9.4962 | 0 | 0.17824 | 0.81 | 0.81 | 0 | 0 | 11.3862 | [13 empty] | 0 | 0.81
[6 empty] | 0.82 | 0.82 | 11.4329 | 454.9787 | 191.3333 | 14 | 11.4329 | 9.5196 | 0 | 0.18086 | 0.82 | 0.82 | 0 | 0 | 11.4329 | [13 empty] | 0 | 0.82
[6 empty] | 0.83 | 0.83 | 11.4796 | 457.3121 | 193.6667 | 14 | 11.4796 | 9.5429 | 0 | 0.18349 | 0.83 | 0.83 | 0 | 0 | 11.4796 | [13 empty] | 0 | 0.83
[6 empty] | 0.84 | 0.84 | 11.5262 | 459.6454 | 196 | 14 | 11.5262 | 9.5662 | 0 | 0.18613 | 0.84 | 0.84 | 0 | 0 | 11.5262 | [13 empty] | 0 | 0.84
[6 empty] | 0.85 | 0.85 | 11.5729 | 461.9787 | 198.3333 | 14 | 11.5729 | 9.5896 | 0 | 0.18877 | 0.85 | 0.85 | 0 | 0 | 11.5729 | [13 empty] | 0 | 0.85
[6 empty] | 0.86 | 0.86 | 11.6196 | 464.3121 | 200.6667 | 14 | 11.6196 | 9.6129 | 0 | 0.19143 | 0.86 | 0.86 | 0 | 0 | 11.6196 | [13 empty] | 0 | 0.86
[6 empty] | 0.87 | 0.87 | 11.6662 | 466.6454 | 203 | 14 | 11.6662 | 9.6362 | 0 | 0.1941 | 0.87 | 0.87 | 0 | 0 | 11.6662 | [13 empty] | 0 | 0.87
[6 empty] | 0.88 | 0.88 | 11.7129 | 468.9787 | 205.3333 | 14 | 11.7129 | 9.6596 | 0 | 0.19679 | 0.88 | 0.88 | 0 | 0 | 11.7129 | [13 empty] | 0 | 0.88
[6 empty] | 0.89 | 0.89 | 11.7596 | 471.3121 | 207.6667 | 14 | 11.7596 | 9.6829 | 0 | 0.19948 | 0.89 | 0.89 | 0 | 0 | 11.7596 | [13 empty] | 0 | 0.89
[6 empty] | 0.9 | 0.9 | 11.8062 | 473.6454 | 210 | 14 | 11.8062 | 9.7062 | 0 | 0.20218 | 0.9 | 0.9 | 0 | 0 | 11.8062 | [13 empty] | 0 | 0.9
[6 empty] | 0.91 | 0.91 | 11.8529 | 475.9787 | 212.3333 | 14 | 11.8529 | 9.7296 | 0 | 0.20489 | 0.91 | 0.91 | 0 | 0 | 11.8529 | [13 empty] | 0 | 0.91
[6 empty] | 0.92 | 0.92 | 11.8996 | 478.3121 | 214.6667 | 14 | 11.8996 | 9.7529 | 0 | 0.20762 | 0.92 | 0.92 | 0 | 0 | 11.8996 | [13 empty] | 0 | 0.92
[6 empty] | 0.93 | 0.93 | 11.9462 | 480.6454 | 217 | 14 | 11.9462 | 9.7762 | 0 | 0.21035 | 0.93 | 0.93 | 0 | 0 | 11.9462 | [13 empty] | 0 | 0.93
[6 empty] | 0.94 | 0.94 | 11.9929 | 482.9787 | 219.3333 | 14 | 11.9929 | 9.7996 | 0 | 0.2131 | 0.94 | 0.94 | 0 | 0 | 11.9929 | [13 empty] | 0 | 0.94
[6 empty] | 0.95 | 0.95 | 12.0396 | 485.3121 | 221.6667 | 14 | 12.0396 | 9.8229 | 0 | 0.21585 | 0.95 | 0.95 | 0 | 0 | 12.0396 | [13 empty] | 0 | 0.95
[6 empty] | 0.96 | 0.96 | 12.0862 | 487.6454 | 224 | 14 | 12.0862 | 9.8462 | 0 | 0.21862 | 0.96 | 0.96 | 0 | 0 | 12.0862 | [13 empty] | 0 | 0.96
[6 empty] | 0.97 | 0.97 | 12.1329 | 489.9787 | 226.3333 | 14 | 12.1329 | 9.8696 | 0 | 0.2214 | 0.97 | 0.97 | 0 | 0 | 12.1329 | [13 empty] | 0 | 0.97
[6 empty] | 0.98 | 0.98 | 12.1796 | 492.3121 | 228.6667 | 14 | 12.1796 | 9.8929 | 0 | 0.22418 | 0.98 | 0.98 | 0 | 0 | 12.1796 | [13 empty] | 0 | 0.98
[6 empty] | 0.99 | 0.99 | 12.2262 | 494.6454 | 231 | 14 | 12.2262 | 9.9162 | 0 | 0.22698 | 0.99 | 0.99 | 0 | 0 | 12.2262 | [13 empty] | 0 | 0.99
[6 empty] | 1 | 1 | 12.2729 | 496.9787 | 233.3333 | 14 | 12.2729 | 9.9396 | 0 | 0.22979 | 1 | 1 | 0 | 0 | 12.2729 | [13 empty] | 0 | 1
[6 empty] | 1.01 | 1.01 | 12.3196 | 499.3121 | 235.6667 | 14 | 12.3196 | 9.9629 | 0 | 0.23261 | 1.01 | 1.01 | 0 | 0 | 12.3196 | [13 empty] | 0 | 1.01
[6 empty] | 1.02 | 1.02 | 12.3662 | 501.6454 | 238 | 14 | 12.3662 | 9.9862 | 0 | 0.23544 | 1.02 | 1.02 | 0 | 0 | 12.3662 | [13 empty] | 0 | 1.02
[6 empty] | 1.03 | 1.03 | 12.4129 | 503.9787 | 240.3333 | 14 | 12.4129 | 10.0096 | 0 | 0.23828 | 1.03 | 1.03 | 0 | 0 | 12.4129 | [13 empty] | 0 | 1.03
[6 empty] | 1.04 | 1.04 | 12.4596 | 506.3121 | 242.6667 | 14 | 12.4596 | 10.0329 | 0 | 0.24114 | 1.04 | 1.04 | 0 | 0 | 12.4596 | [13 empty] | 0 | 1.04
[6 empty] | 1.05 | 1.05 | 12.5062 | 508.6454 | 245 | 14 | 12.5062 | 10.0562 | 0 | 0.244 | 1.05 | 1.05 | 0 | 0 | 12.5062 | [13 empty] | 0 | 1.05
[6 empty] | 1.06 | 1.06 | 12.5529 | 510.9787 | 247.3333 | 14 | 12.5529 | 10.0796 | 0 | 0.24687 | 1.06 | 1.06 | 0 | 0 | 12.5529 | [13 empty] | 0 | 1.06
[6 empty] | 1.07 | 1.07 | 12.5996 | 513.3121 | 249.6667 | 14 | 12.5996 | 10.1029 | 0 | 0.24976 | 1.07 | 1.07 | 0 | 0 | 12.5996 | [13 empty] | 0 | 1.07
[6 empty] | 1.08 | 1.08 | 12.6462 | 515.6454 | 252 | 14 | 12.6462 | 10.1262 | 0 | 0.25265 | 1.08 | 1.08 | 0 | 0 | 12.6462 | [13 empty] | 0 | 1.08
[6 empty] | 1.09 | 1.09 | 12.6929 | 517.9787 | 254.3333 | 14 | 12.6929 | 10.1496 | 0 | 0.25556 | 1.09 | 1.09 | 0 | 0 | 12.6929 | [13 empty] | 0 | 1.09
[6 empty] | 1.1 | 1.1 | 12.7396 | 520.3121 | 256.6667 | 14 | 12.7396 | 10.1729 | 0 | 0.25847 | 1.1 | 1.1 | 0 | 0 | 12.7396 | [13 empty] | 0 | 1.1
[6 empty] | 1.11 | 1.11 | 12.7862 | 522.6454 | 259 | 14 | 12.7862 | 10.1962 | 0 | 0.2614 | 1.11 | 1.11 | 0 | 0 | 12.7862 | [13 empty] | 0 | 1.11
[6 empty] | 1.12 | 1.12 | 12.8329 | 524.9787 | 261.3333 | 14 | 12.8329 | 10.2196 | 0 | 0.26434 | 1.12 | 1.12 | 0 | 0 | 12.8329 | [13 empty] | 0 | 1.12
[6 empty] | 1.13 | 1.13 | 12.8796 | 527.3121 | 263.6667 | 14 | 12.8796 | 10.2429 | 0 | 0.26728 | 1.13 | 1.13 | 0 | 0 | 12.8796 | [13 empty] | 0 | 1.13
[6 empty] | 1.14 | 1.14 | 12.9262 | 529.6454 | 266 | 14 | 12.9262 | 10.2662 | 0 | 0.27024 | 1.14 | 1.14 | 0 | 0 | 12.9262 | [13 empty] | 0 | 1.14
[6 empty] | 1.15 | 1.15 | 12.9729 | 531.9787 | 268.3333 | 14 | 12.9729 | 10.2896 | 0 | 0.27321 | 1.15 | 1.15 | 0 | 0 | 12.9729 | [13 empty] | 0 | 1.15
[6 empty] | 1.16 | 1.16 | 13.0196 | 534.3121 | 270.6667 | 14 | 13.0196 | 10.3129 | 0 | 0.27619 | 1.16 | 1.16 | 0 | 0 | 13.0196 | [13 empty] | 0 | 1.16
[6 empty] | 1.17 | 1.17 | 13.0662 | 536.6454 | 273 | 14 | 13.0662 | 10.3362 | 0 | 0.27918 | 1.17 | 1.17 | 0 | 0 | 13.0662 | [13 empty] | 0 | 1.17
[6 empty] | 1.18 | 1.18 | 13.1129 | 538.9787 | 275.3333 | 14 | 13.1129 | 10.3596 | 0 | 0.28219 | 1.18 | 1.18 | 0 | 0 | 13.1129 | [13 empty] | 0 | 1.18
[6 empty] | 1.19 | 1.19 | 13.1596 | 541.3121 | 277.6667 | 14 | 13.1596 | 10.3829 | 0 | 0.2852 | 1.19 | 1.19 | 0 | 0 | 13.1596 | [13 empty] | 0 | 1.19
[6 empty] | 1.2 | 1.2 | 13.2062 | 543.6454 | 280 | 14 | 13.2062 | 10.4062 | 0 | 0.28822 | 1.2 | 1.2 | 0 | 0 | 13.2062 | [13 empty] | 0 | 1.2
[6 empty] | 1.21 | 1.21 | 13.2529 | 545.9787 | 282.3333 | 14 | 13.2529 | 10.4296 | 0 | 0.29125 | 1.21 | 1.21 | 0 | 0 | 13.2529 | [13 empty] | 0 | 1.21
[6 empty] | 1.22 | 1.22 | 13.2996 | 548.3121 | 284.6667 | 14 | 13.2996 | 10.4529 | 0 | 0.2943 | 1.22 | 1.22 | 0 | 0 | 13.2996 | [13 empty] | 0 | 1.22
[6 empty] | 1.23 | 1.23 | 13.3462 | 550.6454 | 287 | 14 | 13.3462 | 10.4762 | 0 | 0.29735 | 1.23 | 1.23 | 0 | 0 | 13.3462 | [13 empty] | 0 | 1.23
[6 empty] | 1.24 | 1.24 | 13.3929 | 552.9787 | 289.3333 | 14 | 13.3929 | 10.4996 | 0 | 0.30042 | 1.24 | 1.24 | 0 | 0 | 13.3929 | [13 empty] | 0 | 1.24
[6 empty] | 1.25 | 1.25 | 13.4396 | 555.3121 | 291.6667 | 14 | 13.4396 | 10.5229 | 0 | 0.3035 | 1.25 | 1.25 | 0 | 0 | 13.4396 | [13 empty] | 0 | 1.25
[6 empty] | 1.26 | 1.26 | 13.4862 | 557.6454 | 294 | 14 | 13.4862 | 10.5462 | 0 | 0.30658 | 1.26 | 1.26 | 0 | 0 | 13.4862 | [13 empty] | 0 | 1.26
[6 empty] | 1.27 | 1.27 | 13.5329 | 559.9787 | 296.3333 | 14 | 13.5329 | 10.5696 | 0 | 0.30968 | 1.27 | 1.27 | 0 | 0 | 13.5329 | [13 empty] | 0 | 1.27
[6 empty] | 1.28 | 1.28 | 13.5796 | 562.3121 | 298.6667 | 14 | 13.5796 | 10.5929 | 0 | 0.31279 | 1.28 | 1.28 | 0 | 0 | 13.5796 | [13 empty] | 0 | 1.28
[6 empty] | 1.29 | 1.29 | 13.6262 | 564.6454 | 301 | 14 | 13.6262 | 10.6162 | 0 | 0.31591 | 1.29 | 1.29 | 0 | 0 | 13.6262 | [13 empty] | 0 | 1.29
[6 empty] | 1.3 | 1.3 | 13.6729 | 566.9787 | 303.3333 | 14 | 13.6729 | 10.6396 | 0 | 0.31904 | 1.3 | 1.3 | 0 | 0 | 13.6729 | [13 empty] | 0 | 1.3
[6 empty] | 1.31 | 1.31 | 13.7196 | 569.3121 | 305.6667 | 14 | 13.7196 | 10.6629 | 0 | 0.32218 | 1.31 | 1.31 | 0 | 0 | 13.7196 | [13 empty] | 0 | 1.31
[6 empty] | 1.32 | 1.32 | 13.7662 | 571.6454 | 308 | 14 | 13.7662 | 10.6862 | 0 | 0.32533 | 1.32 | 1.32 | 0 | 0 | 13.7662 | [13 empty] | 0 | 1.32
[6 empty] | 1.33 | 1.33 | 13.8129 | 573.9787 | 310.3333 | 14 | 13.8129 | 10.7096 | 0 | 0.32849 | 1.33 | 1.33 | 0 | 0 | 13.8129 | [13 empty] | 0 | 1.33
[6 empty] | 1.34 | 1.34 | 13.8596 | 576.3121 | 312.6667 | 14 | 13.8596 | 10.7329 | 0 | 0.33166 | 1.34 | 1.34 | 0 | 0 | 13.8596 | [13 empty] | 0 | 1.34
[6 empty] | 1.35 | 1.35 | 13.9062 | 578.6454 | 315 | 14 | 13.9062 | 10.7562 | 0 | 0.33485 | 1.35 | 1.35 | 0 | 0 | 13.9062 | [13 empty] | 0 | 1.35
[6 empty] | 1.36 | 1.36 | 13.9529 | 580.9787 | 317.3333 | 14 | 13.9529 | 10.7796 | 0 | 0.33804 | 1.36 | 1.36 | 0 | 0 | 13.9529 | [13 empty] | 0 | 1.36
[6 empty] | 1.37 | 1.37 | 13.9996 | 583.3121 | 319.6667 | 14 | 13.9996 | 10.8029 | 0 | 0.34124 | 1.37 | 1.37 | 0 | 0 | 13.9996 | [13 empty] | 0 | 1.37
[6 empty] | 1.38 | 1.38 | 14.0462 | 585.6454 | 322 | 14 | 14.0462 | 10.8262 | 0 | 0.34446 | 1.38 | 1.38 | 0 | 0 | 14.0462 | [13 empty] | 0 | 1.38
[6 empty] | 1.39 | 1.39 | 14.0929 | 587.9787 | 324.3333 | 14 | 14.0929 | 10.8496 | 0 | 0.34769 | 1.39 | 1.39 | 0 | 0 | 14.0929 | [13 empty] | 0 | 1.39
[6 empty] | 1.4 | 1.4 | 14.1396 | 590.3121 | 326.6667 | 14 | 14.1396 | 10.8729 | 0 | 0.35092 | 1.4 | 1.4 | 0 | 0 | 14.1396 | [13 empty] | 0 | 1.4
[6 empty] | 1.41 | 1.41 | 14.1862 | 592.6454 | 329 | 14 | 14.1862 | 10.8962 | 0 | 0.35417 | 1.41 | 1.41 | 0 | 0 | 14.1862 | [13 empty] | 0 | 1.41
[6 empty] | 1.42 | 1.42 | 14.2329 | 594.9787 | 331.3333 | 14 | 14.2329 | 10.9196 | 0 | 0.35743 | 1.42 | 1.42 | 0 | 0 | 14.2329 | [13 empty] | 0 | 1.42
[6 empty] | 1.43 | 1.43 | 14.2796 | 597.3121 | 333.6667 | 14 | 14.2796 | 10.9429 | 0 | 0.3607 | 1.43 | 1.43 | 0 | 0 | 14.2796 | [13 empty] | 0 | 1.43
[6 empty] | 1.44 | 1.44 | 14.3262 | 599.6454 | 336 | 14 | 14.3262 | 10.9662 | 0 | 0.36398 | 1.44 | 1.44 | 0 | 0 | 14.3262 | [13 empty] | 0 | 1.44
[6 empty] | 1.45 | 1.45 | 14.3729 | 601.9787 | 338.3333 | 14 | 14.3729 | 10.9896 | 0 | 0.36727 | 1.45 | 1.45 | 0 | 0 | 14.3729 | [13 empty] | 0 | 1.45
[6 empty] | 1.46 | 1.46 | 14.4196 | 604.3121 | 340.6667 | 14 | 14.4196 | 11.0129 | 0 | 0.37057 | 1.46 | 1.46 | 0 | 0 | 14.4196 | [13 empty] | 0 | 1.46
[6 empty] | 1.47 | 1.47 | 14.4662 | 606.6454 | 343 | 14 | 14.4662 | 11.0362 | 0 | 0.37388 | 1.47 | 1.47 | 0 | 0 | 14.4662 | [13 empty] | 0 | 1.47
[6 empty] | 1.48 | 1.48 | 14.5129 | 608.9787 | 345.3333 | 14 | 14.5129 | 11.0596 | 0 | 0.3772 | 1.48 | 1.48 | 0 | 0 | 14.5129 | [13 empty] | 0 | 1.48
[6 empty] | 1.49 | 1.49 | 14.5596 | 611.3121 | 347.6667 | 14 | 14.5596 | 11.0829 | 0 | 0.38053 | 1.49 | 1.49 | 0 | 0 | 14.5596 | [13 empty] | 0 | 1.49
[6 empty] | 1.5 | 1.5 | 14.6062 | 613.6454 | 350 | 14 | 14.6062 | 11.1062 | 0 | 0.38388 | 1.5 | 1.5 | 0 | 0 | 14.6062 | [13 empty] | 0 | 1.5
[6 empty] | 1.51 | 1.51 | 14.6529 | 615.9787 | 352.3333 | 14 | 14.6529 | 11.1296 | 0 | 0.38723 | 1.51 | 1.51 | 0 | 0 | 14.6529 | [13 empty] | 0 | 1.51
[6 empty] | 1.52 | 1.52 | 14.6996 | 618.3121 | 354.6667 | 14 | 14.6996 | 11.1529 | 0 | 0.39059 | 1.52 | 1.52 | 0 | 0 | 14.6996 | [13 empty] | 0 | 1.52
[6 empty] | 1.53 | 1.53 | 14.7462 | 620.6454 | 357 | 14 | 14.7462 | 11.1762 | 0 | 0.39397 | 1.53 | 1.53 | 0 | 0 | 14.7462 | [13 empty] | 0 | 1.53
[6 empty] | 1.54 | 1.54 | 14.7929 | 622.9787 | 359.3333 | 14 | 14.7929 | 11.1996 | 0 | 0.39736 | 1.54 | 1.54 | 0 | 0 | 14.7929 | [13 empty] | 0 | 1.54
[6 empty] | 1.55 | 1.55 | 14.8396 | 625.3121 | 361.6667 | 14 | 14.8396 | 11.2229 | 0 | 0.40075 | 1.55 | 1.55 | 0 | 0 | 14.8396 | [13 empty] | 0 | 1.55
[6 empty] | 1.56 | 1.56 | 14.8862 | 627.6454 | 364 | 14 | 14.8862 | 11.2462 | 0 | 0.40416 | 1.56 | 1.56 | 0 | 0 | 14.8862 | [13 empty] | 0 | 1.56
[6 empty] | 1.57 | 1.57 | 14.9329 | 629.9787 | 366.3333 | 14 | 14.9329 | 11.2696 | 0 | 0.40758 | 1.57 | 1.57 | 0 | 0 | 14.9329 | [13 empty] | 0 | 1.57
[6 empty] | 1.58 | 1.58 | 14.9796 | 632.3121 | 368.6667 | 14 | 14.9796 | 11.2929 | 0 | 0.41101 | 1.58 | 1.58 | 0 | 0 | 14.9796 | [13 empty] | 0 | 1.58
[6 empty] | 1.59 | 1.59 | 15.0262 | 634.6454 | 371 | 14 | 15.0262 | 11.3162 | 0 | 0.41445 | 1.59 | 1.59 | 0 | 0 | 15.0262 | [13 empty] | 0 | 1.59
[6 empty] | 1.6 | 1.6 | 15.0729 | 636.9787 | 373.3333 | 14 | 15.0729 | 11.3396 | 0 | 0.4179 | 1.6 | 1.6 | 0 | 0 | 15.0729 | [13 empty] | 0 | 1.6
[6 empty] | 1.61 | 1.61 | 15.1196 | 639.3121 | 375.6667 | 14 | 15.1196 | 11.3629 | 0 | 0.42136 | 1.61 | 1.61 | 0 | 0 | 15.1196 | [13 empty] | 0 | 1.61
[6 empty] | 1.62 | 1.62 | 15.1662 | 641.6454 | 378 | 14 | 15.1662 | 11.3862 | 0 | 0.42483 | 1.62 | 1.62 | 0 | 0 | 15.1662 | [13 empty] | 0 | 1.62
[6 empty] | 1.63 | 1.63 | 15.2129 | 643.9787 | 380.3333 | 14 | 15.2129 | 11.4096 | 0 | 0.42831 | 1.63 | 1.63 | 0 | 0 | 15.2129 | [13 empty] | 0 | 1.63
[6 empty] | 1.64 | 1.64 | 15.2596 | 646.3121 | 382.6667 | 14 | 15.2596 | 11.4329 | 0 | 0.4318 | 1.64 | 1.64 | 0 | 0 | 15.2596 | [13 empty] | 0 | 1.64
[6 empty] | 1.65 | 1.65 | 15.3062 | 648.6454 | 385 | 14 | 15.3062 | 11.4562 | 0 | 0.43531 | 1.65 | 1.65 | 0 | 0 | 15.3062 | [13 empty] | 0 | 1.65
[6 empty] | 1.66 | 1.66 | 15.3529 | 650.9787 | 387.3333 | 14 | 15.3529 | 11.4796 | 0 | 0.43882 | 1.66 | 1.66 | 0 | 0 | 15.3529 | [13 empty] | 0 | 1.66
[6 empty] | 1.67 | 1.67 | 15.3996 | 653.3121 | 389.6667 | 14 | 15.3996 | 11.5029 | 0 | 0.44235 | 1.67 | 1.67 | 0 | 0 | 15.3996 | [13 empty] | 0 | 1.67
[6 empty] | 1.68 | 1.68 | 15.4462 | 655.6454 | 392 | 14 | 15.4462 | 11.5262 | 0 | 0.44588 | 1.68 | 1.68 | 0 | 0 | 15.4462 | [13 empty] | 0 | 1.68
[6 empty] | 1.69 | 1.69 | 15.4929 | 657.9787 | 394.3333 | 14 | 15.4929 | 11.5496 | 0 | 0.44943 | 1.69 | 1.69 | 0 | 0 | 15.4929 | [13 empty] | 0 | 1.69
[6 empty] | 1.7 | 1.7 | 15.5396 | 660.3121 | 396.6667 | 14 | 15.5396 | 11.5729 | 0 | 0.45299 | 1.7 | 1.7 | 0 | 0 | 15.5396 | [13 empty] | 0 | 1.7
[6 empty] | 1.71 | 1.71 | 15.5862 | 662.6454 | 399 | 14 | 15.5862 | 11.5962 | 0 | 0.45655 | 1.71 | 1.71 | 0 | 0 | 15.5862 | [13 empty] | 0 | 1.71
[6 empty] | 1.72 | 1.72 | 15.6329 | 664.9787 | 401.3333 | 14 | 15.6329 | 11.6196 | 0 | 0.46013 | 1.72 | 1.72 | 0 | 0 | 15.6329 | [13 empty] | 0 | 1.72
[6 empty] | 1.73 | 1.73 | 15.6796 | 667.3121 | 403.6667 | 14 | 15.6796 | 11.6429 | 0 | 0.46372 | 1.73 | 1.73 | 0 | 0 | 15.6796 | [13 empty] | 0 | 1.73
[6 empty] | 1.74 | 1.74 | 15.7262 | 669.6454 | 406 | 14 | 15.7262 | 11.6662 | 0 | 0.46732 | 1.74 | 1.74 | 0 | 0 | 15.7262 | [13 empty] | 0 | 1.74
[6 empty] | 1.75 | 1.75 | 15.7729 | 671.9787 | 408.3333 | 14 | 15.7729 | 11.6896 | 0 | 0.47093 | 1.75 | 1.75 | 0 | 0 | 15.7729 | [13 empty] | 0 | 1.75
[6 empty] | 1.76 | 1.76 | 15.8196 | 674.3121 | 410.6667 | 14 | 15.8196 | 11.7129 | 0 | 0.47455 | 1.76 | 1.76 | 0 | 0 | 15.8196 | [13 empty] | 0 | 1.76
[6 empty] | 1.77 | 1.77 | 15.8662 | 676.6454 | 413 | 14 | 15.8662 | 11.7362 | 0 | 0.47818 | 1.77 | 1.77 | 0 | 0 | 15.8662 | [13 empty] | 0 | 1.77
[6 empty] | 1.78 | 1.78 | 15.9129 | 678.9787 | 415.3333 | 14 | 15.9129 | 11.7596 | 0 | 0.48183 | 1.78 | 1.78 | 0 | 0 | 15.9129 | [13 empty] | 0 | 1.78
[6 empty] | 1.79 | 1.79 | 15.9596 | 681.3121 | 417.6667 | 14 | 15.9596 | 11.7829 | 0 | 0.48548 | 1.79 | 1.79 | 0 | 0 | 15.9596 | [13 empty] | 0 | 1.79
[6 empty] | 1.8 | 1.8 | 16.0062 | 683.6454 | 420 | 14 | 16.0062 | 11.8062 | 0 | 0.48914 | 1.8 | 1.8 | 0 | 0 | 16.0062 | [13 empty] | 0 | 1.8
[6 empty] | 1.81 | 1.81 | 5 | 685.9787 | 416.8069 | -52.3174 | 5 | 11.7688 | 0 | 0.48487 | 1.81 | 1.81 | 0 | 0 | 5 | [13 empty] | 0 | 1.81
[6 empty] | 1.82 | 1.82 | 5 | 677.2592 | 408.1745 | -51.2711 | 5 | 11.7319 | 0 | 0.48067 | 1.82 | 1.82 | 0 | 0 | 5 | [13 empty] | 0 | 1.82
[6 empty] | 1.83 | 1.83 | 5 | 668.714 | 399.7148 | -50.2457 | 5 | 11.6953 | 0 | 0.47657 | 1.83 | 1.83 | 0 | 0 | 5 | [13 empty] | 0 | 1.83
[6 empty] | 1.84 | 1.84 | 5 | 660.3397 | 391.4242 | -49.2408 | 5 | 11.6591 | 0 | 0.47254 | 1.84 | 1.84 | 0 | 0 | 5 | [13 empty] | 0 | 1.84
[6 empty] | 1.85 | 1.85 | 5 | 652.1329 | 383.2995 | -48.256 | 5 | 11.6233 | 0 | 0.4686 | 1.85 | 1.85 | 0 | 0 | 5 | [13 empty] | 0 | 1.85
[6 empty] | 1.86 | 1.86 | 5 | 644.0903 | 375.3373 | -47.2908 | 5 | 11.5879 | 0 | 0.46473 | 1.86 | 1.86 | 0 | 0 | 5 | [13 empty] | 0 | 1.86
[6 empty] | 1.87 | 1.87 | 5 | 636.2085 | 367.5343 | -46.345 | 5 | 11.5528 | 0 | 0.46094 | 1.87 | 1.87 | 0 | 0 | 5 | [13 empty] | 0 | 1.87
[6 empty] | 1.88 | 1.88 | 5 | 628.4843 | 359.8874 | -45.4181 | 5 | 11.5181 | 0 | 0.45723 | 1.88 | 1.88 | 0 | 0 | 5 | [13 empty] | 0 | 1.88
[6 empty] | 1.89 | 1.89 | 5 | 620.9146 | 352.3934 | -44.5098 | 5 | 11.4838 | 0 | 0.45359 | 1.89 | 1.89 | 0 | 0 | 5 | [13 empty] | 0 | 1.89
[6 empty] | 1.9 | 1.9 | 5 | 613.4963 | 345.0493 | -43.6196 | 5 | 11.4499 | 0 | 0.45002 | 1.9 | 1.9 | 0 | 0 | 5 | [13 empty] | 0 | 1.9
[6 empty] | 1.91 | 1.91 | 5 | 606.2264 | 337.852 | -42.7472 | 5 | 11.4163 | 0 | 0.44653 | 1.91 | 1.91 | 0 | 0 | 5 | [13 empty] | 0 | 1.91
[6 empty] | 1.92 | 1.92 | 5 | 599.1019 | 330.7988 | -41.8922 | 5 | 11.3831 | 0 | 0.44311 | 1.92 | 1.92 | 0 | 0 | 5 | [13 empty] | 0 | 1.92
[6 empty] | 1.93 | 1.93 | 5 | 592.1198 | 323.8865 | -41.0544 | 5 | 11.3502 | 0 | 0.43975 | 1.93 | 1.93 | 0 | 0 | 5 | [13 empty] | 0 | 1.93
[6 empty] | 1.94 | 1.94 | 5 | 585.2774 | 317.1126 | -40.2333 | 5 | 11.3176 | 0 | 0.43646 | 1.94 | 1.94 | 0 | 0 | 5 | [13 empty] | 0 | 1.94
[6 empty] | 1.95 | 1.95 | 5 | 578.5719 | 310.4741 | -39.4286 | 5 | 11.2854 | 0 | 0.43324 | 1.95 | 1.95 | 0 | 0 | 5 | [13 empty] | 0 | 1.95
[6 empty] | 1.96 | 1.96 | 5 | 572.0004 | 303.9683 | -38.6401 | 5 | 11.2535 | 0 | 0.43008 | 1.96 | 1.96 | 0 | 0 | 5 | [13 empty] | 0 | 1.96
[6 empty] | 1.97 | 1.97 | 5 | 565.5604 | 297.5927 | -37.8673 | 5 | 11.2219 | 0 | 0.42698 | 1.97 | 1.97 | 0 | 0 | 5 | [13 empty] | 0 | 1.97
[6 empty] | 1.98 | 1.98 | 5 | 559.2492 | 291.3446 | -37.1099 | 5 | 11.1906 | 0 | 0.42395 | 1.98 | 1.98 | 0 | 0 | 5 | [13 empty] | 0 | 1.98
[6 empty] | 1.99 | 1.99 | 5 | 553.0642 | 285.2215 | -36.3677 | 5 | 11.1596 | 0 | 0.42098 | 1.99 | 1.99 | 0 | 0 | 5 | [13 empty] | 0 | 1.99
[6 empty] | 2 | 2 | 5 | 547.0029 | 279.2208 | -35.6404 | 5 | 11.129 | 0 | 0.41806 | 2 | 2 | 0 | 0 | 5 | [13 empty] | 0 | 2
[6 empty] | 2.01 | 2.01 | 5 | 541.0629 | 273.3402 | -34.9275 | 5 | 11.0987 | 0 | 0.41521 | 2.01 | 2.01 | 0 | 0 | 5 | [13 empty] | 0 | 2.01
[6 empty] | 2.02 | 2.02 | 5 | 535.2416 | 267.5771 | -34.229 | 5 | 11.0686 | 0 | 0.41241 | 2.02 | 2.02 | 0 | 0 | 5 | [13 empty] | 0 | 2.02
[6 empty] | 2.03 | 2.03 | 5 | 529.5368 | 261.9293 | -33.5444 | 5 | 11.0389 | 0 | 0.40967 | 2.03 | 2.03 | 0 | 0 | 5 | [13 empty] | 0 | 2.03
[6 empty] | 2.04 | 2.04 | 5 | 523.9461 | 256.3945 | -32.8735 | 5 | 11.0094 | 0 | 0.40698 | 2.04 | 2.04 | 0 | 0 | 5 | [13 empty] | 0 | 2.04
[6 empty] | 2.05 | 2.05 | 5 | 518.4671 | 250.9704 | -32.2161 | 5 | 10.9802 | 0 | 0.40435 | 2.05 | 2.05 | 0 | 0 | 5 | [13 empty] | 0 | 2.05
[6 empty] | 2.06 | 2.06 | 5 | 513.0978 | 245.6547 | -31.5717 | 5 | 10.9514 | 0 | 0.40176 | 2.06 | 2.06 | 0 | 0 | 5 | [13 empty] | 0 | 2.06
[6 empty] | 2.07 | 2.07 | 5 | 507.8358 | 240.4454 | -30.9403 | 5 | 10.9227 | 0 | 0.39923 | 2.07 | 2.07 | 0 | 0 | 5 | [13 empty] | 0 | 2.07
[6 empty] | 2.08 | 2.08 | 5 | 502.6791 | 235.3403 | -30.3215 | 5 | 10.8944 | 0 | 0.39676 | 2.08 | 2.08 | 0 | 0 | 5 | [13 empty] | 0 | 2.08
[6 empty] | 2.09 | 2.09 | 5 | 497.6255 | 230.3372 | -29.7151 | 5 | 10.8663 | 0 | 0.39433 | 2.09 | 2.09 | 0 | 0 | 5 | [13 empty] | 0 | 2.09
[6 empty] | 2.1 | 2.1 | 5 | 492.673 | 225.4342 | -29.1208 | 5 | 10.8385 | 0 | 0.39195 | 2.1 | 2.1 | 0 | 0 | 5 | [13 empty] | 0 | 2.1
[6 empty] | 2.11 | 2.11 | 5 | 487.8196 | 220.6293 | -28.5383 | 5 | 10.811 | 0 | 0.38961 | 2.11 | 2.11 | 0 | 0 | 5 | [13 empty] | 0 | 2.11
[6 empty] | 2.12 | 2.12 | 5 | 483.0632 | 215.9205 | -27.9676 | 5 | 10.7837 | 0 | 0.38733 | 2.12 | 2.12 | 0 | 0 | 5 | [13 empty] | 0 | 2.12
[6 empty] | 2.13 | 2.13 | 5 | 478.4019 | 211.3058 | -27.4082 | 5 | 10.7567 | 0 | 0.38509 | 2.13 | 2.13 | 0 | 0 | 5 | [13 empty] | 0 | 2.13
[6 empty] | 2.14 | 2.14 | 5 | 473.8339 | 206.7835 | -26.8601 | 5 | 10.7299 | 0 | 0.38289 | 2.14 | 2.14 | 0 | 0 | 5 | [13 empty] | 0 | 2.14
[6 empty] | 2.15 | 2.15 | 5 | 469.3572 | 202.3515 | -26.3229 | 5 | 10.7034 | 0 | 0.38074 | 2.15 | 2.15 | 0 | 0 | 5 | [13 empty] | 0 | 2.15
[6 empty] | 2.16 | 2.16 | 5 | 464.9701 | 198.0083 | -25.7964 | 5 | 10.6771 | 0 | 0.37863 | 2.16 | 2.16 | 0 | 0 | 5 | [13 empty] | 0 | 2.16
[6 empty] | 2.17 | 2.17 | 5 | 460.6707 | 193.7519 | -25.2805 | 5 | 10.6511 | 0 | 0.37656 | 2.17 | 2.17 | 0 | 0 | 5 | [13 empty] | 0 | 2.17
[6 empty] | 2.18 | 2.18 | 5 | 456.4572 | 189.5806 | -24.7749 | 5 | 10.6252 | 0 | 0.37454 | 2.18 | 2.18 | 0 | 0 | 5 | [13 empty] | 0 | 2.18
[6 empty] | 2.19 | 2.19 | 5 | 452.3281 | 185.4927 | -24.2794 | 5 | 10.5997 | 0 | 0.37255 | 2.19 | 2.19 | 0 | 0 | 5 | [13 empty] | 0 | 2.19
[6 empty] | 2.2 | 2.2 | 5 | 448.2815 | 181.4866 | -23.7938 | 5 | 10.5743 | 0 | 0.37061 | 2.2 | 2.2 | 0 | 0 | 5 | [13 empty] | 0 | 2.2
[6 empty] | 2.21 | 2.21 | 5 | 444.3159 | 177.5607 | -23.3179 | 5 | 10.5492 | 0 | 0.3687 | 2.21 | 2.21 | 0 | 0 | 5 | [13 empty] | 0 | 2.21
[6 empty] | 2.22 | 2.22 | 5 | 440.4296 | 173.7132 | -22.8516 | 5 | 10.5243 | 0 | 0.36683 | 2.22 | 2.22 | 0 | 0 | 5 | [13 empty] | 0 | 2.22
[6 empty] | 2.23 | 2.23 | 5 | 436.621 | 169.9427 | -22.3945 | 5 | 10.4997 | 0 | 0.365 | 2.23 | 2.23 | 0 | 0 | 5 | [13 empty] | 0 | 2.23
[6 empty] | 2.24 | 2.24 | 5 | 432.8886 | 166.2476 | -21.9466 | 5 | 10.4752 | 0 | 0.36321 | 2.24 | 2.24 | 0 | 0 | 5 | [13 empty] | 0 | 2.24
[6 empty] | 2.25 | 2.25 | 5 | 429.2308 | 162.6264 | -21.5077 | 5 | 10.451 | 0 | 0.36145 | 2.25 | 2.25 | 0 | 0 | 5 | [13 empty] | 0 | 2.25
[6 empty] | 2.26 | 2.26 | 5 | 425.6462 | 159.0776 | -21.0775 | 5 | 10.427 | 0 | 0.35973 | 2.26 | 2.26 | 0 | 0 | 5 | [13 empty] | 0 | 2.26
[6 empty] | 2.27 | 2.27 | 5 | 422.1333 | 155.5998 | -20.656 | 5 | 10.4032 | 0 | 0.35804 | 2.27 | 2.27 | 0 | 0 | 5 | [13 empty] | 0 | 2.27
[6 empty] | 2.28 | 2.28 | 5 | 418.6906 | 152.1916 | -20.2429 | 5 | 10.3796 | 0 | 0.35638 | 2.28 | 2.28 | 0 | 0 | 5 | [13 empty] | 0 | 2.28
[6 empty] | 2.29 | 2.29 | 5 | 415.3168 | 148.8515 | -19.838 | 5 | 10.3562 | 0 | 0.35476 | 2.29 | 2.29 | 0 | 0 | 5 | [13 empty] | 0 | 2.29
[6 empty] | 2.3 | 2.3 | 5 | 412.0105 | 145.5783 | -19.4413 | 5 | 10.333 | 0 | 0.35317 | 2.3 | 2.3 | 0 | 0 | 5 | [13 empty] | 0 | 2.3
[6 empty] | 2.31 | 2.31 | 5 | 408.7702 | 142.3705 | -19.0524 | 5 | 10.31 | 0 | 0.35161 | 2.31 | 2.31 | 0 | 0 | 5 | [13 empty] | 0 | 2.31
[6 empty] | 2.32 | 2.32 | 5 | 405.5948 | 139.2268 | -18.6714 | 5 | 10.2873 | 0 | 0.35009 | 2.32 | 2.32 | 0 | 0 | 5 | [13 empty] | 0 | 2.32
[6 empty] | 2.33 | 2.33 | 5 | 402.4829 | 136.146 | -18.298 | 5 | 10.2647 | 0 | 0.34859 | 2.33 | 2.33 | 0 | 0 | 5 | [13 empty] | 0 | 2.33
[6 empty] | 2.34 | 2.34 | 5 | 399.4333 | 133.1269 | -17.932 | 5 | 10.2423 | 0 | 0.34713 | 2.34 | 2.34 | 0 | 0 | 5 | [13 empty] | 0 | 2.34
[6 empty] | 2.35 | 2.35 | 5 | 396.4446 | 130.1681 | -17.5734 | 5 | 10.22 | 0 | 0.34569 | 2.35 | 2.35 | 0 | 0 | 5 | [13 empty] | 0 | 2.35
[6 empty] | 2.36 | 2.36 | 5 | 393.5157 | 127.2685 | -17.2219 | 5 | 10.198 | 0 | 0.34428 | 2.36 | 2.36 | 0 | 0 | 5 | [13 empty] | 0 | 2.36
[6 empty] | 2.37 | 2.37 | 5 | 390.6454 | 124.4269 | -16.8774 | 5 | 10.1762 | 0 | 0.3429 | 2.37 | 2.37 | 0 | 0 | 5 | [13 empty] | 0 | 2.37
[6 empty] | 2.38 | 2.38 | 5 | 387.8325 | 121.6421 | -16.5399 | 5 | 10.1545 | 0 | 0.34155 | 2.38 | 2.38 | 0 | 0 | 5 | [13 empty] | 0 | 2.38
[6 empty] | 2.39 | 2.39 | 5 | 385.0759 | 118.913 | -16.2091 | 5 | 10.133 | 0 | 0.34023 | 2.39 | 2.39 | 0 | 0 | 5 | [13 empty] | 0 | 2.39
[6 empty] | 2.4 | 2.4 | 5 | 382.3743 | 116.2385 | -15.8849 | 5 | 10.1117 | 0 | 0.33893 | 2.4 | 2.4 | 0 | 0 | 5 | [13 empty] | 0 | 2.4
[6 empty] | 2.41 | 2.41 | 5 | 379.7268 | 113.6175 | -15.5672 | 5 | 10.0906 | 0 | 0.33765 | 2.41 | 2.41 | 0 | 0 | 5 | [13 empty] | 0 | 2.41
[6 empty] | 2.42 | 2.42 | 5 | 377.1323 | 111.0489 | -15.2559 | 5 | 10.0697 | 0 | 0.33641 | 2.42 | 2.42 | 0 | 0 | 5 | [13 empty] | 0 | 2.42
[6 empty] | 2.43 | 2.43 | 5 | 374.5897 | 108.5317 | -14.9508 | 5 | 10.0489 | 0 | 0.33518 | 2.43 | 2.43 | 0 | 0 | 5 | [13 empty] | 0 | 2.43
[6 empty] | 2.44 | 2.44 | 5 | 372.0979 | 106.0648 | -14.6517 | 5 | 10.0283 | 0 | 0.33399 | 2.44 | 2.44 | 0 | 0 | 5 | [13 empty] | 0 | 2.44
[6 empty] | 2.45 | 2.45 | 5 | 369.6559 | 103.6473 | -14.3587 | 5 | 10.0078 | 0 | 0.33281 | 2.45 | 2.45 | 0 | 0 | 5 | [13 empty] | 0 | 2.45
[6 empty] | 2.46 | 2.46 | 5 | 367.2628 | 101.2781 | -14.0715 | 5 | 9.9876 | 0 | 0.33166 | 2.46 | 2.46 | 0 | 0 | 5 | [13 empty] | 0 | 2.46
[6 empty] | 2.47 | 2.47 | 5 | 364.9175 | 98.9563 | -13.7901 | 5 | 9.9675 | 0 | 0.33054 | 2.47 | 2.47 | 0 | 0 | 5 | [13 empty] | 0 | 2.47
[6 empty] | 2.48 | 2.48 | 5 | 362.6192 | 96.6809 | -13.5143 | 5 | 9.9475 | 0 | 0.32943 | 2.48 | 2.48 | 0 | 0 | 5 | [13 empty] | 0 | 2.48
[6 empty] | 2.49 | 2.49 | 5 | 360.3668 | 94.4511 | -13.244 | 5 | 9.9277 | 0 | 0.32835 | 2.49 | 2.49 | 0 | 0 | 5 | [13 empty] | 0 | 2.49
[6 empty] | 2.5 | 2.5 | 5 | 358.1595 | 92.2658 | -12.9791 | 5 | 9.9081 | 0 | 0.32729 | 2.5 | 2.5 | 0 | 0 | 5 | [13 empty] | 0 | 2.5
[6 empty] | 2.51 | 2.51 | 5 | 355.9963 | 90.1242 | -12.7196 | 5 | 9.8886 | 0 | 0.32625 | 2.51 | 2.51 | 0 | 0 | 5 | [13 empty] | 0 | 2.51
[6 empty] | 2.52 | 2.52 | 5 | 353.8764 | 88.0255 | -12.4652 | 5 | 9.8693 | 0 | 0.32523 | 2.52 | 2.52 | 0 | 0 | 5 | [13 empty] | 0 | 2.52
[6 empty] | 2.53 | 2.53 | 5 | 351.7988 | 85.9688 | -12.2159 | 5 | 9.8501 | 0 | 0.32423 | 2.53 | 2.53 | 0 | 0 | 5 | [13 empty] | 0 | 2.53
[6 empty] | 2.54 | 2.54 | 5 | 349.7628 | 83.9531 | -11.9715 | 5 | 9.8311 | 0 | 0.32325 | 2.54 | 2.54 | 0 | 0 | 5 | [13 empty] | 0 | 2.54
[6 empty] | 2.55 | 2.55 | 5 | 347.7676 | 81.9778 | -11.7321 | 5 | 9.8122 | 0 | 0.32229 | 2.55 | 2.55 | 0 | 0 | 5 | [13 empty] | 0 | 2.55
[6 empty] | 2.56 | 2.56 | 5 | 345.8122 | 80.042 | -11.4975 | 5 | 9.7935 | 0 | 0.32135 | 2.56 | 2.56 | 0 | 0 | 5 | [13 empty] | 0 | 2.56
[6 empty] | 2.57 | 2.57 | 5 | 343.896 | 78.145 | -11.2675 | 5 | 9.7749 | 0 | 0.32043 | 2.57 | 2.57 | 0 | 0 | 5 | [13 empty] | 0 | 2.57
[6 empty] | 2.58 | 2.58 | 5 | 342.0181 | 76.2858 | -11.0422 | 5 | 9.7565 | 0 | 0.31953 | 2.58 | 2.58 | 0 | 0 | 5 | [13 empty] | 0 | 2.58
[6 empty] | 2.59 | 2.59 | 5 | 340.1777 | 74.4639 | -10.8213 | 5 | 9.7382 | 0 | 0.31864 | 2.59 | 2.59 | 0 | 0 | 5 | [13 empty] | 0 | 2.59
[6 empty] | 2.6 | 2.6 | 5 | 338.3742 | 72.6783 | -10.6049 | 5 | 9.72 | 0 | 0.31778 | 2.6 | 2.6 | 0 | 0 | 5 | [13 empty] | 0 | 2.6
[6 empty] | 2.61 | 2.61 | 5 | 336.6067 | 70.9285 | -10.3928 | 5 | 9.702 | 0 | 0.31693 | 2.61 | 2.61 | 0 | 0 | 5 | [13 empty] | 0 | 2.61
[6 empty] | 2.62 | 2.62 | 5 | 334.8745 | 69.2137 | -10.1849 | 5 | 9.6841 | 0 | 0.31609 | 2.62 | 2.62 | 0 | 0 | 5 | [13 empty] | 0 | 2.62
[6 empty] | 2.63 | 2.63 | 5 | 333.1771 | 67.5332 | -9.9812 | 5 | 9.6664 | 0 | 0.31528 | 2.63 | 2.63 | 0 | 0 | 5 | [13 empty] | 0 | 2.63
[6 empty] | 2.64 | 2.64 | 5 | 331.5135 | 65.8863 | -9.7816 | 5 | 9.6488 | 0 | 0.31448 | 2.64 | 2.64 | 0 | 0 | 5 | [13 empty] | 0 | 2.64
[6 empty] | 2.65 | 2.65 | 5 | 329.8832 | 64.2723 | -9.586 | 5 | 9.6313 | 0 | 0.31369 | 2.65 | 2.65 | 0 | 0 | 5 | [13 empty] | 0 | 2.65
[6 empty] | 2.66 | 2.66 | 5 | 328.2856 | 62.6906 | -9.3943 | 5 | 9.614 | 0 | 0.31293 | 2.66 | 2.66 | 0 | 0 | 5 | [13 empty] | 0 | 2.66
[6 empty] | 2.67 | 2.67 | 5 | 326.7199 | 61.1406 | -9.2064 | 5 | 9.5968 | 0 | 0.31217 | 2.67 | 2.67 | 0 | 0 | 5 | [13 empty] | 0 | 2.67
[6 empty] | 2.68 | 2.68 | 5 | 325.1855 | 59.6215 | -9.0223 | 5 | 9.5797 | 0 | 0.31144 | 2.68 | 2.68 | 0 | 0 | 5 | [13 empty] | 0 | 2.68
[6 empty] | 2.69 | 2.69 | 5 | 323.6818 | 58.1329 | -8.8418 | 5 | 9.5627 | 0 | 0.31071 | 2.69 | 2.69 | 0 | 0 | 5 | [13 empty] | 0 | 2.69
[6 empty] | 2.7 | 2.7 | 5 | 322.2081 | 56.674 | -8.665 | 5 | 9.5459 | 0 | 0.31001 | 2.7 | 2.7 | 0 | 0 | 5 | [13 empty] | 0 | 2.7
[6 empty] | 2.71 | 2.71 | 5 | 320.764 | 55.2442 | -8.4917 | 5 | 9.5292 | 0 | 0.30931 | 2.71 | 2.71 | 0 | 0 | 5 | [13 empty] | 0 | 2.71
[6 empty] | 2.72 | 2.72 | 5 | 319.3487 | 53.8431 | -8.3218 | 5 | 9.5126 | 0 | 0.30863 | 2.72 | 2.72 | 0 | 0 | 5 | [13 empty] | 0 | 2.72
[6 empty] | 2.73 | 2.73 | 5 | 317.9617 | 52.47 | -8.1554 | 5 | 9.4961 | 0 | 0.30796 | 2.73 | 2.73 | 0 | 0 | 5 | [13 empty] | 0 | 2.73
[6 empty] | 2.74 | 2.74 | 5 | 316.6025 | 51.1244 | -7.9923 | 5 | 9.4797 | 0 | 0.30731 | 2.74 | 2.74 | 0 | 0 | 5 | [13 empty] | 0 | 2.74
[6 empty] | 2.75 | 2.75 | 5 | 315.2704 | 49.8056 | -7.8325 | 5 | 9.4635 | 0 | 0.30667 | 2.75 | 2.75 | 0 | 0 | 5 | [13 empty] | 0 | 2.75
[6 empty] | 2.76 | 2.76 | 5 | 313.965 | 48.5133 | -7.6758 | 5 | 9.4474 | 0 | 0.30604 | 2.76 | 2.76 | 0 | 0 | 5 | [13 empty] | 0 | 2.76
[6 empty] | 2.77 | 2.77 | 5 | 312.6857 | 47.2468 | -7.5223 | 5 | 9.4314 | 0 | 0.30543 | 2.77 | 2.77 | 0 | 0 | 5 | [13 empty] | 0 | 2.77
[6 empty] | 2.78 | 2.78 | 5 | 311.432 | 46.0056 | -7.3718 | 5 | 9.4155 | 0 | 0.30483 | 2.78 | 2.78 | 0 | 0 | 5 | [13 empty] | 0 | 2.78
[6 empty] | 2.79 | 2.79 | 5 | 310.2034 | 44.7892 | -7.2244 | 5 | 9.3997 | 0 | 0.30423 | 2.79 | 2.79 | 0 | 0 | 5 | [13 empty] | 0 | 2.79
[6 empty] | 2.8 | 2.8 | 5 | 308.9993 | 43.5972 | -7.0799 | 5 | 9.3841 | 0 | 0.30366 | 2.8 | 2.8 | 0 | 0 | 5 | [13 empty] | 0 | 2.8
[6 empty] | 2.81 | 2.81 | 5 | 307.8193 | 42.429 | -6.9383 | 5 | 9.3685 | 0 | 0.30309 | 2.81 | 2.81 | 0 | 0 | 5 | [13 empty] | 0 | 2.81
[6 empty] | 2.82 | 2.82 | 5 | 306.6629 | 41.2842 | -6.7996 | 5 | 9.3531 | 0 | 0.30253 | 2.82 | 2.82 | 0 | 0 | 5 | [13 empty] | 0 | 2.82
[6 empty] | 2.83 | 2.83 | 5 | 305.5297 | 40.1623 | -6.6636 | 5 | 9.3378 | 0 | 0.30199 | 2.83 | 2.83 | 0 | 0 | 5 | [13 empty] | 0 | 2.83
[6 empty] | 2.84 | 2.84 | 5 | 304.4191 | 39.0628 | -6.5303 | 5 | 9.3226 | 0 | 0.30145 | 2.84 | 2.84 | 0 | 0 | 5 | [13 empty] | 0 | 2.84
[6 empty] | 2.85 | 2.85 | 5 | 303.3307 | 37.9853 | -6.3997 | 5 | 9.3074 | 0 | 0.30093 | 2.85 | 2.85 | 0 | 0 | 5 | [13 empty] | 0 | 2.85
[6 empty] | 2.86 | 2.86 | 5 | 302.2641 | 36.9294 | -6.2717 | 5 | 9.2924 | 0 | 0.30042 | 2.86 | 2.86 | 0 | 0 | 5 | [13 empty] | 0 | 2.86
[6 empty] | 2.87 | 2.87 | 5 | 301.2188 | 35.8945 | -6.1463 | 5 | 9.2775 | 0 | 0.29992 | 2.87 | 2.87 | 0 | 0 | 5 | [13 empty] | 0 | 2.87
[6 empty] | 2.88 | 2.88 | 5 | 300.1944 | 34.8804 | -6.0233 | 5 | 9.2627 | 0 | 0.29942 | 2.88 | 2.88 | 0 | 0 | 5 | [13 empty] | 0 | 2.88
[6 empty] | 2.89 | 2.89 | 5 | 299.1905 | 33.8865 | -5.9029 | 5 | 9.248 | 0 | 0.29894 | 2.89 | 2.89 | 0 | 0 | 5 | [13 empty] | 0 | 2.89
[6 empty] | 2.9 | 2.9 | 5 | 298.2067 | 32.9126 | -5.7848 | 5 | 9.2334 | 0 | 0.29847 | 2.9 | 2.9 | 0 | 0 | 5 | [13 empty] | 0 | 2.9
[6 empty] | 2.91 | 2.91 | 5 | 297.2426 | 31.9581 | -5.6691 | 5 | 9.2189 | 0 | 0.298 | 2.91 | 2.91 | 0 | 0 | 5 | [13 empty] | 0 | 2.91
[6 empty] | 2.92 | 2.92 | 5 | 296.2977 | 31.0227 | -5.5557 | 5 | 9.2045 | 0 | 0.29755 | 2.92 | 2.92 | 0 | 0 | 5 | [13 empty] | 0 | 2.92
[6 empty] | 2.93 | 2.93 | 5 | 295.3718 | 30.106 | -5.4446 | 5 | 9.1902 | 0 | 0.29711 | 2.93 | 2.93 | 0 | 0 | 5 | [13 empty] | 0 | 2.93
[6 empty] | 2.94 | 2.94 | 5 | 294.4643 | 29.2076 | -5.3357 | 5 | 9.176 | 0 | 0.29667 | 2.94 | 2.94 | 0 | 0 | 5 | [13 empty] | 0 | 2.94
[6 empty] | 2.95 | 2.95 | 5 | 293.5751 | 28.3272 | -5.229 | 5 | 9.1619 | 0 | 0.29624 | 2.95 | 2.95 | 0 | 0 | 5 | [13 empty] | 0 | 2.95
[6 empty] | 2.96 | 2.96 | 5 | 292.7036 | 27.4644 | -5.1244 | 5 | 9.1479 | 0 | 0.29582 | 2.96 | 2.96 | 0 | 0 | 5 | [13 empty] | 0 | 2.96
[6 empty] | 2.97 | 2.97 | 5 | 291.8495 | 26.6189 | -5.0219 | 5 | 9.134 | 0 | 0.29541 | 2.97 | 2.97 | 0 | 0 | 5 | [13 empty] | 0 | 2.97
[6 empty] | 2.98 | 2.98 | 5 | 291.0125 | 25.7903 | -4.9215 | 5 | 9.1202 | 0 | 0.29501 | 2.98 | 2.98 | 0 | 0 | 5 | [13 empty] | 0 | 2.98
[6 empty] | 2.99 | 2.99 | 5 | 290.1922 | 24.9782 | -4.8231 | 5 | 9.1064 | 0 | 0.29462 | 2.99 | 2.99 | 0 | 0 | 5 | [13 empty] | 0 | 2.99
[6 empty] | 3 | 3 | 5 | 289.3884 | 24.1824 | -4.7266 | 5 | 9.0928 | 0 | 0.29423 | 3 | 3 | 0 | 0 | 5 | [13 empty] | 0 | 3
[6 empty] | 3.01 | 3.01 | 5 | 288.6006 | 23.4025 | -4.6321 | 5 | 9.0792 | 0 | 0.29385 | 3.01 | 3.01 | 0 | 0 | 5 | [14 empty] | 3.01
[6 empty] | 3.02 | 3.02 | 5 | 287.8286 | 22.6382 | -4.5394 | 5 | 9.0658 | 0 | 0.29348 | 3.02 | 3.02 | 0 | 0 | 5 | [14 empty] | 3.02
[6 empty] | 3.03 | 3.03 | 5 | 287.072 | 21.8892 | -4.4486 | 5 | 9.0524 | 0 | 0.29312 | 3.03 | 3.03 | 0 | 0 | 5 | [14 empty] | 3.03
[6 empty] | 3.04 | 3.04 | 5 | 286.3306 | 21.1552 | -4.3597 | 5 | 9.0391 | 0 | 0.29276 | 3.04 | 3.04 | 0 | 0 | 5 | [14 empty] | 3.04
[6 empty] | 3.05 | 3.05 | 5 | 285.604 | 20.4359 | -4.2725 | 5 | 9.0259 | 0 | 0.29241 | 3.05 | 3.05 | 0 | 0 | 5 | [14 empty] | 3.05
[6 empty] | 3.06 | 3.06 | 5 | 284.8919 | 19.7309 | -4.187 | 5 | 9.0128 | 0 | 0.29207 | 3.06 | 3.06 | 0 | 0 | 5 | [14 empty] | 3.06
[6 empty] | 3.07 | 3.07 | 5 | 284.1941 | 19.0401 | -4.1033 | 5 | 8.9998 | 0 | 0.29173 | 3.07 | 3.07 | 0 | 0 | 5 | [14 empty] | 3.07
[6 empty] | 3.08 | 3.08 | 5 | 283.5102 | 18.363 | -4.0212 | 5 | 8.9868 | 0 | 0.2914 | 3.08 | 3.08 | 0 | 0 | 5 | [14 empty] | 3.08
[6 empty] | 3.09 | 3.09 | 5 | 282.84 | 17.6995 | -3.9408 | 5 | 8.974 | 0 | 0.29108 | 3.09 | 3.09 | 0 | 0 | 5 | [14 empty] | 3.09
[6 empty] | 3.1 | 3.1 | 5 | 282.1832 | 17.0493 | -3.862 | 5 | 8.9612 | 0 | 0.29077 | 3.1 | 3.1 | 0 | 0 | 5 | [14 empty] | 3.1
[6 empty] | 3.11 | 3.11 | 5 | 281.5395 | 16.412 | -3.7847 | 5 | 8.9485 | 0 | 0.29046 | 3.11 | 3.11 | 0 | 0 | 5 | [14 empty] | 3.11
[6 empty] | 3.12 | 3.12 | 5 | 280.9087 | 15.7876 | -3.709 | 5 | 8.9359 | 0 | 0.29015 | 3.12 | 3.12 | 0 | 0 | 5 | [14 empty] | 3.12
[6 empty] | 3.13 | 3.13 | 5 | 280.2906 | 15.1756 | -3.6349 | 5 | 8.9233 | 0 | 0.28986 | 3.13 | 3.13 | 0 | 0 | 5 | [14 empty] | 3.13
[6 empty] | 3.14 | 3.14 | 5 | 279.6847 | 14.5758 | -3.5622 | 5 | 8.9109 | 0 | 0.28956 | 3.14 | 3.14 | 0 | 0 | 5 | [14 empty] | 3.14
[6 empty] | 3.15 | 3.15 | 5 | 279.0911 | 13.9881 | -3.4909 | 5 | 8.8985 | 0 | 0.28928 | 3.15 | 3.15 | 0 | 0 | 5 | [14 empty] | 3.15
[6 empty] | 3.16 | 3.16 | 5 | 278.5092 | 13.4121 | -3.4211 | 5 | 8.8862 | 0 | 0.289 | 3.16 | 3.16 | 0 | 0 | 5 | [14 empty] | 3.16
[6 empty] | 3.17 | 3.17 | 5 | 277.939 | 12.8476 | -3.3527 | 5 | 8.874 | 0 | 0.28873 | 3.17 | 3.17 | 0 | 0 | 5 | [14 empty] | 3.17
[6 empty] | 3.18 | 3.18 | 5 | 277.3803 | 12.2944 | -3.2856 | 5 | 8.8618 | 0 | 0.28846 | 3.18 | 3.18 | 0 | 0 | 5 | [14 empty] | 3.18
[6 empty] | 3.19 | 3.19 | 5 | 276.8327 | 11.7523 | -3.2199 | 5 | 8.8498 | 0 | 0.28819 | 3.19 | 3.19 | 0 | 0 | 5 | [14 empty] | 3.19
[6 empty] | 3.2 | 3.2 | 5 | 276.296 | 11.221 | -3.1555 | 5 | 8.8378 | 0 | 0.28794 | 3.2 | 3.2 | 0 | 0 | 5 | [14 empty] | 3.2
[6 empty] | 3.21 | 3.21 | 5 | 275.7701 | 10.7003 | -3.0924 | 5 | 8.8259 | 0 | 0.28768 | 3.21 | 3.21 | 0 | 0 | 5 | [14 empty] | 3.21
[6 empty] | 3.22 | 3.22 | 5 | 275.2547 | 10.1901 | -3.0306 | 5 | 8.814 | 0 | 0.28744 | 3.22 | 3.22 | 0 | 0 | 5 | [14 empty] | 3.22
[6 empty] | 3.23 | 3.23 | 5 | 274.7496 | 9.69 | -2.97 | 5 | 8.8023 | 0 | 0.28719 | 3.23 | 3.23 | 0 | 0 | 5 | [14 empty] | 3.23
[6 empty] | 3.24 | 3.24 | 5 | 274.2546 | 9.2 | -2.9106 | 5 | 8.7906 | 0 | 0.28695 | 3.24 | 3.24 | 0 | 0 | 5 | [14 empty] | 3.24
[6 empty] | 3.25 | 3.25 | 5 | 273.7695 | 8.7197 | -2.8523 | 5 | 8.7789 | 0 | 0.28672 | 3.25 | 3.25 | 0 | 0 | 5 | [14 empty] | 3.25
[6 empty] | 3.26 | 3.26 | 5 | 273.2941 | 8.2491 | -2.7953 | 5 | 8.7674 | 0 | 0.28649 | 3.26 | 3.26 | 0 | 0 | 5 | [14 empty] | 3.26
[6 empty] | 3.27 | 3.27 | 5 | 272.8282 | 7.7879 | -2.7394 | 5 | 8.7559 | 0 | 0.28627 | 3.27 | 3.27 | 0 | 0 | 5 | [14 empty] | 3.27
[6 empty] | 3.28 | 3.28 | 5 | 272.3717 | 7.3359 | -2.6846 | 5 | 8.7445 | 0 | 0.28605 | 3.28 | 3.28 | 0 | 0 | 5 | [14 empty] | 3.28
[6 empty] | 3.29 | 3.29 | 5 | 271.9242 | 6.8929 | -2.6309 | 5 | 8.7331 | 0 | 0.28583 | 3.29 | 3.29 | 0 | 0 | 5 | [14 empty] | 3.29
[6 empty] | 3.3 | 3.3 | 5 | 271.4858 | 6.4588 | -2.5783 | 5 | 8.7218 | 0 | 0.28562 | 3.3 | 3.3 | 0 | 0 | 5 | [14 empty] | 3.3
[6 empty] | 3.31 | 3.31 | 5 | 271.056 | 6.0334 | -2.5267 | 5 | 8.7106 | 0 | 0.28542 | 3.31 | 3.31 | 0 | 0 | 5 | [14 empty] | 3.31
[6 empty] | 3.32 | 3.32 | 5 | 270.6349 | 5.6165 | -2.4762 | 5 | 8.6995 | 0 | 0.28521 | 3.32 | 3.32 | 0 | 0 | 5 | [14 empty] | 3.32
[6 empty] | 3.33 | 3.33 | 5 | 270.2222 | 5.2079 | -2.4267 | 5 | 8.6884 | 0 | 0.28502 | 3.33 | 3.33 | 0 | 0 | 5 | [14 empty] | 3.33
[6 empty] | 3.34 | 3.34 | 5 | 269.8178 | 4.8075 | -2.3781 | 5 | 8.6774 | 0 | 0.28482 | 3.34 | 3.34 | 0 | 0 | 5 | [14 empty] | 3.34
[6 empty] | 3.35 | 3.35 | 5 | 269.4214 | 4.4151 | -2.3306 | 5 | 8.6665 | 0 | 0.28463 | 3.35 | 3.35 | 0 | 0 | 5 | [14 empty] | 3.35
[6 empty] | 3.36 | 3.36 | 5 | 269.033 | 4.0306 | -2.284 | 5 | 8.6556 | 0 | 0.28444 | 3.36 | 3.36 | 0 | 0 | 5 | [14 empty] | 3.36
[6 empty] | 3.37 | 3.37 | 5 | 268.6523 | 3.6537 | -2.2383 | 5 | 8.6448 | 0 | 0.28426 | 3.37 | 3.37 | 0 | 0 | 5 | [14 empty] | 3.37
[6 empty] | 3.38 | 3.38 | 5 | 268.2793 | 3.2844 | -2.1935 | 5 | 8.634 | 0 | 0.28408 | 3.38 | 3.38 | 0 | 0 | 5 | [14 empty] | 3.38
[6 empty] | 3.39 | 3.39 | 5 | 267.9137 | 2.9225 | -2.1496 | 5 | 8.6233 | 0 | 0.28391 | 3.39 | 3.39 | 0 | 0 | 5 | [14 empty] | 3.39
[6 empty] | 3.4 | 3.4 | 5 | 267.5554 | 2.5678 | -2.1067 | 5 | 8.6127 | 0 | 0.28373 | 3.4 | 3.4 | 0 | 0 | 5 | [14 empty] | 3.4
[6 empty] | 3.41 | 3.41 | 5 | 267.2043 | 2.2202 | -2.0645 | 5 | 8.6021 | 0 | 0.28357 | 3.41 | 3.41 | 0 | 0 | 5 | [14 empty] | 3.41
[6 empty] | 3.42 | 3.42 | 5 | 266.8602 | 1.8795 | -2.0232 | 5 | 8.5916 | 0 | 0.2834 | 3.42 | 3.42 | 0 | 0 | 5 | [14 empty] | 3.42
[6 empty] | 3.43 | 3.43 | 5 | 266.523 | 1.5457 | -1.9828 | 5 | 8.5812 | 0 | 0.28324 | 3.43 | 3.43 | 0 | 0 | 5 | [14 empty] | 3.43
[6 empty] | 3.44 | 3.44 | 5 | 266.1926 | 1.2186 | -1.9431 | 5 | 8.5708 | 0 | 0.28308 | 3.44 | 3.44 | 0 | 0 | 5 | [14 empty] | 3.44
[6 empty] | 3.45 | 3.45 | 5 | 265.8687 | 0.89794 | -1.9042 | 5 | 8.5605 | 0 | 0.28292 | 3.45 | 3.45 | 0 | 0 | 5 | [14 empty] | 3.45
[6 empty] | 3.46 | 3.46 | 5 | 265.5513 | 0.58374 | -1.8662 | 5 | 8.5502 | 0 | 0.28277 | 3.46 | 3.46 | 0 | 0 | 5 | [14 empty] | 3.46
[6 empty] | 3.47 | 3.47 | 5 | 265.2403 | 0.27583 | -1.8288 | 5 | 8.54 | 0 | 0.28262 | 3.47 | 3.47 | 0 | 0 | 5 | [14 empty] | 3.47
[6 empty] | 3.48 | 3.48 | 5 | 264.9355 | -0.025931 | -1.7923 | 5 | 8.5299 | 0 | 0.28247 | 3.48 | 3.48 | 0 | 0 | 5 | [14 empty] | 3.48
[6 empty] | 3.49 | 3.49 | 5 | 264.6368 | -0.32165 | -1.7564 | 5 | 8.5198 | 0 | 0.28233 | 3.49 | 3.49 | 0 | 0 | 5 | [14 empty] | 3.49
[6 empty] | 3.5 | 3.5 | 5 | 264.3441 | -0.61146 | -1.7213 | 5 | 8.5098 | 0 | 0.28219 | 3.5 | 3.5 | 0 | 0 | 5 | [14 empty] | 3.5
[6 empty] | 3.51 | 3.51 | 5 | 264.0572 | -0.89548 | -1.6869 | 5 | 8.4998 | 0 | 0.28205 | 3.51 | 3.51 | 0 | 0 | 5 | [14 empty] | 3.51
[6 empty] | 3.52 | 3.52 | 5 | 263.776 | -1.1738 | -1.6531 | 5 | 8.4899 | 0 | 0.28192 | 3.52 | 3.52 | 0 | 0 | 5 | [14 empty] | 3.52
[6 empty] | 0.00058824 | 3.53 | 7.6033 | 263.5005 | -0.1449 | 14 | 7.6033 | 8.4874 | 0 | 0.28366 | 3.53 | 3.53 | 0 | 0 | 7.6033 | [14 empty] | 3.53
[6 empty] | 0.010588 | 3.54 | 7.65 | 265.8338 | 2.1884 | 14 | 7.65 | 8.485 | 0 | 0.28541 | 3.54 | 3.54 | 0 | 0 | 7.65 | [14 empty] | 3.54
[6 empty] | 0.020588 | 3.55 | 7.6967 | 268.1672 | 4.5218 | 14 | 7.6967 | 8.4828 | 0 | 0.28717 | 3.55 | 3.55 | 0 | 0 | 7.6967 | [14 empty] | 3.55
[6 empty] | 0.030588 | 3.56 | 7.7433 | 270.5005 | 6.8551 | 14 | 7.7433 | 8.4807 | 0 | 0.28894 | 3.56 | 3.56 | 0 | 0 | 7.7433 | [14 empty] | 3.56
[6 empty] | 0.040588 | 3.57 | 7.79 | 272.8338 | 9.1884 | 14 | 7.79 | 8.4788 | 0 | 0.29073 | 3.57 | 3.57 | 0 | 0 | 7.79 | [14 empty] | 3.57
[6 empty] | 0.050588 | 3.58 | 7.8367 | 275.1672 | 11.5218 | 14 | 7.8367 | 8.477 | 0 | 0.29252 | 3.58 | 3.58 | 0 | 0 | 7.8367 | [14 empty] | 3.58
[6 empty] | 0.060588 | 3.59 | 7.8833 | 277.5005 | 13.8551 | 14 | 7.8833 | 8.4754 | 0 | 0.29432 | 3.59 | 3.59 | 0 | 0 | 7.8833 | [14 empty] | 3.59
[6 empty] | 0.070588 | 3.6 | 7.93 | 279.8338 | 16.1884 | 14 | 7.93 | 8.4739 | 0 | 0.29614 | 3.6 | 3.6 | 0 | 0 | 7.93 | [14 empty] | 3.6
[6 empty] | 0.080588 | 3.61 | 7.9767 | 282.1672 | 18.5218 | 14 | 7.9767 | 8.4725 | 0 | 0.29797 | 3.61 | 3.61 | 0 | 0 | 7.9767 | [14 empty] | 3.61
[6 empty] | 0.090588 | 3.62 | 8.0233 | 284.5005 | 20.8551 | 14 | 8.0233 | 8.4712 | 0 | 0.2998 | 3.62 | 3.62 | 0 | 0 | 8.0233 | [14 empty] | 3.62
[6 empty] | 0.10059 | 3.63 | 8.07 | 286.8338 | 23.1884 | 14 | 8.07 | 8.4701 | 0 | 0.30165 | 3.63 | 3.63 | 0 | 0 | 8.07 | [14 empty] | 3.63
[6 empty] | 0.11059 | 3.64 | 8.1167 | 289.1672 | 25.5218 | 14 | 8.1167 | 8.4692 | 0 | 0.30351 | 3.64 | 3.64 | 0 | 0 | 8.1167 | [14 empty] | 3.64
[6 empty] | 0.12059 | 3.65 | 8.1633 | 291.5005 | 27.8551 | 14 | 8.1633 | 8.4683 | 0 | 0.30538 | 3.65 | 3.65 | 0 | 0 | 8.1633 | [14 empty] | 3.65
[6 empty] | 0.13059 | 3.66 | 8.21 | 293.8338 | 30.1884 | 14 | 8.21 | 8.4676 | 0 | 0.30726 | 3.66 | 3.66 | 0 | 0 | 8.21 | [14 empty] | 3.66
[6 empty] | 0.14059 | 3.67 | 8.2567 | 296.1672 | 32.5218 | 14 | 8.2567 | 8.4671 | 0 | 0.30915 | 3.67 | 3.67 | 0 | 0 | 8.2567 | [14 empty] | 3.67
[6 empty] | 0.15059 | 3.68 | 8.3033 | 298.5005 | 34.8551 | 14 | 8.3033 | 8.4666 | 0 | 0.31105 | 3.68 | 3.68 | 0 | 0 | 8.3033 | [14 empty] | 3.68
[6 empty] | 0.16059 | 3.69 | 8.35 | 300.8338 | 37.1884 | 14 | 8.35 | 8.4663 | 0 | 0.31296 | 3.69 | 3.69 | 0 | 0 | 8.35 | [14 empty] | 3.69
[6 empty] | 0.17059 | 3.7 | 8.3967 | 303.1672 | 39.5218 | 14 | 8.3967 | 8.4661 | 0 | 0.31488 | 3.7 | 3.7 | 0 | 0 | 8.3967 | [14 empty] | 3.7
[6 empty] | 0.18059 | 3.71 | 8.4433 | 305.5005 | 41.8551 | 14 | 8.4433 | 8.4661 | 0 | 0.31681 | 3.71 | 3.71 | 0 | 0 | 8.4433 | [14 empty] | 3.71
[6 empty] | 0.19059 | 3.72 | 8.49 | 307.8338 | 44.1884 | 14 | 8.49 | 8.4661 | 0 | 0.31876 | 3.72 | 3.72 | 0 | 0 | 8.49 | [14 empty] | 3.72
[6 empty] | 0.20059 | 3.73 | 8.5367 | 310.1672 | 46.5218 | 14 | 8.5367 | 8.4663 | 0 | 0.32071 | 3.73 | 3.73 | 0 | 0 | 8.5367 | [14 empty] | 3.73
[6 empty] | 0.21059 | 3.74 | 8.5833 | 312.5005 | 48.8551 | 14 | 8.5833 | 8.4666 | 0 | 0.32267 | 3.74 | 3.74 | 0 | 0 | 8.5833 | [14 empty] | 3.74
[6 empty] | 0.22059 | 3.75 | 8.63 | 314.8338 | 51.1884 | 14 | 8.63 | 8.4671 | 0 | 0.32465 | 3.75 | 3.75 | 0 | 0 | 8.63 | [14 empty] | 3.75
[6 empty] | 0.23059 | 3.76 | 8.6767 | 317.1672 | 53.5218 | 14 | 8.6767 | 8.4676 | 0 | 0.32664 | 3.76 | 3.76 | 0 | 0 | 8.6767 | [14 empty] | 3.76
[6 empty] | 0.24059 | 3.77 | 8.7233 | 319.5005 | 55.8551 | 14 | 8.7233 | 8.4683 | 0 | 0.32863 | 3.77 | 3.77 | 0 | 0 | 8.7233 | [14 empty] | 3.77
[6 empty] | 0.25059 | 3.78 | 8.77 | 321.8338 | 58.1884 | 14 | 8.77 | 8.4691 | 0 | 0.33064 | 3.78 | 3.78 | 0 | 0 | 8.77 | [14 empty] | 3.78
[6 empty] | 0.26059 | 3.79 | 8.8167 | 324.1672 | 60.5218 | 14 | 8.8167 | 8.47 | 0 | 0.33266 | 3.79 | 3.79 | 0 | 0 | 8.8167 | [14 empty] | 3.79
[6 empty] | 0.27059 | 3.8 | 8.8633 | 326.5005 | 62.8551 | 14 | 8.8633 | 8.471 | 0 | 0.33469 | 3.8 | 3.8 | 0 | 0 | 8.8633 | [14 empty] | 3.8
[6 empty] | 0.28059 | 3.81 | 8.91 | 328.8338 | 65.1884 | 14 | 8.91 | 8.4722 | 0 | 0.33673 | 3.81 | 3.81 | 0 | 0 | 8.91 | [14 empty] | 3.81
[6 empty] | 0.29059 | 3.82 | 8.9567 | 331.1672 | 67.5218 | 14 | 8.9567 | 8.4734 | 0 | 0.33878 | 3.82 | 3.82 | 0 | 0 | 8.9567 | [14 empty] | 3.82
[6 empty] | 0.30059 | 3.83 | 9.0033 | 333.5005 | 69.8551 | 14 | 9.0033 | 8.4748 | 0 | 0.34084 | 3.83 | 3.83 | 0 | 0 | 9.0033 | [14 empty] | 3.83
[6 empty] | 0.31059 | 3.84 | 9.05 | 335.8338 | 72.1884 | 14 | 9.05 | 8.4763 | 0 | 0.34291 | 3.84 | 3.84 | 0 | 0 | 9.05 | [14 empty] | 3.84
[6 empty] | 0.32059 | 3.85 | 9.0967 | 338.1672 | 74.5218 | 14 | 9.0967 | 8.4779 | 0 | 0.34499 | 3.85 | 3.85 | 0 | 0 | 9.0967 | [14 empty] | 3.85
[6 empty] | 0.33059 | 3.86 | 9.1433 | 340.5005 | 76.8551 | 14 | 9.1433 | 8.4796 | 0 | 0.34708 | 3.86 | 3.86 | 0 | 0 | 9.1433 | [14 empty] | 3.86
[6 empty] | 0.34059 | 3.87 | 9.19 | 342.8338 | 79.1884 | 14 | 9.19 | 8.4815 | 0 | 0.34919 | 3.87 | 3.87 | 0 | 0 | 9.19 | [14 empty] | 3.87
[6 empty] | 0.35059 | 3.88 | 9.2367 | 345.1672 | 81.5218 | 14 | 9.2367 | 8.4834 | 0 | 0.3513 | 3.88 | 3.88 | 0 | 0 | 9.2367 | [14 empty] | 3.88
[6 empty] | 0.36059 | 3.89 | 9.2833 | 347.5005 | 83.8551 | 14 | 9.2833 | 8.4855 | 0 | 0.35343 | 3.89 | 3.89 | 0 | 0 | 9.2833 | [14 empty] | 3.89
[6 empty] | 0.37059 | 3.9 | 9.33 | 349.8338 | 86.1884 | 14 | 9.33 | 8.4876 | 0 | 0.35556 | 3.9 | 3.9 | 0 | 0 | 9.33 | [14 empty] | 3.9
[6 empty] | 0.38059 | 3.91 | 9.3767 | 352.1672 | 88.5218 | 14 | 9.3767 | 8.4899 | 0 | 0.35771 | 3.91 | 3.91 | 0 | 0 | 9.3767 | [14 empty] | 3.91
[6 empty] | 0.39059 | 3.92 | 9.4233 | 354.5005 | 90.8551 | 14 | 9.4233 | 8.4923 | 0 | 0.35987 | 3.92 | 3.92 | 0 | 0 | 9.4233 | [14 empty] | 3.92
[6 empty] | 0.40059 | 3.93 | 9.47 | 356.8338 | 93.1884 | 14 | 9.47 | 8.4948 | 0 | 0.36203 | 3.93 | 3.93 | 0 | 0 | 9.47 | [14 empty] | 3.93
[6 empty] | 0.41059 | 3.94 | 9.5167 | 359.1672 | 95.5218 | 14 | 9.5167 | 8.4973 | 0 | 0.36421 | 3.94 | 3.94 | 0 | 0 | 9.5167 | [14 empty] | 3.94
[6 empty] | 0.42059 | 3.95 | 9.5633 | 361.5005 | 97.8551 | 14 | 9.5633 | 8.5 | 0 | 0.3664 | 3.95 | 3.95 | 0 | 0 | 9.5633 | [14 empty] | 3.95
[6 empty] | 0.43059 | 3.96 | 9.61 | 363.8338 | 100.1884 | 14 | 9.61 | 8.5028 | 0 | 0.3686 | 3.96 | 3.96 | 0 | 0 | 9.61 | [14 empty] | 3.96
[6 empty] | 0.44059 | 3.97 | 9.6567 | 366.1672 | 102.5218 | 14 | 9.6567 | 8.5057 | 0 | 0.37081 | 3.97 | 3.97 | 0 | 0 | 9.6567 | [14 empty] | 3.97
[6 empty] | 0.45059 | 3.98 | 9.7033 | 368.5005 | 104.8551 | 14 | 9.7033 | 8.5087 | 0 | 0.37303 | 3.98 | 3.98 | 0 | 0 | 9.7033 | [14 empty] | 3.98
[6 empty] | 0.46059 | 3.99 | 9.75 | 370.8338 | 107.1884 | 14 | 9.75 | 8.5118 | 0 | 0.37526 | 3.99 | 3.99 | 0 | 0 | 9.75 | [14 empty] | 3.99
[6 empty] | 0.47059 | 4 | 9.7967 | 373.1672 | 109.5218 | 14 | 9.7967 | 8.515 | 0 | 0.37751 | 4 | 4 | 0 | 0 | 9.7967 | [14 empty] | 4
[6 empty] | 0.48059 | 4.01 | 9.8433 | 375.5005 | 111.8551 | 14 | 9.8433 | 8.5183 | 0 | 0.37976 | 4.01 | 4.01 | 0 | 0 | 9.8433 | [14 empty] | 4.01
[6 empty] | 0.49059 | 4.02 | 9.89 | 377.8338 | 114.1884 | 14 | 9.89 | 8.5217 | 0 | 0.38202 | 4.02 | 4.02 | 0 | 0 | 9.89 | [14 empty] | 4.02
[6 empty] | 0.50059 | 4.03 | 9.9367 | 380.1672 | 116.5218 | 14 | 9.9367 | 8.5252 | 0 | 0.3843 | 4.03 | 4.03 | 0 | 0 | 9.9367 | [14 empty] | 4.03
[6 empty] | 0.51059 | 4.04 | 9.9833 | 382.5005 | 118.8551 | 14 | 9.9833 | 8.5288 | 0 | 0.38658 | 4.04 | 4.04 | 0 | 0 | 9.9833 | [14 empty] | 4.04
[6 empty] | 0.52059 | 4.05 | 10.03 | 384.8338 | 121.1884 | 14 | 10.03 | 8.5325 | 0 | 0.38888 | 4.05 | 4.05 | 0 | 0 | 10.03 | [14 empty] | 4.05
[6 empty] | 0.53059 | 4.06 | 10.0767 | 387.1672 | 123.5218 | 14 | 10.0767 | 8.5363 | 0 | 0.39119 | 4.06 | 4.06 | 0 | 0 | 10.0767 | [14 empty] | 4.06
[6 empty] | 0.54059 | 4.07 | 10.1233 | 389.5005 | 125.8551 | 14 | 10.1233 | 8.5402 | 0 | 0.3935 | 4.07 | 4.07 | 0 | 0 | 10.1233 | [14 empty] | 4.07
[6 empty] | 0.55059 | 4.08 | 10.17 | 391.8338 | 128.1884 | 14 | 10.17 | 8.5442 | 0 | 0.39583 | 4.08 | 4.08 | 0 | 0 | 10.17 | [14 empty] | 4.08
[6 empty] | 0.56059 | 4.09 | 10.2167 | 394.1672 | 130.5218 | 14 | 10.2167 | 8.5483 | 0 | 0.39817 | 4.09 | 4.09 | 0 | 0 | 10.2167 | [14 empty] | 4.09
[6 empty] | 0.57059 | 4.1 | 10.2633 | 396.5005 | 132.8551 | 14 | 10.2633 | 8.5525 | 0 | 0.40052 | 4.1 | 4.1 | 0 | 0 | 10.2633 | [14 empty] | 4.1
[6 empty] | 0.58059 | 4.11 | 10.31 | 398.8338 | 135.1884 | 14 | 10.31 | 8.5567 | 0 | 0.40288 | 4.11 | 4.11 | 0 | 0 | 10.31 | [14 empty] | 4.11
[6 empty] | 0.59059 | 4.12 | 10.3567 | 401.1672 | 137.5218 | 14 | 10.3567 | 8.5611 | 0 | 0.40525 | 4.12 | 4.12 | 0 | 0 | 10.3567 | [14 empty] | 4.12
[6 empty] | 0.60059 | 4.13 | 10.4033 | 403.5005 | 139.8551 | 14 | 10.4033 | 8.5655 | 0 | 0.40763 | 4.13 | 4.13 | 0 | 0 | 10.4033 | [14 empty] | 4.13
[6 empty] | 0.61059 | 4.14 | 10.45 | 405.8338 | 142.1884 | 14 | 10.45 | 8.5701 | 0 | 0.41002 | 4.14 | 4.14 | 0 | 0 | 10.45 | [14 empty] | 4.14
[6 empty] | 0.62059 | 4.15 | 10.4967 | 408.1672 | 144.5218 | 14 | 10.4967 | 8.5747 | 0 | 0.41243 | 4.15 | 4.15 | 0 | 0 | 10.4967 | [14 empty] | 4.15
[6 empty] | 0.63059 | 4.16 | 10.5433 | 410.5005 | 146.8551 | 14 | 10.5433 | 8.5794 | 0 | 0.41484 | 4.16 | 4.16 | 0 | 0 | 10.5433 | [14 empty] | 4.16
[6 empty] | 0.64059 | 4.17 | 10.59 | 412.8338 | 149.1884 | 14 | 10.59 | 8.5842 | 0 | 0.41726 | 4.17 | 4.17 | 0 | 0 | 10.59 | [14 empty] | 4.17
[6 empty] | 0.65059 | 4.18 | 10.6367 | 415.1672 | 151.5218 | 14 | 10.6367 | 8.5891 | 0 | 0.4197 | 4.18 | 4.18 | 0 | 0 | 10.6367 | [14 empty] | 4.18
[6 empty] | 0.66059 | 4.19 | 10.6833 | 417.5005 | 153.8551 | 14 | 10.6833 | 8.5941 | 0 | 0.42214 | 4.19 | 4.19 | 0 | 0 | 10.6833 | [14 empty] | 4.19
[6 empty] | 0.67059 | 4.2 | 10.73 | 419.8338 | 156.1884 | 14 | 10.73 | 8.5992 | 0 | 0.4246 | 4.2 | 4.2 | 0 | 0 | 10.73 | [14 empty] | 4.2
[6 empty] | 0.68059 | 4.21 | 10.7767 | 422.1672 | 158.5218 | 14 | 10.7767 | 8.6044 | 0 | 0.42707 | 4.21 | 4.21 | 0 | 0 | 10.7767 | [14 empty] | 4.21
[6 empty] | 0.69059 | 4.22 | 10.8233 | 424.5005 | 160.8551 | 14 | 10.8233 | 8.6096 | 0 | 0.42954 | 4.22 | 4.22 | 0 | 0 | 10.8233 | [14 empty] | 4.22
[6 empty] | 0.70059 | 4.23 | 10.87 | 426.8338 | 163.1884 | 14 | 10.87 | 8.6149 | 0 | 0.43203 | 4.23 | 4.23 | 0 | 0 | 10.87 | [14 empty] | 4.23
[6 empty] | 0.71059 | 4.24 | 10.9167 | 429.1672 | 165.5218 | 14 | 10.9167 | 8.6204 | 0 | 0.43453 | 4.24 | 4.24 | 0 | 0 | 10.9167 | [14 empty] | 4.24
[6 empty] | 0.72059 | 4.25 | 10.9633 | 431.5005 | 167.8551 | 14 | 10.9633 | 8.6259 | 0 | 0.43704 | 4.25 | 4.25 | 0 | 0 | 10.9633 | [14 empty] | 4.25
[6 empty] | 0.73059 | 4.26 | 11.01 | 433.8338 | 170.1884 | 14 | 11.01 | 8.6314 | 0 | 0.43956 | 4.26 | 4.26 | 0 | 0 | 11.01 | [14 empty] | 4.26
[6 empty] | 0.74059 | 4.27 | 11.0567 | 436.1672 | 172.5218 | 14 | 11.0567 | 8.6371 | 0 | 0.44209 | 4.27 | 4.27 | 0 | 0 | 11.0567 | [14 empty] | 4.27
[6 empty] | 0.75059 | 4.28 | 11.1033 | 438.5005 | 174.8551 | 14 | 11.1033 | 8.6429 | 0 | 0.44463 | 4.28 | 4.28 | 0 | 0 | 11.1033 | [14 empty] | 4.28
[6 empty] | 0.76059 | 4.29 | 11.15 | 440.8338 | 177.1884 | 14 | 11.15 | 8.6487 | 0 | 0.44718 | 4.29 | 4.29 | 0 | 0 | 11.15 | [14 empty] | 4.29
[6 empty] | 0.77059 | 4.3 | 11.1967 | 443.1672 | 179.5218 | 14 | 11.1967 | 8.6546 | 0 | 0.44975 | 4.3 | 4.3 | 0 | 0 | 11.1967 | [14 empty] | 4.3
[6 empty] | 0.78059 | 4.31 | 11.2433 | 445.5005 | 181.8551 | 14 | 11.2433 | 8.6606 | 0 | 0.45232 | 4.31 | 4.31 | 0 | 0 | 11.2433 | [14 empty] | 4.31
[6 empty] | 0.79059 | 4.32 | 11.29 | 447.8338 | 184.1884 | 14 | 11.29 | 8.6667 | 0 | 0.45491 | 4.32 | 4.32 | 0 | 0 | 11.29 | [14 empty] | 4.32
[6 empty] | 0.80059 | 4.33 | 11.3367 | 450.1672 | 186.5218 | 14 | 11.3367 | 8.6728 | 0 | 0.4575 | 4.33 | 4.33 | 0 | 0 | 11.3367 | [14 empty] | 4.33
[6 empty] | 0.81059 | 4.34 | 11.3833 | 452.5005 | 188.8551 | 14 | 11.3833 | 8.679 | 0 | 0.46011 | 4.34 | 4.34 | 0 | 0 | 11.3833 | [14 empty] | 4.34
[6 empty] | 0.82059 | 4.35 | 11.43 | 454.8338 | 191.1884 | 14 | 11.43 | 8.6854 | 0 | 0.46272 | 4.35 | 4.35 | 0 | 0 | 11.43 | [14 empty] | 4.35
[6 empty] | 0.83059 | 4.36 | 11.4767 | 457.1672 | 193.5218 | 14 | 11.4767 | 8.6917 | 0 | 0.46535 | 4.36 | 4.36 | 0 | 0 | 11.4767 | [14 empty] | 4.36
[6 empty] | 0.84059 | 4.37 | 11.5233 | 459.5005 | 195.8551 | 14 | 11.5233 | 8.6982 | 0 | 0.46799 | 4.37 | 4.37 | 0 | 0 | 11.5233 | [14 empty] | 4.37
[6 empty] | 0.85059 | 4.38 | 11.57 | 461.8338 | 198.1884 | 14 | 11.57 | 8.7047 | 0 | 0.47064 | 4.38 | 4.38 | 0 | 0 | 11.57 | [14 empty] | 4.38
[6 empty] | 0.86059 | 4.39 | 11.6167 | 464.1672 | 200.5218 | 14 | 11.6167 | 8.7114 | 0 | 0.47329 | 4.39 | 4.39 | 0 | 0 | 11.6167 | [14 empty] | 4.39
[6 empty] | 0.87059 | 4.4 | 11.6633 | 466.5005 | 202.8551 | 14 | 11.6633 | 8.7181 | 0 | 0.47596 | 4.4 | 4.4 | 0 | 0 | 11.6633 | [14 empty] | 4.4
[6 empty] | 0.88059 | 4.41 | 11.71 | 468.8338 | 205.1884 | 14 | 11.71 | 8.7248 | 0 | 0.47864 | 4.41 | 4.41 | 0 | 0 | 11.71 | [14 empty] | 4.41
[6 empty] | 0.89059 | 4.42 | 11.7567 | 471.1672 | 207.5218 | 14 | 11.7567 | 8.7317 | 0 | 0.48134 | 4.42 | 4.42 | 0 | 0 | 11.7567 | [14 empty] | 4.42
[6 empty] | 0.90059 | 4.43 | 11.8033 | 473.5005 | 209.8551 | 14 | 11.8033 | 8.7386 | 0 | 0.48404 | 4.43 | 4.43 | 0 | 0 | 11.8033 | [14 empty] | 4.43
[6 empty] | 0.91059 | 4.44 | 11.85 | 475.8338 | 212.1884 | 14 | 11.85 | 8.7456 | 0 | 0.48675 | 4.44 | 4.44 | 0 | 0 | 11.85 | [14 empty] | 4.44
[6 empty] | 0.92059 | 4.45 | 11.8967 | 478.1672 | 214.5218 | 14 | 11.8967 | 8.7526 | 0 | 0.48947 | 4.45 | 4.45 | 0 | 0 | 11.8967 | [14 empty] | 4.45
[6 empty] | 0.93059 | 4.46 | 11.9433 | 480.5005 | 216.8551 | 14 | 11.9433 | 8.7598 | 0 | 0.49221 | 4.46 | 4.46 | 0 | 0 | 11.9433 | [14 empty] | 4.46
[6 empty] | 0.94059 | 4.47 | 11.99 | 482.8338 | 219.1884 | 14 | 11.99 | 8.767 | 0 | 0.49495 | 4.47 | 4.47 | 0 | 0 | 11.99 | [14 empty] | 4.47
[6 empty] | 0.95059 | 4.48 | 12.0367 | 485.1672 | 221.5218 | 14 | 12.0367 | 8.7743 | 0 | 0.49771 | 4.48 | 4.48 | 0 | 0 | 12.0367 | [14 empty] | 4.48
[6 empty] | 0.96059 | 4.49 | 12.0833 | 487.5005 | 223.8551 | 14 | 12.0833 | 8.7816 | 0 | 0.50047 | 4.49 | 4.49 | 0 | 0 | 12.0833 | [14 empty] | 4.49
[6 empty] | 0.97059 | 4.5 | 12.13 | 489.8338 | 226.1884 | 14 | 12.13 | 8.7891 | 0 | 0.50325 | 4.5 | 4.5 | 0 | 0 | 12.13 | [14 empty] | 4.5
[6 empty] | 0.98059 | 4.51 | 12.1767 | 492.1672 | 228.5218 | 14 | 12.1767 | 8.7966 | 0 | 0.50604 | 4.51 | 4.51 | 0 | 0 | 12.1767 | [14 empty] | 4.51
[6 empty] | 0.99059 | 4.52 | 12.2233 | 494.5005 | 230.8551 | 14 | 12.2233 | 8.8041 | 0 | 0.50883 | 4.52 | 4.52 | 0 | 0 | 12.2233 | [14 empty] | 4.52
[6 empty] | 1.0006 | 4.53 | 12.27 | 496.8338 | 233.1884 | 14 | 12.27 | 8.8117 | 0 | 0.51164 | 4.53 | 4.53 | 0 | 0 | 12.27 | [14 empty] | 4.53
[6 empty] | 1.0106 | 4.54 | 12.3167 | 499.1672 | 235.5218 | 14 | 12.3167 | 8.8195 | 0 | 0.51446 | 4.54 | 4.54 | 0 | 0 | 12.3167 | [14 empty] | 4.54
[6 empty] | 1.0206 | 4.55 | 12.3633 | 501.5005 | 237.8551 | 14 | 12.3633 | 8.8272 | 0 | 0.51729 | 4.55 | 4.55 | 0 | 0 | 12.3633 | [14 empty] | 4.55
[6 empty] | 1.0306 | 4.56 | 12.41 | 503.8338 | 240.1884 | 14 | 12.41 | 8.8351 | 0 | 0.52013 | 4.56 | 4.56 | 0 | 0 | 12.41 | [14 empty] | 4.56
[6 empty] | 1.0406 | 4.57 | 12.4567 | 506.1672 | 242.5218 | 14 | 12.4567 | 8.843 | 0 | 0.52298 | 4.57 | 4.57 | 0 | 0 | 12.4567 | [14 empty] | 4.57
[6 empty] | 1.0506 | 4.58 | 12.5033 | 508.5005 | 244.8551 | 14 | 12.5033 | 8.8509 | 0 | 0.52585 | 4.58 | 4.58 | 0 | 0 | 12.5033 | [14 empty] | 4.58
[6 empty] | 1.0606 | 4.59 | 12.55 | 510.8338 | 247.1884 | 14 | 12.55 | 8.859 | 0 | 0.52872 | 4.59 | 4.59 | 0 | 0 | 12.55 | [14 empty] | 4.59
[6 empty] | 1.0706 | 4.6 | 12.5967 | 513.1672 | 249.5218 | 14 | 12.5967 | 8.8671 | 0 | 0.5316 | 4.6 | 4.6 | 0 | 0 | 12.5967 | [14 empty] | 4.6
[6 empty] | 1.0806 | 4.61 | 12.6433 | 515.5005 | 251.8551 | 14 | 12.6433 | 8.8753 | 0 | 0.5345 | 4.61 | 4.61 | 0 | 0 | 12.6433 | [14 empty] | 4.61
[6 empty] | 1.0906 | 4.62 | 12.69 | 517.8338 | 254.1884 | 14 | 12.69 | 8.8835 | 0 | 0.5374 | 4.62 | 4.62 | 0 | 0 | 12.69 | [14 empty] | 4.62
[6 empty] | 1.1006 | 4.63 | 12.7367 | 520.1672 | 256.5218 | 14 | 12.7367 | 8.8918 | 0 | 0.54032 | 4.63 | 4.63 | 0 | 0 | 12.7367 | [14 empty] | 4.63
[6 empty] | 1.1106 | 4.64 | 12.7833 | 522.5005 | 258.8551 | 14 | 12.7833 | 8.9002 | 0 | 0.54324 | 4.64 | 4.64 | 0 | 0 | 12.7833 | [14 empty] | 4.64
[6 empty] | 1.1206 | 4.65 | 12.83 | 524.8338 | 261.1884 | 14 | 12.83 | 8.9086 | 0 | 0.54618 | 4.65 | 4.65 | 0 | 0 | 12.83 | [14 empty] | 4.65
[6 empty] | 1.1306 | 4.66 | 12.8767 | 527.1672 | 263.5218 | 14 | 12.8767 | 8.9171 | 0 | 0.54913 | 4.66 | 4.66 | 0 | 0 | 12.8767 | [14 empty] | 4.66
[6 empty] | 1.1406 | 4.67 | 12.9233 | 529.5005 | 265.8551 | 14 | 12.9233 | 8.9257 | 0 | 0.55208 | 4.67 | 4.67 | 0 | 0 | 12.9233 | [14 empty] | 4.67
[6 empty] | 1.1506 | 4.68 | 12.97 | 531.8338 | 268.1884 | 14 | 12.97 | 8.9343 | 0 | 0.55505 | 4.68 | 4.68 | 0 | 0 | 12.97 | [14 empty] | 4.68
[6 empty] | 1.1606 | 4.69 | 13.0167 | 534.1672 | 270.5218 | 14 | 13.0167 | 8.943 | 0 | 0.55803 | 4.69 | 4.69 | 0 | 0 | 13.0167 | [14 empty] | 4.69
[6 empty] | 1.1706 | 4.7 | 13.0633 | 536.5005 | 272.8551 | 14 | 13.0633 | 8.9517 | 0 | 0.56102 | 4.7 | 4.7 | 0 | 0 | 13.0633 | [14 empty] | 4.7
[6 empty] | 1.1806 | 4.71 | 13.11 | 538.8338 | 275.1884 | 14 | 13.11 | 8.9605 | 0 | 0.56402 | 4.71 | 4.71 | 0 | 0 | 13.11 | [14 empty] | 4.71
[6 empty] | 1.1906 | 4.72 | 13.1567 | 541.1672 | 277.5218 | 14 | 13.1567 | 8.9694 | 0 | 0.56704 | 4.72 | 4.72 | 0 | 0 | 13.1567 | [14 empty] | 4.72
[6 empty] | 1.2006 | 4.73 | 13.2033 | 543.5005 | 279.8551 | 14 | 13.2033 | 8.9783 | 0 | 0.57006 | 4.73 | 4.73 | 0 | 0 | 13.2033 | [14 empty] | 4.73
[6 empty] | 1.2106 | 4.74 | 13.25 | 545.8338 | 282.1884 | 14 | 13.25 | 8.9873 | 0 | 0.57309 | 4.74 | 4.74 | 0 | 0 | 13.25 | [14 empty] | 4.74
[6 empty] | 1.2206 | 4.75 | 13.2967 | 548.1672 | 284.5218 | 14 | 13.2967 | 8.9964 | 0 | 0.57613 | 4.75 | 4.75 | 0 | 0 | 13.2967 | [14 empty] | 4.75
[6 empty] | 1.2306 | 4.76 | 13.3433 | 550.5005 | 286.8551 | 14 | 13.3433 | 9.0055 | 0 | 0.57919 | 4.76 | 4.76 | 0 | 0 | 13.3433 | [14 empty] | 4.76
[6 empty] | 1.2406 | 4.77 | 13.39 | 552.8338 | 289.1884 | 14 | 13.39 | 9.0147 | 0 | 0.58225 | 4.77 | 4.77 | 0 | 0 | 13.39 | [14 empty] | 4.77
[6 empty] | 1.2506 | 4.78 | 13.4367 | 555.1672 | 291.5218 | 14 | 13.4367 | 9.0239 | 0 | 0.58533 | 4.78 | 4.78 | 0 | 0 | 13.4367 | [14 empty] | 4.78
[6 empty] | 1.2606 | 4.79 | 13.4833 | 557.5005 | 293.8551 | 14 | 13.4833 | 9.0332 | 0 | 0.58842 | 4.79 | 4.79 | 0 | 0 | 13.4833 | [14 empty] | 4.79
[6 empty] | 1.2706 | 4.8 | 13.53 | 559.8338 | 296.1884 | 14 | 13.53 | 9.0425 | 0 | 0.59151 | 4.8 | 4.8 | 0 | 0 | 13.53 | [14 empty] | 4.8
[6 empty] | 1.2806 | 4.81 | 13.5767 | 562.1672 | 298.5218 | 14 | 13.5767 | 9.052 | 0 | 0.59462 | 4.81 | 4.81 | 0 | 0 | 13.5767 | [14 empty] | 4.81
[6 empty] | 1.2906 | 4.82 | 13.6233 | 564.5005 | 300.8551 | 14 | 13.6233 | 9.0614 | 0 | 0.59774 | 4.82 | 4.82 | 0 | 0 | 13.6233 | [14 empty] | 4.82
[6 empty] | 1.3006 | 4.83 | 13.67 | 566.8338 | 303.1884 | 14 | 13.67 | 9.0709 | 0 | 0.60087 | 4.83 | 4.83 | 0 | 0 | 13.67 | [14 empty] | 4.83
[6 empty] | 1.3106 | 4.84 | 13.7167 | 569.1672 | 305.5218 | 14 | 13.7167 | 9.0805 | 0 | 0.60401 | 4.84 | 4.84 | 0 | 0 | 13.7167 | [14 empty] | 4.84
[6 empty] | 1.3206 | 4.85 | 13.7633 | 571.5005 | 307.8551 | 14 | 13.7633 | 9.0902 | 0 | 0.60716 | 4.85 | 4.85 | 0 | 0 | 13.7633 | [14 empty] | 4.85
[6 empty] | 1.3306 | 4.86 | 13.81 | 573.8338 | 310.1884 | 14 | 13.81 | 9.0998 | 0 | 0.61032 | 4.86 | 4.86 | 0 | 0 | 13.81 | [14 empty] | 4.86
[6 empty] | 1.3406 | 4.87 | 13.8567 | 576.1672 | 312.5218 | 14 | 13.8567 | 9.1096 | 0 | 0.61349 | 4.87 | 4.87 | 0 | 0 | 13.8567 | [14 empty] | 4.87
[6 empty] | 1.3506 | 4.88 | 13.9033 | 578.5005 | 314.8551 | 14 | 13.9033 | 9.1194 | 0 | 0.61667 | 4.88 | 4.88 | 0 | 0 | 13.9033 | [14 empty] | 4.88
[6 empty] | 1.3606 | 4.89 | 13.95 | 580.8338 | 317.1884 | 14 | 13.95 | 9.1293 | 0 | 0.61987 | 4.89 | 4.89 | 0 | 0 | 13.95 | [14 empty] | 4.89
[6 empty] | 1.3706 | 4.9 | 13.9967 | 583.1672 | 319.5218 | 14 | 13.9967 | 9.1392 | 0 | 0.62307 | 4.9 | 4.9 | 0 | 0 | 13.9967 | [14 empty] | 4.9
[6 empty] | 1.3806 | 4.91 | 14.0433 | 585.5005 | 321.8551 | 14 | 14.0433 | 9.1491 | 0 | 0.62629 | 4.91 | 4.91 | 0 | 0 | 14.0433 | [14 empty] | 4.91
[6 empty] | 1.3906 | 4.92 | 14.09 | 587.8338 | 324.1884 | 14 | 14.09 | 9.1592 | 0 | 0.62951 | 4.92 | 4.92 | 0 | 0 | 14.09 | [14 empty] | 4.92
[6 empty] | 1.4006 | 4.93 | 14.1367 | 590.1672 | 326.5218 | 14 | 14.1367 | 9.1692 | 0 | 0.63275 | 4.93 | 4.93 | 0 | 0 | 14.1367 | [14 empty] | 4.93
[6 empty] | 1.4106 | 4.94 | 14.1833 | 592.5005 | 328.8551 | 14 | 14.1833 | 9.1794 | 0 | 0.63599 | 4.94 | 4.94 | 0 | 0 | 14.1833 | [14 empty] | 4.94
[6 empty] | 1.4206 | 4.95 | 14.23 | 594.8338 | 331.1884 | 14 | 14.23 | 9.1895 | 0 | 0.63925 | 4.95 | 4.95 | 0 | 0 | 14.23 | [14 empty] | 4.95
[6 empty] | 1.4306 | 4.96 | 14.2767 | 597.1672 | 333.5218 | 14 | 14.2767 | 9.1998 | 0 | 0.64252 | 4.96 | 4.96 | 0 | 0 | 14.2767 | [14 empty] | 4.96
[6 empty] | 1.4406 | 4.97 | 14.3233 | 599.5005 | 335.8551 | 14 | 14.3233 | 9.2101 | 0 | 0.6458 | 4.97 | 4.97 | 0 | 0 | 14.3233 | [14 empty] | 4.97
[6 empty] | 1.4506 | 4.98 | 14.37 | 601.8338 | 338.1884 | 14 | 14.37 | 9.2204 | 0 | 0.64909 | 4.98 | 4.98 | 0 | 0 | 14.37 | [14 empty] | 4.98
[6 empty] | 1.4606 | 4.99 | 14.4167 | 604.1672 | 340.5218 | 14 | 14.4167 | 9.2308 | 0 | 0.65239 | 4.99 | 4.99 | 0 | 0 | 14.4167 | [14 empty] | 4.99
[6 empty] | 1.4706 | 5 | 14.4633 | 606.5005 | 342.8551 | 14 | 14.4633 | 9.2412 | 0 | 0.6557 | 5 | 5 | 0 | 0 | 14.4633 | [14 empty] | 5
[6 empty] | 1.4806 | 5.01 | 14.51 | 608.8338 | 345.1884 | 14 | 14.51 | 9.2517 | 0 | 0.65902 | 5.01 | 5.01 | 0 | 0 | 14.51 | [14 empty] | 5.01
[6 empty] | 1.4906 | 5.02 | 14.5567 | 611.1672 | 347.5218 | 14 | 14.5567 | 9.2623 | 0 | 0.66235 | 5.02 | 5.02 | 0 | 0 | 14.5567 | [14 empty] | 5.02
[6 empty] | 1.5006 | 5.03 | 14.6033 | 613.5005 | 349.8551 | 14 | 14.6033 | 9.2729 | 0 | 0.66569 | 5.03 | 5.03 | 0 | 0 | 14.6033 | [14 empty] | 5.03
[6 empty] | 1.5106 | 5.04 | 14.65 | 615.8338 | 352.1884 | 14 | 14.65 | 9.2835 | 0 | 0.66905 | 5.04 | 5.04 | 0 | 0 | 14.65 | [14 empty] | 5.04
[6 empty] | 1.5206 | 5.05 | 14.6967 | 618.1672 | 354.5218 | 14 | 14.6967 | 9.2942 | 0 | 0.67241 | 5.05 | 5.05 | 0 | 0 | 14.6967 | [14 empty] | 5.05
[6 empty] | 1.5306 | 5.06 | 14.7433 | 620.5005 | 356.8551 | 14 | 14.7433 | 9.305 | 0 | 0.67578 | 5.06 | 5.06 | 0 | 0 | 14.7433 | [14 empty] | 5.06
[6 empty] | 1.5406 | 5.07 | 14.79 | 622.8338 | 359.1884 | 14 | 14.79 | 9.3158 | 0 | 0.67917 | 5.07 | 5.07 | 0 | 0 | 14.79 | [14 empty] | 5.07
[6 empty] | 1.5506 | 5.08 | 14.8367 | 625.1672 | 361.5218 | 14 | 14.8367 | 9.3266 | 0 | 0.68257 | 5.08 | 5.08 | 0 | 0 | 14.8367 | [14 empty] | 5.08
[6 empty] | 1.5606 | 5.09 | 14.8833 | 627.5005 | 363.8551 | 14 | 14.8833 | 9.3375 | 0 | 0.68597 | 5.09 | 5.09 | 0 | 0 | 14.8833 | [14 empty] | 5.09
[6 empty] | 1.5706 | 5.1 | 14.93 | 629.8338 | 366.1884 | 14 | 14.93 | 9.3485 | 0 | 0.68939 | 5.1 | 5.1 | 0 | 0 | 14.93 | [14 empty] | 5.1
[6 empty] | 1.5806 | 5.11 | 14.9767 | 632.1672 | 368.5218 | 14 | 14.9767 | 9.3595 | 0 | 0.69282 | 5.11 | 5.11 | 0 | 0 | 14.9767 | [14 empty] | 5.11
[6 empty] | 1.5906 | 5.12 | 15.0233 | 634.5005 | 370.8551 | 14 | 15.0233 | 9.3705 | 0 | 0.69626 | 5.12 | 5.12 | 0 | 0 | 15.0233 | [14 empty] | 5.12
[6 empty] | 1.6006 | 5.13 | 15.07 | 636.8338 | 373.1884 | 14 | 15.07 | 9.3816 | 0 | 0.69971 | 5.13 | 5.13 | 0 | 0 | 15.07 | [14 empty] | 5.13
[6 empty] | 1.6106 | 5.14 | 15.1167 | 639.1672 | 375.5218 | 14 | 15.1167 | 9.3927 | 0 | 0.70317 | 5.14 | 5.14 | 0 | 0 | 15.1167 | [14 empty] | 5.14
[6 empty] | 1.6206 | 5.15 | 15.1633 | 641.5005 | 377.8551 | 14 | 15.1633 | 9.4039 | 0 | 0.70664 | 5.15 | 5.15 | 0 | 0 | 15.1633 | [14 empty] | 5.15
[6 empty] | 1.6306 | 5.16 | 15.21 | 643.8338 | 380.1884 | 14 | 15.21 | 9.4151 | 0 | 0.71012 | 5.16 | 5.16 | 0 | 0 | 15.21 | [14 empty] | 5.16
[6 empty] | 1.6406 | 5.17 | 15.2567 | 646.1672 | 382.5218 | 14 | 15.2567 | 9.4264 | 0 | 0.71361 | 5.17 | 5.17 | 0 | 0 | 15.2567 | [14 empty] | 5.17
[6 empty] | 1.6506 | 5.18 | 15.3033 | 648.5005 | 384.8551 | 14 | 15.3033 | 9.4377 | 0 | 0.71711 | 5.18 | 5.18 | 0 | 0 | 15.3033 | [14 empty] | 5.18
[6 empty] | 1.6606 | 5.19 | 15.35 | 650.8338 | 387.1884 | 14 | 15.35 | 9.4491 | 0 | 0.72063 | 5.19 | 5.19 | 0 | 0 | 15.35 | [14 empty] | 5.19
[6 empty] | 1.6706 | 5.2 | 15.3967 | 653.1672 | 389.5218 | 14 | 15.3967 | 9.4605 | 0 | 0.72415 | 5.2 | 5.2 | 0 | 0 | 15.3967 | [14 empty] | 5.2
[6 empty] | 1.6806 | 5.21 | 15.4433 | 655.5005 | 391.8551 | 14 | 15.4433 | 9.472 | 0 | 0.72769 | 5.21 | 5.21 | 0 | 0 | 15.4433 | [14 empty] | 5.21
[6 empty] | 1.6906 | 5.22 | 15.49 | 657.8338 | 394.1884 | 14 | 15.49 | 9.4835 | 0 | 0.73123 | 5.22 | 5.22 | 0 | 0 | 15.49 | [14 empty] | 5.22
[6 empty] | 1.7006 | 5.23 | 15.5367 | 660.1672 | 396.5218 | 14 | 15.5367 | 9.495 | 0 | 0.73479 | 5.23 | 5.23 | 0 | 0 | 15.5367 | [14 empty] | 5.23
[6 empty] | 1.7106 | 5.24 | 15.5833 | 662.5005 | 398.8551 | 14 | 15.5833 | 9.5066 | 0 | 0.73836 | 5.24 | 5.24 | 0 | 0 | 15.5833 | [14 empty] | 5.24
[6 empty] | 1.7206 | 5.25 | 15.63 | 664.8338 | 401.1884 | 14 | 15.63 | 9.5183 | 0 | 0.74193 | 5.25 | 5.25 | 0 | 0 | 15.63 | [14 empty] | 5.25
[6 empty] | 1.7306 | 5.26 | 15.6767 | 667.1672 | 403.5218 | 14 | 15.6767 | 9.53 | 0 | 0.74552 | 5.26 | 5.26 | 0 | 0 | 15.6767 | [14 empty] | 5.26
[6 empty] | 1.7406 | 5.27 | 15.7233 | 669.5005 | 405.8551 | 14 | 15.7233 | 9.5417 | 0 | 0.74912 | 5.27 | 5.27 | 0 | 0 | 15.7233 | [14 empty] | 5.27
[6 empty] | 1.7506 | 5.28 | 15.77 | 671.8338 | 408.1884 | 14 | 15.77 | 9.5535 | 0 | 0.75273 | 5.28 | 5.28 | 0 | 0 | 15.77 | [14 empty] | 5.28
[6 empty] | 1.7606 | 5.29 | 15.8167 | 674.1672 | 410.5218 | 14 | 15.8167 | 9.5653 | 0 | 0.75635 | 5.29 | 5.29 | 0 | 0 | 15.8167 | [14 empty] | 5.29
[6 empty] | 1.7706 | 5.3 | 15.8633 | 676.5005 | 412.8551 | 14 | 15.8633 | 9.5771 | 0 | 0.75998 | 5.3 | 5.3 | 0 | 0 | 15.8633 | [14 empty] | 5.3
[6 empty] | 1.7806 | 5.31 | 15.91 | 678.8338 | 415.1884 | 14 | 15.91 | 9.5891 | 0 | 0.76363 | 5.31 | 5.31 | 0 | 0 | 15.91 | [14 empty] | 5.31
[6 empty] | 1.7906 | 5.32 | 15.9567 | 681.1672 | 417.5218 | 14 | 15.9567 | 9.601 | 0 | 0.76728 | 5.32 | 5.32 | 0 | 0 | 15.9567 | [14 empty] | 5.32
[6 empty] | 1.8006 | 5.33 | 5 | 683.5005 | 414.3534 | -52.0201 | 5 | 9.5924 | 0 | 0.76302 | 5.33 | 5.33 | 0 | 0 | 5 | [14 empty] | 5.33
[6 empty] | 1.8106 | 5.34 | 5 | 674.8305 | 405.7701 | -50.9797 | 5 | 9.5838 | 0 | 0.75886 | 5.34 | 5.34 | 0 | 0 | 5 | [14 empty] | 5.34
[6 empty] | 1.8206 | 5.35 | 5 | 666.3339 | 397.3585 | -49.9601 | 5 | 9.5752 | 0 | 0.75477 | 5.35 | 5.35 | 0 | 0 | 5 | [14 empty] | 5.35
[6 empty] | 1.8306 | 5.36 | 5 | 658.0072 | 389.1151 | -48.9609 | 5 | 9.5667 | 0 | 0.75077 | 5.36 | 5.36 | 0 | 0 | 5 | [14 empty] | 5.36
[6 empty] | 1.8406 | 5.37 | 5 | 649.8471 | 381.0365 | -47.9816 | 5 | 9.5582 | 0 | 0.74685 | 5.37 | 5.37 | 0 | 0 | 5 | [14 empty] | 5.37
[6 empty] | 1.8506 | 5.38 | 5 | 641.8501 | 373.1195 | -47.022 | 5 | 9.5498 | 0 | 0.743 | 5.38 | 5.38 | 0 | 0 | 5 | [14 empty] | 5.38
[6 empty] | 1.8606 | 5.39 | 5 | 634.0131 | 365.3609 | -46.0816 | 5 | 9.5414 | 0 | 0.73924 | 5.39 | 5.39 | 0 | 0 | 5 | [14 empty] | 5.39
[6 empty] | 1.8706 | 5.4 | 5 | 626.3329 | 357.7575 | -45.1599 | 5 | 9.533 | 0 | 0.73555 | 5.4 | 5.4 | 0 | 0 | 5 | [14 empty] | 5.4
[6 empty] | 1.8806 | 5.41 | 5 | 618.8062 | 350.3061 | -44.2567 | 5 | 9.5246 | 0 | 0.73193 | 5.41 | 5.41 | 0 | 0 | 5 | [14 empty] | 5.41
[6 empty] | 1.8906 | 5.42 | 5 | 611.4301 | 343.0037 | -43.3716 | 5 | 9.5163 | 0 | 0.72838 | 5.42 | 5.42 | 0 | 0 | 5 | [14 empty] | 5.42
[6 empty] | 1.9006 | 5.43 | 5 | 604.2015 | 335.8474 | -42.5042 | 5 | 9.508 | 0 | 0.72491 | 5.43 | 5.43 | 0 | 0 | 5 | [14 empty] | 5.43
[6 empty] | 1.9106 | 5.44 | 5 | 597.1175 | 328.8342 | -41.6541 | 5 | 9.4997 | 0 | 0.7215 | 5.44 | 5.44 | 0 | 0 | 5 | [14 empty] | 5.44
[6 empty] | 1.9206 | 5.45 | 5 | 590.1751 | 321.9613 | -40.821 | 5 | 9.4914 | 0 | 0.71816 | 5.45 | 5.45 | 0 | 0 | 5 | [14 empty] | 5.45
[6 empty] | 1.9306 | 5.46 | 5 | 583.3716 | 315.2258 | -40.0046 | 5 | 9.4832 | 0 | 0.71489 | 5.46 | 5.46 | 0 | 0 | 5 | [14 empty] | 5.46
[6 empty] | 1.9406 | 5.47 | 5 | 576.7042 | 308.625 | -39.2045 | 5 | 9.4751 | 0 | 0.71169 | 5.47 | 5.47 | 0 | 0 | 5 | [14 empty] | 5.47
[6 empty] | 1.9506 | 5.48 | 5 | 570.1701 | 302.1563 | -38.4204 | 5 | 9.4669 | 0 | 0.70855 | 5.48 | 5.48 | 0 | 0 | 5 | [14 empty] | 5.48
[6 empty] | 1.9606 | 5.49 | 5 | 563.7667 | 295.8169 | -37.652 | 5 | 9.4588 | 0 | 0.70547 | 5.49 | 5.49 | 0 | 0 | 5 | [14 empty] | 5.49
[6 empty] | 1.9706 | 5.5 | 5 | 557.4914 | 289.6044 | -36.899 | 5 | 9.4507 | 0 | 0.70245 | 5.5 | 5.5 | 0 | 0 | 5 | [14 empty] | 5.5
[6 empty] | 1.9806 | 5.51 | 5 | 551.3415 | 283.516 | -36.161 | 5 | 9.4426 | 0 | 0.6995 | 5.51 | 5.51 | 0 | 0 | 5 | [14 empty] | 5.51
[6 empty] | 1.9906 | 5.52 | 5 | 545.3147 | 277.5495 | -35.4378 | 5 | 9.4346 | 0 | 0.6966 | 5.52 | 5.52 | 0 | 0 | 5 | [14 empty] | 5.52
[6 empty] | 2.0006 | 5.53 | 5 | 539.4084 | 271.7022 | -34.729 | 5 | 9.4266 | 0 | 0.69376 | 5.53 | 5.53 | 0 | 0 | 5 | [14 empty] | 5.53
[6 empty] | 2.0106 | 5.54 | 5 | 533.6202 | 265.9719 | -34.0344 | 5 | 9.4186 | 0 | 0.69098 | 5.54 | 5.54 | 0 | 0 | 5 | [14 empty] | 5.54
[6 empty] | 2.0206 | 5.55 | 5 | 527.9478 | 260.3563 | -33.3537 | 5 | 9.4107 | 0 | 0.68825 | 5.55 | 5.55 | 0 | 0 | 5 | [14 empty] | 5.55
[6 empty] | 2.0306 | 5.56 | 5 | 522.3889 | 254.8529 | -32.6867 | 5 | 9.4027 | 0 | 0.68558 | 5.56 | 5.56 | 0 | 0 | 5 | [14 empty] | 5.56
[6 empty] | 2.0406 | 5.57 | 5 | 516.9411 | 249.4596 | -32.0329 | 5 | 9.3949 | 0 | 0.68296 | 5.57 | 5.57 | 0 | 0 | 5 | [14 empty] | 5.57
[6 empty] | 2.0506 | 5.58 | 5 | 511.6023 | 244.1742 | -31.3923 | 5 | 9.387 | 0 | 0.6804 | 5.58 | 5.58 | 0 | 0 | 5 | [14 empty] | 5.58
[6 empty] | 2.0606 | 5.59 | 5 | 506.3702 | 238.9944 | -30.7644 | 5 | 9.3792 | 0 | 0.67788 | 5.59 | 5.59 | 0 | 0 | 5 | [14 empty] | 5.59
[6 empty] | 2.0706 | 5.6 | 5 | 501.2428 | 233.9183 | -30.1491 | 5 | 9.3714 | 0 | 0.67542 | 5.6 | 5.6 | 0 | 0 | 5 | [14 empty] | 5.6
[6 empty] | 2.0806 | 5.61 | 5 | 496.218 | 228.9437 | -29.5462 | 5 | 9.3636 | 0 | 0.673 | 5.61 | 5.61 | 0 | 0 | 5 | [14 empty] | 5.61
[6 empty] | 2.0906 | 5.62 | 5 | 491.2936 | 224.0686 | -28.9552 | 5 | 9.3558 | 0 | 0.67063 | 5.62 | 5.62 | 0 | 0 | 5 | [14 empty] | 5.62
[6 empty] | 2.1006 | 5.63 | 5 | 486.4677 | 219.291 | -28.3761 | 5 | 9.3481 | 0 | 0.66831 | 5.63 | 5.63 | 0 | 0 | 5 | [14 empty] | 5.63
[6 empty] | 2.1106 | 5.64 | 5 | 481.7384 | 214.6089 | -27.8086 | 5 | 9.3404 | 0 | 0.66604 | 5.64 | 5.64 | 0 | 0 | 5 | [14 empty] | 5.64
[6 empty] | 2.1206 | 5.65 | 5 | 477.1036 | 210.0205 | -27.2524 | 5 | 9.3327 | 0 | 0.66381 | 5.65 | 5.65 | 0 | 0 | 5 | [14 empty] | 5.65
[6 empty] | 2.1306 | 5.66 | 5 | 472.5615 | 205.5238 | -26.7074 | 5 | 9.3251 | 0 | 0.66163 | 5.66 | 5.66 | 0 | 0 | 5 | [14 empty] | 5.66
[6 empty] | 2.1406 | 5.67 | 5 | 468.1103 | 201.1171 | -26.1732 | 5 | 9.3175 | 0 | 0.65949 | 5.67 | 5.67 | 0 | 0 | 5 | [14 empty] | 5.67
[6 empty] | 2.1506 | 5.68 | 5 | 463.7481 | 196.7985 | -25.6498 | 5 | 9.3099 | 0 | 0.65739 | 5.68 | 5.68 | 0 | 0 | 5 | [14 empty] | 5.68
[6 empty] | 2.1606 | 5.69 | 5 | 459.4731 | 192.5663 | -25.1368 | 5 | 9.3023 | 0 | 0.65534 | 5.69 | 5.69 | 0 | 0 | 5 | [14 empty] | 5.69
[6 empty] | 2.1706 | 5.7 | 5 | 455.2837 | 188.4188 | -24.634 | 5 | 9.2948 | 0 | 0.65332 | 5.7 | 5.7 | 0 | 0 | 5 | [14 empty] | 5.7
[6 empty] | 2.1806 | 5.71 | 5 | 451.178 | 184.3541 | -24.1414 | 5 | 9.2873 | 0 | 0.65135 | 5.71 | 5.71 | 0 | 0 | 5 | [14 empty] | 5.71
[6 empty] | 2.1906 | 5.72 | 5 | 447.1544 | 180.3708 | -23.6585 | 5 | 9.2798 | 0 | 0.64942 | 5.72 | 5.72 | 0 | 0 | 5 | [14 empty] | 5.72
[6 empty] | 2.2006 | 5.73 | 5 | 443.2114 | 176.4672 | -23.1854 | 5 | 9.2724 | 0 | 0.64752 | 5.73 | 5.73 | 0 | 0 | 5 | [14 empty] | 5.73
[6 empty] | 2.2106 | 5.74 | 5 | 439.3471 | 172.6416 | -22.7217 | 5 | 9.2649 | 0 | 0.64566 | 5.74 | 5.74 | 0 | 0 | 5 | [14 empty] | 5.74
[6 empty] | 2.2206 | 5.75 | 5 | 435.5602 | 168.8925 | -22.2672 | 5 | 9.2575 | 0 | 0.64384 | 5.75 | 5.75 | 0 | 0 | 5 | [14 empty] | 5.75
[6 empty] | 2.2306 | 5.76 | 5 | 431.849 | 165.2184 | -21.8219 | 5 | 9.2501 | 0 | 0.64206 | 5.76 | 5.76 | 0 | 0 | 5 | [14 empty] | 5.76
[6 empty] | 2.2406 | 5.77 | 5 | 428.212 | 161.6178 | -21.3854 | 5 | 9.2428 | 0 | 0.64031 | 5.77 | 5.77 | 0 | 0 | 5 | [14 empty] | 5.77
[6 empty] | 2.2506 | 5.78 | 5 | 424.6478 | 158.0892 | -20.9577 | 5 | 9.2355 | 0 | 0.6386 | 5.78 | 5.78 | 0 | 0 | 5 | [14 empty] | 5.78
[6 empty] | 2.2606 | 5.79 | 5 | 421.1548 | 154.6312 | -20.5386 | 5 | 9.2282 | 0 | 0.63692 | 5.79 | 5.79 | 0 | 0 | 5 | [14 empty] | 5.79
[6 empty] | 2.2706 | 5.8 | 5 | 417.7317 | 151.2423 | -20.1278 | 5 | 9.2209 | 0 | 0.63527 | 5.8 | 5.8 | 0 | 0 | 5 | [14 empty] | 5.8
[6 empty] | 2.2806 | 5.81 | 5 | 414.3771 | 147.9212 | -19.7252 | 5 | 9.2136 | 0 | 0.63366 | 5.81 | 5.81 | 0 | 0 | 5 | [14 empty] | 5.81
[6 empty] | 2.2906 | 5.82 | 5 | 411.0895 | 144.6666 | -19.3307 | 5 | 9.2064 | 0 | 0.63208 | 5.82 | 5.82 | 0 | 0 | 5 | [14 empty] | 5.82
[6 empty] | 2.3006 | 5.83 | 5 | 407.8677 | 141.477 | -18.9441 | 5 | 9.1992 | 0 | 0.63053 | 5.83 | 5.83 | 0 | 0 | 5 | [14 empty] | 5.83
[6 empty] | 2.3106 | 5.84 | 5 | 404.7104 | 138.3512 | -18.5652 | 5 | 9.192 | 0 | 0.62901 | 5.84 | 5.84 | 0 | 0 | 5 | [14 empty] | 5.84
[6 empty] | 2.3206 | 5.85 | 5 | 401.6162 | 135.2879 | -18.1939 | 5 | 9.1849 | 0 | 0.62753 | 5.85 | 5.85 | 0 | 0 | 5 | [14 empty] | 5.85
[6 empty] | 2.3306 | 5.86 | 5 | 398.5839 | 132.2859 | -17.8301 | 5 | 9.1777 | 0 | 0.62607 | 5.86 | 5.86 | 0 | 0 | 5 | [14 empty] | 5.86
[6 empty] | 2.3406 | 5.87 | 5 | 395.6122 | 129.344 | -17.4735 | 5 | 9.1706 | 0 | 0.62464 | 5.87 | 5.87 | 0 | 0 | 5 | [14 empty] | 5.87
[6 empty] | 2.3506 | 5.88 | 5 | 392.6999 | 126.4609 | -17.124 | 5 | 9.1636 | 0 | 0.62324 | 5.88 | 5.88 | 0 | 0 | 5 | [14 empty] | 5.88
[6 empty] | 2.3606 | 5.89 | 5 | 389.8459 | 123.6354 | -16.7815 | 5 | 9.1565 | 0 | 0.62187 | 5.89 | 5.89 | 0 | 0 | 5 | [14 empty] | 5.89
[6 empty] | 2.3706 | 5.9 | 5 | 387.049 | 120.8664 | -16.4459 | 5 | 9.1495 | 0 | 0.62052 | 5.9 | 5.9 | 0 | 0 | 5 | [14 empty] | 5.9
[6 empty] | 2.3806 | 5.91 | 5 | 384.308 | 118.1529 | -16.117 | 5 | 9.1425 | 0 | 0.61921 | 5.91 | 5.91 | 0 | 0 | 5 | [14 empty] | 5.91
[6 empty] | 2.3906 | 5.92 | 5 | 381.6219 | 115.4936 | -15.7946 | 5 | 9.1355 | 0 | 0.61791 | 5.92 | 5.92 | 0 | 0 | 5 | [14 empty] | 5.92
[6 empty] | 2.4006 | 5.93 | 5 | 378.9894 | 112.8875 | -15.4787 | 5 | 9.1285 | 0 | 0.61665 | 5.93 | 5.93 | 0 | 0 | 5 |  | y: 11.2887 |  | 5.93
[6 empty] | 2.4106 | 5.94 | 5 | 376.4097 | 110.3335 | -15.1692 | 5 | 9.1216 | 0 | 0.61541 | 5.94 | 5.94 | 0 | 0 | 5 |  | y: 11.0333 |  | 5.94
[6 empty] | 2.4206 | 5.95 | 5 | 373.8815 | 107.8306 | -14.8658 | 5 | 9.1147 | 0 | 0.61419 | 5.95 | 5.95 | 0 | 0 | 5 |  | y: 10.7831 |  | 5.95
[6 empty] | 2.4306 | 5.96 | 5 | 371.4038 | 105.3777 | -14.5685 | 5 | 9.1078 | 0 | 0.613 | 5.96 | 5.96 | 0 | 0 | 5 |  | y: 10.5378 |  | 5.96
[6 empty] | 2.4406 | 5.97 | 5 | 368.9758 | 102.9739 | -14.2771 | 5 | 9.1009 | 0 | 0.61184 | 5.97 | 5.97 | 0 | 0 | 5 |  | y: 10.2974 |  | 5.97
[6 empty] | 2.4506 | 5.98 | 5 | 366.5962 | 100.6182 | -13.9915 | 5 | 9.094 | 0 | 0.61069 | 5.98 | 5.98 | 0 | 0 | 5 |  | y: 10.0618 |  | 5.98
[6 empty] | 2.4606 | 5.99 | 5 | 364.2643 | 98.3096 | -13.7117 | 5 | 9.0872 | 0 | 0.60957 | 5.99 | 5.99 | 0 | 0 | 5 |  | y: 9.831 |  | 5.99
[6 empty] | 2.4706 | 6 | 5 | 361.979 | 96.0472 | -13.4375 | 5 | 9.0804 | 0 | 0.60847 | 6 | 6 | 0 | 0 | 5 |  | y: 9.6047 |  | 6
Menu
square
ramp
Passive Pressure -Volume Curve
Blue = Static Green = Dynamic

## VC (ramp)
Volume Control
Lung Model
|  | Elastance (cm H2O/L) | 20
|  | Compliance (L/cm H2O) | 0.05 |  | 50
|  | Resistance (cm H2O/L/s) | 10 |  | 10
|  | Time Constant (s) | 0.5
|  | Tidal Lung Pressure (cm H2O) | 8.4
Effort Model
|  | Delay Before Effort (s) | 0 |  | 0
|  | Duration of Effort (s) | 0.25 |  | 5
|  | Freq (hz) | 1
|  | Period (s) | 0.5
|  | max Pmus (cm H2O) | 0 | 49 | 0
Ventilator Settings
Flow Waveform
|  | Peak Flow (L/min) | 41 |  | 41
|  | Peak Flow (L/s) | 0.68333
|  | Tidal Volume (mL) | 420 |  | 420
|  | Tidal Volume (L) | 0.42
|  | Frequency (breaths/min) | 25 |  | 25
|  | PEEP (cm H2O) | 5 |  | 5
|  | autoPEEP (cm H2O) | 0.83806
|  | Pause Time (s) | 0 |  | 0
|  | Peak Airway Pressure (cm H2O) | 15.7057
|  | Total Cycle Time (s) | 2.4
|  | Inspirtory Time (s) | 1.2293
|  | Expiratory Time (s) | 1.2
|  | I:E | 1:1
|  | Duty Cycle (%) | 51.2195
|  | Plateau Pressure (cm H2O) | 14.2381
|  | Tidal (Driving) Pressure (cm H2O) | 8.4
|  | Peak Inspiratory Flow (L/min) | 41
|  | Peak Expiratory Flow (L/min) | -55.8386
|  | Mean Inspiratory Pressure (cm H2O) | 14.8818
|  | Mean Airway Pressure (cm H2O) | 10.0754
|  | Inspiratory Flow Time (s) | 1.2293
Passive |  | Ventilator Elastic Work (J) | 0.41359
|  | Ventilator Resistive Work (J) | 0.18763
|  | Total Work (J) | 0.60122
Active |  | Ventilator Total Work (J) | 0.60647
|  | Patient Inspiratory Work (J) | 0
|  | Work Shift Index (%) | 0
|  | Ventilator Inspiratory Power (W) | 0.25051
|  | Ventilator Inspiratory Power (J/min) | 15.0305
[19 empty] | Peak | Elastance
[19 empty] | 15.7057 | 0.02 | …………………….Pressure Volume Curve Data……………….
[12 empty] | Airway | Mean |  |  |  |  | Muscle | total | [5 empty] | Seg |  | Seg |  | Seg |  | Seg |  | Muscle
[8 empty] | Pressure | Lung Volume | Volume | Flow | Press | Press | Pat. | Vent | Absolute | Relative | Pressure | Pressure |  | Pstat |  | Pdyn |  | L1 |  | L2 |  | U1 |  | U2 |  | flow
[6 empty] | BreathTime | Time | (cm H2O) | (mL) | (mL) | (L/m) | (cm H2O) | (cm H2O) | Work [J] | Work [J] | Time | Time | (cm H2O) | (cm H2O) |  | X | Y | X | Y | x | y | x | y | x | y | x | y | (L/s) | Time
[6 empty] | 0 | 0 | 12.6714 | 291.9029 | 0 | 41 | 12.6714 | 12.6714 | 0 | 0.0084943 | 0 | 0 | 0 | 12.6714 |  | 5 | 0 | 12.6714 | 0 | 0 | 0 | 5 | 0 | 0 | 420 |  |  | 0 | 0
[6 empty] | 0.01 | 0.01 | 12.7525 | 298.7362 | 6.8055 | 40.6665 | 12.7525 | 12.7119 | 0 | 0.016973 | 0.01 | 0.01 | 0 | 12.7525 |  | 13.4 | 420 | 13.4 | 420 | 12.6714 | 0 | 12.6714 | 0 | 13.4 | 420 |  |  | 0 | 0.01
[6 empty] | 0.02 | 0.02 | 12.8324 | 305.514 | 13.5555 | 40.3329 | 12.8324 | 12.7521 | 0 | 0.025436 | 0.02 | 0.02 | 0 | 12.8324 | [13 empty] | 0 | 0.02
[6 empty] | 0.03 | 0.03 | 12.9113 | 312.2361 | 20.2499 | 39.9994 | 12.9113 | 12.7919 | 0 | 0.033879 | 0.03 | 0.03 | 0 | 12.9113 | [13 empty] | 0 | 0.03
[6 empty] | 0.04 | 0.04 | 12.989 | 318.9027 | 26.8886 | 39.6659 | 12.989 | 12.8313 | 0 | 0.042303 | 0.04 | 0.04 | 0 | 12.989 | [13 empty] | 0 | 0.04
[6 empty] | 0.05 | 0.05 | 13.0657 | 325.5137 | 33.4718 | 39.3323 | 13.0657 | 12.8704 | 0 | 0.050706 | 0.05 | 0.05 | 0 | 13.0657 | [13 empty] | 0 | 0.05
[6 empty] | 0.06 | 0.06 | 13.1412 | 332.0691 | 39.9994 | 38.9988 | 13.1412 | 12.9091 | 0 | 0.059085 | 0.06 | 0.06 | 0 | 13.1412 | [13 empty] | 0 | 0.06
[6 empty] | 0.07 | 0.07 | 13.2156 | 338.5689 | 46.4714 | 38.6653 | 13.2156 | 12.9474 | 0 | 0.067439 | 0.07 | 0.07 | 0 | 13.2156 | [13 empty] | 0 | 0.07
[6 empty] | 0.08 | 0.08 | 13.2889 | 345.0131 | 52.8878 | 38.3317 | 13.2889 | 12.9853 | 0 | 0.075768 | 0.08 | 0.08 | 0 | 13.2889 | [13 empty] | 0 | 0.08
[6 empty] | 0.09 | 0.09 | 13.3611 | 351.4017 | 59.2487 | 37.9982 | 13.3611 | 13.0229 | 0 | 0.084069 | 0.09 | 0.09 | 0 | 13.3611 | [13 empty] | 0 | 0.09
[6 empty] | 0.1 | 0.1 | 13.4321 | 357.7347 | 65.5539 | 37.6647 | 13.4321 | 13.0601 | 0 | 0.092341 | 0.1 | 0.1 | 0 | 13.4321 | [13 empty] | 0 | 0.1
[6 empty] | 0.11 | 0.11 | 13.5021 | 364.0122 | 71.8036 | 37.3312 | 13.5021 | 13.0969 | 0 | 0.10058 | 0.11 | 0.11 | 0 | 13.5021 | [13 empty] | 0 | 0.11
[6 empty] | 0.12 | 0.12 | 13.571 | 370.234 | 77.9976 | 36.9976 | 13.571 | 13.1334 | 0 | 0.10879 | 0.12 | 0.12 | 0 | 13.571 | [13 empty] | 0 | 0.12
[6 empty] | 0.13 | 0.13 | 13.6387 | 376.4003 | 84.1361 | 36.6641 | 13.6387 | 13.1695 | 0 | 0.11697 | 0.13 | 0.13 | 0 | 13.6387 | [13 empty] | 0 | 0.13
[6 empty] | 0.14 | 0.14 | 13.7053 | 382.511 | 90.219 | 36.3306 | 13.7053 | 13.2052 | 0 | 0.12511 | 0.14 | 0.14 | 0 | 13.7053 | [13 empty] | 0 | 0.14
[6 empty] | 0.15 | 0.15 | 13.7708 | 388.5661 | 96.2463 | 35.997 | 13.7708 | 13.2406 | 0 | 0.13321 | 0.15 | 0.15 | 0 | 13.7708 | [13 empty] | 0 | 0.15
[6 empty] | 0.16 | 0.16 | 13.8352 | 394.5656 | 102.218 | 35.6635 | 13.8352 | 13.2755 | 0 | 0.14128 | 0.16 | 0.16 | 0 | 13.8352 | [13 empty] | 0 | 0.16
[6 empty] | 0.17 | 0.17 | 13.8985 | 400.5095 | 108.1341 | 35.33 | 13.8985 | 13.3102 | 0 | 0.14931 | 0.17 | 0.17 | 0 | 13.8985 | [13 empty] | 0 | 0.17
[6 empty] | 0.18 | 0.18 | 13.9607 | 406.3978 | 113.9946 | 34.9964 | 13.9607 | 13.3444 | 0 | 0.1573 | 0.18 | 0.18 | 0 | 13.9607 | [13 empty] | 0 | 0.18
[6 empty] | 0.19 | 0.19 | 14.0218 | 412.2306 | 119.7996 | 34.6629 | 14.0218 | 13.3783 | 0 | 0.16524 | 0.19 | 0.19 | 0 | 14.0218 | [13 empty] | 0 | 0.19
[6 empty] | 0.2 | 0.2 | 14.0817 | 418.0077 | 125.5489 | 34.3294 | 14.0817 | 13.4118 | 0 | 0.17315 | 0.2 | 0.2 | 0 | 14.0817 | [13 empty] | 0 | 0.2
[6 empty] | 0.21 | 0.21 | 14.1406 | 423.7293 | 131.2427 | 33.9958 | 14.1406 | 13.4449 | 0 | 0.18101 | 0.21 | 0.21 | 0 | 14.1406 | [13 empty] | 0 | 0.21
[6 empty] | 0.22 | 0.22 | 14.1983 | 429.3952 | 136.8809 | 33.6623 | 14.1983 | 13.4776 | 0 | 0.18882 | 0.22 | 0.22 | 0 | 14.1983 | [13 empty] | 0 | 0.22
[6 empty] | 0.23 | 0.23 | 14.2549 | 435.0056 | 142.4635 | 33.3288 | 14.2549 | 13.51 | 0 | 0.19659 | 0.23 | 0.23 | 0 | 14.2549 | [13 empty] | 0 | 0.23
[6 empty] | 0.24 | 0.24 | 14.3104 | 440.5604 | 147.9905 | 32.9952 | 14.3104 | 13.542 | 0 | 0.20431 | 0.24 | 0.24 | 0 | 14.3104 | [13 empty] | 0 | 0.24
[6 empty] | 0.25 | 0.25 | 14.3648 | 446.0596 | 153.4619 | 32.6617 | 14.3648 | 13.5737 | 0 | 0.21198 | 0.25 | 0.25 | 0 | 14.3648 | [13 empty] | 0 | 0.25
[6 empty] | 0.26 | 0.26 | 14.4181 | 451.5032 | 158.8777 | 32.3282 | 14.4181 | 13.605 | 0 | 0.2196 | 0.26 | 0.26 | 0 | 14.4181 | [13 empty] | 0 | 0.26
[6 empty] | 0.27 | 0.27 | 14.4703 | 456.8913 | 164.2379 | 31.9946 | 14.4703 | 13.6359 | 0 | 0.22717 | 0.27 | 0.27 | 0 | 14.4703 | [13 empty] | 0 | 0.27
[6 empty] | 0.28 | 0.28 | 14.5213 | 462.2237 | 169.5426 | 31.6611 | 14.5213 | 13.6664 | 0 | 0.23469 | 0.28 | 0.28 | 0 | 14.5213 | [13 empty] | 0 | 0.28
[6 empty] | 0.29 | 0.29 | 14.5713 | 467.5006 | 174.7917 | 31.3276 | 14.5713 | 13.6966 | 0 | 0.24215 | 0.29 | 0.29 | 0 | 14.5713 | [13 empty] | 0 | 0.29
[6 empty] | 0.3 | 0.3 | 14.6201 | 472.7218 | 179.9851 | 30.994 | 14.6201 | 13.7264 | 0 | 0.24956 | 0.3 | 0.3 | 0 | 14.6201 | [13 empty] | 0 | 0.3
[6 empty] | 0.31 | 0.31 | 14.6678 | 477.8875 | 185.123 | 30.6605 | 14.6678 | 13.7558 | 0 | 0.25691 | 0.31 | 0.31 | 0 | 14.6678 | [13 empty] | 0 | 0.31
[6 empty] | 0.32 | 0.32 | 14.7144 | 482.9976 | 190.2053 | 30.327 | 14.7144 | 13.7848 | 0 | 0.26421 | 0.32 | 0.32 | 0 | 14.7144 | [13 empty] | 0 | 0.32
[6 empty] | 0.33 | 0.33 | 14.76 | 488.0521 | 195.232 | 29.9935 | 14.76 | 13.8135 | 0 | 0.27145 | 0.33 | 0.33 | 0 | 14.76 | [13 empty] | 0 | 0.33
[6 empty] | 0.34 | 0.34 | 14.8043 | 493.051 | 200.2031 | 29.6599 | 14.8043 | 13.8418 | 0 | 0.27863 | 0.34 | 0.34 | 0 | 14.8043 | [13 empty] | 0 | 0.34
[6 empty] | 0.35 | 0.35 | 14.8476 | 497.9943 | 205.1186 | 29.3264 | 14.8476 | 13.8698 | 0 | 0.28575 | 0.35 | 0.35 | 0 | 14.8476 | [13 empty] | 0 | 0.35
[6 empty] | 0.36 | 0.36 | 14.8898 | 502.8821 | 209.9786 | 28.9929 | 14.8898 | 13.8973 | 0 | 0.2928 | 0.36 | 0.36 | 0 | 14.8898 | [13 empty] | 0 | 0.36
[6 empty] | 0.37 | 0.37 | 14.9308 | 507.7142 | 214.7829 | 28.6593 | 14.9308 | 13.9245 | 0 | 0.2998 | 0.37 | 0.37 | 0 | 14.9308 | [13 empty] | 0 | 0.37
[6 empty] | 0.38 | 0.38 | 14.9708 | 512.4907 | 219.5317 | 28.3258 | 14.9708 | 13.9513 | 0 | 0.30673 | 0.38 | 0.38 | 0 | 14.9708 | [13 empty] | 0 | 0.38
[6 empty] | 0.39 | 0.39 | 15.0096 | 517.2117 | 224.2249 | 27.9923 | 15.0096 | 13.9778 | 0 | 0.3136 | 0.39 | 0.39 | 0 | 15.0096 | [13 empty] | 0 | 0.39
[6 empty] | 0.4 | 0.4 | 15.0473 | 521.8771 | 228.8624 | 27.6587 | 15.0473 | 14.0039 | 0 | 0.32041 | 0.4 | 0.4 | 0 | 15.0473 | [13 empty] | 0 | 0.4
[6 empty] | 0.41 | 0.41 | 15.0839 | 526.4869 | 233.4444 | 27.3252 | 15.0839 | 14.0296 | 0 | 0.32715 | 0.41 | 0.41 | 0 | 15.0839 | [13 empty] | 0 | 0.41
[6 empty] | 0.42 | 0.42 | 15.1194 | 531.0411 | 237.9708 | 26.9917 | 15.1194 | 14.055 | 0 | 0.33382 | 0.42 | 0.42 | 0 | 15.1194 | [13 empty] | 0 | 0.42
[6 empty] | 0.43 | 0.43 | 15.1538 | 535.5397 | 242.4417 | 26.6581 | 15.1538 | 14.0799 | 0 | 0.34042 | 0.43 | 0.43 | 0 | 15.1538 | [13 empty] | 0 | 0.43
[6 empty] | 0.44 | 0.44 | 15.1871 | 539.9827 | 246.8569 | 26.3246 | 15.1871 | 14.1045 | 0 | 0.34696 | 0.44 | 0.44 | 0 | 15.1871 | [13 empty] | 0 | 0.44
[6 empty] | 0.45 | 0.45 | 15.2192 | 544.3701 | 251.2165 | 25.9911 | 15.2192 | 14.1288 | 0 | 0.35343 | 0.45 | 0.45 | 0 | 15.2192 | [13 empty] | 0 | 0.45
[6 empty] | 0.46 | 0.46 | 15.2503 | 548.702 | 255.5206 | 25.6575 | 15.2503 | 14.1526 | 0 | 0.35983 | 0.46 | 0.46 | 0 | 15.2503 | [13 empty] | 0 | 0.46
[6 empty] | 0.47 | 0.47 | 15.2802 | 552.9782 | 259.769 | 25.324 | 15.2802 | 14.1761 | 0 | 0.36615 | 0.47 | 0.47 | 0 | 15.2802 | [13 empty] | 0 | 0.47
[6 empty] | 0.48 | 0.48 | 15.3091 | 557.1989 | 263.9619 | 24.9905 | 15.3091 | 14.1992 | 0 | 0.37241 | 0.48 | 0.48 | 0 | 15.3091 | [13 empty] | 0 | 0.48
[6 empty] | 0.49 | 0.49 | 15.3368 | 561.364 | 268.0992 | 24.6569 | 15.3368 | 14.222 | 0 | 0.37859 | 0.49 | 0.49 | 0 | 15.3368 | [13 empty] | 0 | 0.49
[6 empty] | 0.5 | 0.5 | 15.3634 | 565.4735 | 272.1809 | 24.3234 | 15.3634 | 14.2444 | 0 | 0.3847 | 0.5 | 0.5 | 0 | 15.3634 | [13 empty] | 0 | 0.5
[6 empty] | 0.51 | 0.51 | 15.3889 | 569.5274 | 276.207 | 23.9899 | 15.3889 | 14.2664 | 0 | 0.39074 | 0.51 | 0.51 | 0 | 15.3889 | [13 empty] | 0 | 0.51
[6 empty] | 0.52 | 0.52 | 15.4132 | 573.5257 | 280.1775 | 23.6563 | 15.4132 | 14.288 | 0 | 0.3967 | 0.52 | 0.52 | 0 | 15.4132 | [13 empty] | 0 | 0.52
[6 empty] | 0.53 | 0.53 | 15.4365 | 577.4684 | 284.0924 | 23.3228 | 15.4365 | 14.3093 | 0 | 0.40258 | 0.53 | 0.53 | 0 | 15.4365 | [13 empty] | 0 | 0.53
[6 empty] | 0.54 | 0.54 | 15.4587 | 581.3556 | 287.9518 | 22.9893 | 15.4587 | 14.3302 | 0 | 0.4084 | 0.54 | 0.54 | 0 | 15.4587 | [13 empty] | 0 | 0.54
[6 empty] | 0.55 | 0.55 | 15.4797 | 585.1871 | 291.7555 | 22.6558 | 15.4797 | 14.3507 | 0 | 0.41413 | 0.55 | 0.55 | 0 | 15.4797 | [13 empty] | 0 | 0.55
[6 empty] | 0.56 | 0.56 | 15.4996 | 588.9631 | 295.5037 | 22.3222 | 15.4996 | 14.3709 | 0 | 0.41979 | 0.56 | 0.56 | 0 | 15.4996 | [13 empty] | 0 | 0.56
[6 empty] | 0.57 | 0.57 | 15.5185 | 592.6834 | 299.1963 | 21.9887 | 15.5185 | 14.3907 | 0 | 0.42537 | 0.57 | 0.57 | 0 | 15.5185 | [13 empty] | 0 | 0.57
[6 empty] | 0.58 | 0.58 | 15.5362 | 596.3482 | 302.8333 | 21.6552 | 15.5362 | 14.4101 | 0 | 0.43087 | 0.58 | 0.58 | 0 | 15.5362 | [13 empty] | 0 | 0.58
[6 empty] | 0.59 | 0.59 | 15.5528 | 599.9574 | 306.4147 | 21.3216 | 15.5528 | 14.4291 | 0 | 0.43629 | 0.59 | 0.59 | 0 | 15.5528 | [13 empty] | 0 | 0.59
[6 empty] | 0.6 | 0.6 | 15.5682 | 603.511 | 309.9405 | 20.9881 | 15.5682 | 14.4478 | 0 | 0.44163 | 0.6 | 0.6 | 0 | 15.5682 | [13 empty] | 0 | 0.6
[6 empty] | 0.61 | 0.61 | 15.5826 | 607.009 | 313.4107 | 20.6546 | 15.5826 | 14.4661 | 0 | 0.44689 | 0.61 | 0.61 | 0 | 15.5826 | [13 empty] | 0 | 0.61
[6 empty] | 0.62 | 0.62 | 15.5959 | 610.4515 | 316.8253 | 20.321 | 15.5959 | 14.484 | 0 | 0.45207 | 0.62 | 0.62 | 0 | 15.5959 | [13 empty] | 0 | 0.62
[6 empty] | 0.63 | 0.63 | 15.608 | 613.8383 | 320.1844 | 19.9875 | 15.608 | 14.5016 | 0 | 0.45717 | 0.63 | 0.63 | 0 | 15.608 | [13 empty] | 0 | 0.63
[6 empty] | 0.64 | 0.64 | 15.6191 | 617.1696 | 323.4878 | 19.654 | 15.6191 | 14.5188 | 0 | 0.46219 | 0.64 | 0.64 | 0 | 15.6191 | [13 empty] | 0 | 0.64
[6 empty] | 0.65 | 0.65 | 15.629 | 620.4452 | 326.7357 | 19.3204 | 15.629 | 14.5356 | 0 | 0.46713 | 0.65 | 0.65 | 0 | 15.629 | [13 empty] | 0 | 0.65
[6 empty] | 0.66 | 0.66 | 15.6378 | 623.6653 | 329.928 | 18.9869 | 15.6378 | 14.552 | 0 | 0.47199 | 0.66 | 0.66 | 0 | 15.6378 | [13 empty] | 0 | 0.66
[6 empty] | 0.67 | 0.67 | 15.6455 | 626.8298 | 333.0647 | 18.6534 | 15.6455 | 14.5681 | 0 | 0.47676 | 0.67 | 0.67 | 0 | 15.6455 | [13 empty] | 0 | 0.67
[6 empty] | 0.68 | 0.68 | 15.6521 | 629.9387 | 336.1458 | 18.3198 | 15.6521 | 14.5838 | 0 | 0.48145 | 0.68 | 0.68 | 0 | 15.6521 | [13 empty] | 0 | 0.68
[6 empty] | 0.69 | 0.69 | 15.6576 | 632.992 | 339.1713 | 17.9863 | 15.6576 | 14.5992 | 0 | 0.48605 | 0.69 | 0.69 | 0 | 15.6576 | [13 empty] | 0 | 0.69
[6 empty] | 0.7 | 0.7 | 15.6619 | 635.9897 | 342.1412 | 17.6528 | 15.6619 | 14.6141 | 0 | 0.49057 | 0.7 | 0.7 | 0 | 15.6619 | [13 empty] | 0 | 0.7
[6 empty] | 0.71 | 0.71 | 15.6652 | 638.9318 | 345.0555 | 17.3192 | 15.6652 | 14.6287 | 0 | 0.49501 | 0.71 | 0.71 | 0 | 15.6652 | [13 empty] | 0 | 0.71
[6 empty] | 0.72 | 0.72 | 15.6673 | 641.8184 | 347.9143 | 16.9857 | 15.6673 | 14.643 | 0 | 0.49936 | 0.72 | 0.72 | 0 | 15.6673 | [13 empty] | 0 | 0.72
[6 empty] | 0.73 | 0.73 | 15.6684 | 644.6493 | 350.7174 | 16.6522 | 15.6684 | 14.6568 | 0 | 0.50362 | 0.73 | 0.73 | 0 | 15.6684 | [13 empty] | 0 | 0.73
[6 empty] | 0.74 | 0.74 | 15.6683 | 647.4247 | 353.465 | 16.3187 | 15.6683 | 14.6703 | 0 | 0.5078 | 0.74 | 0.74 | 0 | 15.6683 | [13 empty] | 0 | 0.74
[6 empty] | 0.75 | 0.75 | 15.6671 | 650.1445 | 356.157 | 15.9851 | 15.6671 | 14.6834 | 0 | 0.5119 | 0.75 | 0.75 | 0 | 15.6671 | [13 empty] | 0 | 0.75
[6 empty] | 0.76 | 0.76 | 15.6648 | 652.8086 | 358.7934 | 15.6516 | 15.6648 | 14.6962 | 0 | 0.51591 | 0.76 | 0.76 | 0 | 15.6648 | [13 empty] | 0 | 0.76
[6 empty] | 0.77 | 0.77 | 15.6614 | 655.4172 | 361.3742 | 15.3181 | 15.6614 | 14.7085 | 0 | 0.51983 | 0.77 | 0.77 | 0 | 15.6614 | [13 empty] | 0 | 0.77
[6 empty] | 0.78 | 0.78 | 15.6568 | 657.9702 | 363.8994 | 14.9845 | 15.6568 | 14.7206 | 0 | 0.52366 | 0.78 | 0.78 | 0 | 15.6568 | [13 empty] | 0 | 0.78
[6 empty] | 0.79 | 0.79 | 15.6512 | 660.4677 | 366.369 | 14.651 | 15.6512 | 14.7322 | 0 | 0.52741 | 0.79 | 0.79 | 0 | 15.6512 | [13 empty] | 0 | 0.79
[6 empty] | 0.8 | 0.8 | 15.6444 | 662.9095 | 368.7831 | 14.3175 | 15.6444 | 14.7434 | 0 | 0.53108 | 0.8 | 0.8 | 0 | 15.6444 | [13 empty] | 0 | 0.8
[6 empty] | 0.81 | 0.81 | 15.6366 | 665.2957 | 371.1415 | 13.9839 | 15.6366 | 14.7543 | 0 | 0.53465 | 0.81 | 0.81 | 0 | 15.6366 | [13 empty] | 0 | 0.81
[6 empty] | 0.82 | 0.82 | 15.6276 | 667.6264 | 373.4444 | 13.6504 | 15.6276 | 14.7649 | 0 | 0.53814 | 0.82 | 0.82 | 0 | 15.6276 | [13 empty] | 0 | 0.82
[6 empty] | 0.83 | 0.83 | 15.6175 | 669.9015 | 375.6917 | 13.3169 | 15.6175 | 14.775 | 0 | 0.54154 | 0.83 | 0.83 | 0 | 15.6175 | [13 empty] | 0 | 0.83
[6 empty] | 0.84 | 0.84 | 15.6063 | 672.1209 | 377.8833 | 12.9833 | 15.6063 | 14.7848 | 0 | 0.54485 | 0.84 | 0.84 | 0 | 15.6063 | [13 empty] | 0 | 0.84
[6 empty] | 0.85 | 0.85 | 15.594 | 674.2848 | 380.0194 | 12.6498 | 15.594 | 14.7942 | 0 | 0.54808 | 0.85 | 0.85 | 0 | 15.594 | [13 empty] | 0 | 0.85
[6 empty] | 0.86 | 0.86 | 15.5806 | 676.3931 | 382.0999 | 12.3163 | 15.5806 | 14.8032 | 0 | 0.55122 | 0.86 | 0.86 | 0 | 15.5806 | [13 empty] | 0 | 0.86
[6 empty] | 0.87 | 0.87 | 15.566 | 678.4458 | 384.1249 | 11.9827 | 15.566 | 14.8119 | 0 | 0.55426 | 0.87 | 0.87 | 0 | 15.566 | [13 empty] | 0 | 0.87
[6 empty] | 0.88 | 0.88 | 15.5504 | 680.443 | 386.0942 | 11.6492 | 15.5504 | 14.8202 | 0 | 0.55723 | 0.88 | 0.88 | 0 | 15.5504 | [13 empty] | 0 | 0.88
[6 empty] | 0.89 | 0.89 | 15.5336 | 682.3845 | 388.0079 | 11.3157 | 15.5336 | 14.8281 | 0 | 0.5601 | 0.89 | 0.89 | 0 | 15.5336 | [13 empty] | 0 | 0.89
[6 empty] | 0.9 | 0.9 | 15.5158 | 684.2704 | 389.8661 | 10.9821 | 15.5158 | 14.8357 | 0 | 0.56289 | 0.9 | 0.9 | 0 | 15.5158 | [13 empty] | 0 | 0.9
[6 empty] | 0.91 | 0.91 | 15.4968 | 686.1008 | 391.6686 | 10.6486 | 15.4968 | 14.8429 | 0 | 0.56558 | 0.91 | 0.91 | 0 | 15.4968 | [13 empty] | 0 | 0.91
[6 empty] | 0.92 | 0.92 | 15.4767 | 687.8756 | 393.4156 | 10.3151 | 15.4767 | 14.8497 | 0 | 0.56819 | 0.92 | 0.92 | 0 | 15.4767 | [13 empty] | 0 | 0.92
[6 empty] | 0.93 | 0.93 | 15.4555 | 689.5947 | 395.107 | 9.9815 | 15.4555 | 14.8561 | 0 | 0.57072 | 0.93 | 0.93 | 0 | 15.4555 | [13 empty] | 0 | 0.93
[6 empty] | 0.94 | 0.94 | 15.4332 | 691.2583 | 396.7428 | 9.648 | 15.4332 | 14.8622 | 0 | 0.57315 | 0.94 | 0.94 | 0 | 15.4332 | [13 empty] | 0 | 0.94
[6 empty] | 0.95 | 0.95 | 15.4097 | 692.8663 | 398.323 | 9.3145 | 15.4097 | 14.8679 | 0 | 0.5755 | 0.95 | 0.95 | 0 | 15.4097 | [13 empty] | 0 | 0.95
[6 empty] | 0.96 | 0.96 | 15.3852 | 694.4188 | 399.8476 | 8.981 | 15.3852 | 14.8732 | 0 | 0.57776 | 0.96 | 0.96 | 0 | 15.3852 | [13 empty] | 0 | 0.96
[6 empty] | 0.97 | 0.97 | 15.3595 | 695.9156 | 401.3167 | 8.6474 | 15.3595 | 14.8782 | 0 | 0.57993 | 0.97 | 0.97 | 0 | 15.3595 | [13 empty] | 0 | 0.97
[6 empty] | 0.98 | 0.98 | 15.3328 | 697.3568 | 402.7301 | 8.3139 | 15.3328 | 14.8828 | 0 | 0.58201 | 0.98 | 0.98 | 0 | 15.3328 | [13 empty] | 0 | 0.98
[6 empty] | 0.99 | 0.99 | 15.3049 | 698.7425 | 404.0879 | 7.9804 | 15.3049 | 14.887 | 0 | 0.58401 | 0.99 | 0.99 | 0 | 15.3049 | [13 empty] | 0 | 0.99
[6 empty] | 1 | 1 | 15.2759 | 700.0725 | 405.3902 | 7.6468 | 15.2759 | 14.8909 | 0 | 0.58592 | 1 | 1 | 0 | 15.2759 | [13 empty] | 0 | 1
[6 empty] | 1.01 | 1.01 | 15.2458 | 701.347 | 406.6369 | 7.3133 | 15.2458 | 14.8943 | 0 | 0.58774 | 1.01 | 1.01 | 0 | 15.2458 | [13 empty] | 0 | 1.01
[6 empty] | 1.02 | 1.02 | 15.2146 | 702.5659 | 407.828 | 6.9798 | 15.2146 | 14.8975 | 0 | 0.58948 | 1.02 | 1.02 | 0 | 15.2146 | [13 empty] | 0 | 1.02
[6 empty] | 1.03 | 1.03 | 15.1823 | 703.7292 | 408.9635 | 6.6462 | 15.1823 | 14.9002 | 0 | 0.59113 | 1.03 | 1.03 | 0 | 15.1823 | [13 empty] | 0 | 1.03
[6 empty] | 1.04 | 1.04 | 15.1489 | 704.8369 | 410.0434 | 6.3127 | 15.1489 | 14.9026 | 0 | 0.59269 | 1.04 | 1.04 | 0 | 15.1489 | [13 empty] | 0 | 1.04
[6 empty] | 1.05 | 1.05 | 15.1143 | 705.889 | 411.0677 | 5.9792 | 15.1143 | 14.9046 | 0 | 0.59417 | 1.05 | 1.05 | 0 | 15.1143 | [13 empty] | 0 | 1.05
[6 empty] | 1.06 | 1.06 | 15.0786 | 706.8855 | 412.0364 | 5.6456 | 15.0786 | 14.9062 | 0 | 0.59556 | 1.06 | 1.06 | 0 | 15.0786 | [13 empty] | 0 | 1.06
[6 empty] | 1.07 | 1.07 | 15.0419 | 707.8265 | 412.9496 | 5.3121 | 15.0419 | 14.9074 | 0 | 0.59687 | 1.07 | 1.07 | 0 | 15.0419 | [13 empty] | 0 | 1.07
[6 empty] | 1.08 | 1.08 | 15.004 | 708.7118 | 413.8071 | 4.9786 | 15.004 | 14.9083 | 0 | 0.59809 | 1.08 | 1.08 | 0 | 15.004 | [13 empty] | 0 | 1.08
[6 empty] | 1.09 | 1.09 | 14.965 | 709.5416 | 414.6091 | 4.645 | 14.965 | 14.9088 | 0 | 0.59923 | 1.09 | 1.09 | 0 | 14.965 | [13 empty] | 0 | 1.09
[6 empty] | 1.1 | 1.1 | 14.9249 | 710.3157 | 415.3555 | 4.3115 | 14.9249 | 14.909 | 0 | 0.60028 | 1.1 | 1.1 | 0 | 14.9249 | [13 empty] | 0 | 1.1
[6 empty] | 1.11 | 1.11 | 14.8837 | 711.0343 | 416.0463 | 3.978 | 14.8837 | 14.9088 | 0 | 0.60125 | 1.11 | 1.11 | 0 | 14.8837 | [13 empty] | 0 | 1.11
[6 empty] | 1.12 | 1.12 | 14.8414 | 711.6973 | 416.6815 | 3.6444 | 14.8414 | 14.9082 | 0 | 0.60213 | 1.12 | 1.12 | 0 | 14.8414 | [13 empty] | 0 | 1.12
[6 empty] | 1.13 | 1.13 | 14.7979 | 712.3047 | 417.2611 | 3.3109 | 14.7979 | 14.9072 | 0 | 0.60293 | 1.13 | 1.13 | 0 | 14.7979 | [13 empty] | 0 | 1.13
[6 empty] | 1.14 | 1.14 | 14.7534 | 712.8566 | 417.7851 | 2.9774 | 14.7534 | 14.9059 | 0 | 0.60365 | 1.14 | 1.14 | 0 | 14.7534 | [13 empty] | 0 | 1.14
[6 empty] | 1.15 | 1.15 | 14.7077 | 713.3528 | 418.2536 | 2.6438 | 14.7077 | 14.9042 | 0 | 0.60429 | 1.15 | 1.15 | 0 | 14.7077 | [13 empty] | 0 | 1.15
[6 empty] | 1.16 | 1.16 | 14.6609 | 713.7934 | 418.6664 | 2.3103 | 14.6609 | 14.9021 | 0 | 0.60484 | 1.16 | 1.16 | 0 | 14.6609 | [13 empty] | 0 | 1.16
[6 empty] | 1.17 | 1.17 | 14.613 | 714.1785 | 419.0237 | 1.9768 | 14.613 | 14.8996 | 0 | 0.60531 | 1.17 | 1.17 | 0 | 14.613 | [13 empty] | 0 | 1.17
[6 empty] | 1.18 | 1.18 | 14.564 | 714.5079 | 419.3253 | 1.6433 | 14.564 | 14.8968 | 0 | 0.6057 | 1.18 | 1.18 | 0 | 14.564 | [13 empty] | 0 | 1.18
[6 empty] | 1.19 | 1.19 | 14.5139 | 714.7818 | 419.5714 | 1.3097 | 14.5139 | 14.8936 | 0 | 0.60601 | 1.19 | 1.19 | 0 | 14.5139 | [13 empty] | 0 | 1.19
[6 empty] | 1.2 | 1.2 | 14.4627 | 715.0001 | 419.7619 | 0.97619 | 14.4627 | 14.8901 | 0 | 0.60625 | 1.2 | 1.2 | 0 | 14.4627 | [13 empty] | 0 | 1.2
[6 empty] | 1.21 | 1.21 | 14.4104 | 715.1628 | 419.8968 | 0.64266 | 14.4104 | 14.8861 | 0 | 0.6064 | 1.21 | 1.21 | 0 | 14.4104 | [13 empty] | 0 | 1.21
[6 empty] | 1.22 | 1.22 | 14.3569 | 715.2699 | 419.9761 | 0.30913 | 14.3569 | 14.8818 | 0 | 0.60647 | 1.22 | 1.22 | 0 | 14.3569 | [13 empty] | 0 | 1.22
[6 empty] | 1.23 | 1.23 | 5 | 715.3214 | 415.3487 | -55.8386 | 5 | 14.8021 | 0 | 0.6019 | 1.23 | 1.23 | 0 | 5 | [13 empty] | 0 | 1.23
[6 empty] | 1.24 | 1.24 | 5 | 706.015 | 406.1353 | -54.7218 | 5 | 14.7237 | 0 | 0.59743 | 1.24 | 1.24 | 0 | 5 | [13 empty] | 0 | 1.24
[6 empty] | 1.25 | 1.25 | 5 | 696.8947 | 397.1062 | -53.6274 | 5 | 14.6465 | 0 | 0.59305 | 1.25 | 1.25 | 0 | 5 | [13 empty] | 0 | 1.25
[6 empty] | 1.26 | 1.26 | 5 | 687.9568 | 388.2577 | -52.5548 | 5 | 14.5706 | 0 | 0.58875 | 1.26 | 1.26 | 0 | 5 | [13 empty] | 0 | 1.26
[6 empty] | 1.27 | 1.27 | 5 | 679.1977 | 379.5861 | -51.5037 | 5 | 14.4958 | 0 | 0.58454 | 1.27 | 1.27 | 0 | 5 | [13 empty] | 0 | 1.27
[6 empty] | 1.28 | 1.28 | 5 | 670.6137 | 371.088 | -50.4736 | 5 | 14.4222 | 0 | 0.58041 | 1.28 | 1.28 | 0 | 5 | [13 empty] | 0 | 1.28
[6 empty] | 1.29 | 1.29 | 5 | 662.2014 | 362.7599 | -49.4642 | 5 | 14.3497 | 0 | 0.57637 | 1.29 | 1.29 | 0 | 5 | [13 empty] | 0 | 1.29
[6 empty] | 1.3 | 1.3 | 5 | 653.9574 | 354.5983 | -48.4749 | 5 | 14.2784 | 0 | 0.57241 | 1.3 | 1.3 | 0 | 5 | [13 empty] | 0 | 1.3
[6 empty] | 1.31 | 1.31 | 5 | 645.8783 | 346.5999 | -47.5054 | 5 | 14.2081 | 0 | 0.56852 | 1.31 | 1.31 | 0 | 5 | [13 empty] | 0 | 1.31
[6 empty] | 1.32 | 1.32 | 5 | 637.9607 | 338.7615 | -46.5553 | 5 | 14.1388 | 0 | 0.56472 | 1.32 | 1.32 | 0 | 5 | [13 empty] | 0 | 1.32
[6 empty] | 1.33 | 1.33 | 5 | 630.2015 | 331.0799 | -45.6242 | 5 | 14.0706 | 0 | 0.56099 | 1.33 | 1.33 | 0 | 5 | [13 empty] | 0 | 1.33
[6 empty] | 1.34 | 1.34 | 5 | 622.5975 | 323.5519 | -44.7117 | 5 | 14.0034 | 0 | 0.55733 | 1.34 | 1.34 | 0 | 5 | [13 empty] | 0 | 1.34
[6 empty] | 1.35 | 1.35 | 5 | 615.1455 | 316.1745 | -43.8175 | 5 | 13.9372 | 0 | 0.55375 | 1.35 | 1.35 | 0 | 5 | [13 empty] | 0 | 1.35
[6 empty] | 1.36 | 1.36 | 5 | 607.8426 | 308.9446 | -42.9411 | 5 | 13.872 | 0 | 0.55024 | 1.36 | 1.36 | 0 | 5 | [13 empty] | 0 | 1.36
[6 empty] | 1.37 | 1.37 | 5 | 600.6858 | 301.8593 | -42.0823 | 5 | 13.8077 | 0 | 0.5468 | 1.37 | 1.37 | 0 | 5 | [13 empty] | 0 | 1.37
[6 empty] | 1.38 | 1.38 | 5 | 593.672 | 294.9158 | -41.2406 | 5 | 13.7443 | 0 | 0.54343 | 1.38 | 1.38 | 0 | 5 | [13 empty] | 0 | 1.38
[6 empty] | 1.39 | 1.39 | 5 | 586.7986 | 288.1111 | -40.4158 | 5 | 13.6819 | 0 | 0.54012 | 1.39 | 1.39 | 0 | 5 | [13 empty] | 0 | 1.39
[6 empty] | 1.4 | 1.4 | 5 | 580.0626 | 281.4424 | -39.6075 | 5 | 13.6203 | 0 | 0.53689 | 1.4 | 1.4 | 0 | 5 | [13 empty] | 0 | 1.4
[6 empty] | 1.41 | 1.41 | 5 | 573.4614 | 274.9072 | -38.8154 | 5 | 13.5596 | 0 | 0.53371 | 1.41 | 1.41 | 0 | 5 | [13 empty] | 0 | 1.41
[6 empty] | 1.42 | 1.42 | 5 | 566.9921 | 268.5027 | -38.0391 | 5 | 13.4997 | 0 | 0.5306 | 1.42 | 1.42 | 0 | 5 | [13 empty] | 0 | 1.42
[6 empty] | 1.43 | 1.43 | 5 | 560.6523 | 262.2262 | -37.2783 | 5 | 13.4407 | 0 | 0.52756 | 1.43 | 1.43 | 0 | 5 | [13 empty] | 0 | 1.43
[6 empty] | 1.44 | 1.44 | 5 | 554.4393 | 256.0753 | -36.5327 | 5 | 13.3825 | 0 | 0.52457 | 1.44 | 1.44 | 0 | 5 | [13 empty] | 0 | 1.44
[6 empty] | 1.45 | 1.45 | 5 | 548.3505 | 250.0474 | -35.8021 | 5 | 13.3251 | 0 | 0.52164 | 1.45 | 1.45 | 0 | 5 | [13 empty] | 0 | 1.45
[6 empty] | 1.46 | 1.46 | 5 | 542.3835 | 244.1401 | -35.086 | 5 | 13.2685 | 0 | 0.51877 | 1.46 | 1.46 | 0 | 5 | [13 empty] | 0 | 1.46
[6 empty] | 1.47 | 1.47 | 5 | 536.5358 | 238.3509 | -34.3843 | 5 | 13.2126 | 0 | 0.51596 | 1.47 | 1.47 | 0 | 5 | [13 empty] | 0 | 1.47
[6 empty] | 1.48 | 1.48 | 5 | 530.8051 | 232.6775 | -33.6966 | 5 | 13.1575 | 0 | 0.51321 | 1.48 | 1.48 | 0 | 5 | [13 empty] | 0 | 1.48
[6 empty] | 1.49 | 1.49 | 5 | 525.189 | 227.1175 | -33.0227 | 5 | 13.1031 | 0 | 0.51051 | 1.49 | 1.49 | 0 | 5 | [13 empty] | 0 | 1.49
[6 empty] | 1.5 | 1.5 | 5 | 519.6852 | 221.6688 | -32.3622 | 5 | 13.0494 | 0 | 0.50786 | 1.5 | 1.5 | 0 | 5 | [13 empty] | 0 | 1.5
[6 empty] | 1.51 | 1.51 | 5 | 514.2915 | 216.329 | -31.715 | 5 | 12.9965 | 0 | 0.50527 | 1.51 | 1.51 | 0 | 5 | [13 empty] | 0 | 1.51
[6 empty] | 1.52 | 1.52 | 5 | 509.0057 | 211.0961 | -31.0807 | 5 | 12.9442 | 0 | 0.50273 | 1.52 | 1.52 | 0 | 5 | [13 empty] | 0 | 1.52
[6 empty] | 1.53 | 1.53 | 5 | 503.8255 | 205.9677 | -30.4591 | 5 | 12.8926 | 0 | 0.50024 | 1.53 | 1.53 | 0 | 5 | [13 empty] | 0 | 1.53
[6 empty] | 1.54 | 1.54 | 5 | 498.749 | 200.942 | -29.8499 | 5 | 12.8417 | 0 | 0.4978 | 1.54 | 1.54 | 0 | 5 | [13 empty] | 0 | 1.54
[6 empty] | 1.55 | 1.55 | 5 | 493.7741 | 196.0168 | -29.2529 | 5 | 12.7914 | 0 | 0.49541 | 1.55 | 1.55 | 0 | 5 | [13 empty] | 0 | 1.55
[6 empty] | 1.56 | 1.56 | 5 | 488.8986 | 191.19 | -28.6678 | 5 | 12.7418 | 0 | 0.49307 | 1.56 | 1.56 | 0 | 5 | [13 empty] | 0 | 1.56
[6 empty] | 1.57 | 1.57 | 5 | 484.1206 | 186.4598 | -28.0945 | 5 | 12.6928 | 0 | 0.49077 | 1.57 | 1.57 | 0 | 5 | [13 empty] | 0 | 1.57
[6 empty] | 1.58 | 1.58 | 5 | 479.4382 | 181.8243 | -27.5326 | 5 | 12.6444 | 0 | 0.48852 | 1.58 | 1.58 | 0 | 5 | [13 empty] | 0 | 1.58
[6 empty] | 1.59 | 1.59 | 5 | 474.8494 | 177.2814 | -26.9819 | 5 | 12.5966 | 0 | 0.48631 | 1.59 | 1.59 | 0 | 5 | [13 empty] | 0 | 1.59
[6 empty] | 1.6 | 1.6 | 5 | 470.3524 | 172.8294 | -26.4423 | 5 | 12.5495 | 0 | 0.48415 | 1.6 | 1.6 | 0 | 5 | [13 empty] | 0 | 1.6
[6 empty] | 1.61 | 1.61 | 5 | 465.9454 | 168.4664 | -25.9134 | 5 | 12.5029 | 0 | 0.48203 | 1.61 | 1.61 | 0 | 5 | [13 empty] | 0 | 1.61
[6 empty] | 1.62 | 1.62 | 5 | 461.6265 | 164.1907 | -25.3952 | 5 | 12.4568 | 0 | 0.47996 | 1.62 | 1.62 | 0 | 5 | [13 empty] | 0 | 1.62
[6 empty] | 1.63 | 1.63 | 5 | 457.394 | 160.0005 | -24.8873 | 5 | 12.4114 | 0 | 0.47792 | 1.63 | 1.63 | 0 | 5 | [13 empty] | 0 | 1.63
[6 empty] | 1.64 | 1.64 | 5 | 453.2461 | 155.8941 | -24.3895 | 5 | 12.3664 | 0 | 0.47593 | 1.64 | 1.64 | 0 | 5 | [13 empty] | 0 | 1.64
[6 empty] | 1.65 | 1.65 | 5 | 449.1812 | 151.8698 | -23.9017 | 5 | 12.3221 | 0 | 0.47397 | 1.65 | 1.65 | 0 | 5 | [13 empty] | 0 | 1.65
[6 empty] | 1.66 | 1.66 | 5 | 445.1975 | 147.926 | -23.4237 | 5 | 12.2782 | 0 | 0.47206 | 1.66 | 1.66 | 0 | 5 | [13 empty] | 0 | 1.66
[6 empty] | 1.67 | 1.67 | 5 | 441.2936 | 144.0611 | -22.9552 | 5 | 12.2349 | 0 | 0.47018 | 1.67 | 1.67 | 0 | 5 | [13 empty] | 0 | 1.67
[6 empty] | 1.68 | 1.68 | 5 | 437.4677 | 140.2735 | -22.4961 | 5 | 12.1921 | 0 | 0.46834 | 1.68 | 1.68 | 0 | 5 | [13 empty] | 0 | 1.68
[6 empty] | 1.69 | 1.69 | 5 | 433.7184 | 136.5616 | -22.0462 | 5 | 12.1498 | 0 | 0.46654 | 1.69 | 1.69 | 0 | 5 | [13 empty] | 0 | 1.69
[6 empty] | 1.7 | 1.7 | 5 | 430.044 | 132.924 | -21.6053 | 5 | 12.108 | 0 | 0.46477 | 1.7 | 1.7 | 0 | 5 | [13 empty] | 0 | 1.7
[6 empty] | 1.71 | 1.71 | 5 | 426.4431 | 129.3591 | -21.1732 | 5 | 12.0667 | 0 | 0.46304 | 1.71 | 1.71 | 0 | 5 | [13 empty] | 0 | 1.71
[6 empty] | 1.72 | 1.72 | 5 | 422.9142 | 125.8655 | -20.7497 | 5 | 12.0258 | 0 | 0.46135 | 1.72 | 1.72 | 0 | 5 | [13 empty] | 0 | 1.72
[6 empty] | 1.73 | 1.73 | 5 | 419.456 | 122.4418 | -20.3347 | 5 | 11.9854 | 0 | 0.45968 | 1.73 | 1.73 | 0 | 5 | [13 empty] | 0 | 1.73
[6 empty] | 1.74 | 1.74 | 5 | 416.0668 | 119.0866 | -19.928 | 5 | 11.9455 | 0 | 0.45806 | 1.74 | 1.74 | 0 | 5 | [13 empty] | 0 | 1.74
[6 empty] | 1.75 | 1.75 | 5 | 412.7455 | 115.7985 | -19.5295 | 5 | 11.906 | 0 | 0.45646 | 1.75 | 1.75 | 0 | 5 | [13 empty] | 0 | 1.75
[6 empty] | 1.76 | 1.76 | 5 | 409.4906 | 112.5761 | -19.1389 | 5 | 11.867 | 0 | 0.45489 | 1.76 | 1.76 | 0 | 5 | [13 empty] | 0 | 1.76
[6 empty] | 1.77 | 1.77 | 5 | 406.3008 | 109.4182 | -18.7561 | 5 | 11.8284 | 0 | 0.45336 | 1.77 | 1.77 | 0 | 5 | [13 empty] | 0 | 1.77
[6 empty] | 1.78 | 1.78 | 5 | 403.1748 | 106.3235 | -18.381 | 5 | 11.7903 | 0 | 0.45186 | 1.78 | 1.78 | 0 | 5 | [13 empty] | 0 | 1.78
[6 empty] | 1.79 | 1.79 | 5 | 400.1113 | 103.2906 | -18.0134 | 5 | 11.7526 | 0 | 0.45039 | 1.79 | 1.79 | 0 | 5 | [13 empty] | 0 | 1.79
[6 empty] | 1.8 | 1.8 | 5 | 397.109 | 100.3184 | -17.6531 | 5 | 11.7153 | 0 | 0.44894 | 1.8 | 1.8 | 0 | 5 | [13 empty] | 0 | 1.8
[6 empty] | 1.81 | 1.81 | 5 | 394.1669 | 97.4056 | -17.3 | 5 | 11.6784 | 0 | 0.44753 | 1.81 | 1.81 | 0 | 5 | [13 empty] | 0 | 1.81
[6 empty] | 1.82 | 1.82 | 5 | 391.2835 | 94.5511 | -16.954 | 5 | 11.6419 | 0 | 0.44614 | 1.82 | 1.82 | 0 | 5 | [13 empty] | 0 | 1.82
[6 empty] | 1.83 | 1.83 | 5 | 388.4579 | 91.7537 | -16.6149 | 5 | 11.6058 | 0 | 0.44478 | 1.83 | 1.83 | 0 | 5 | [13 empty] | 0 | 1.83
[6 empty] | 1.84 | 1.84 | 5 | 385.6887 | 89.0123 | -16.2826 | 5 | 11.5701 | 0 | 0.44345 | 1.84 | 1.84 | 0 | 5 | [13 empty] | 0 | 1.84
[6 empty] | 1.85 | 1.85 | 5 | 382.9749 | 86.3256 | -15.957 | 5 | 11.5348 | 0 | 0.44215 | 1.85 | 1.85 | 0 | 5 | [13 empty] | 0 | 1.85
[6 empty] | 1.86 | 1.86 | 5 | 380.3154 | 83.6927 | -15.6379 | 5 | 11.4998 | 0 | 0.44087 | 1.86 | 1.86 | 0 | 5 | [13 empty] | 0 | 1.86
[6 empty] | 1.87 | 1.87 | 5 | 377.7091 | 81.1125 | -15.3251 | 5 | 11.4652 | 0 | 0.43962 | 1.87 | 1.87 | 0 | 5 | [13 empty] | 0 | 1.87
[6 empty] | 1.88 | 1.88 | 5 | 375.1549 | 78.5838 | -15.0186 | 5 | 11.431 | 0 | 0.43839 | 1.88 | 1.88 | 0 | 5 | [13 empty] | 0 | 1.88
[6 empty] | 1.89 | 1.89 | 5 | 372.6518 | 76.1058 | -14.7182 | 5 | 11.3972 | 0 | 0.43719 | 1.89 | 1.89 | 0 | 5 | [13 empty] | 0 | 1.89
[6 empty] | 1.9 | 1.9 | 5 | 370.1988 | 73.6773 | -14.4239 | 5 | 11.3637 | 0 | 0.43601 | 1.9 | 1.9 | 0 | 5 | [13 empty] | 0 | 1.9
[6 empty] | 1.91 | 1.91 | 5 | 367.7948 | 71.2973 | -14.1354 | 5 | 11.3305 | 0 | 0.43485 | 1.91 | 1.91 | 0 | 5 | [13 empty] | 0 | 1.91
[6 empty] | 1.92 | 1.92 | 5 | 365.4389 | 68.965 | -13.8527 | 5 | 11.2977 | 0 | 0.43372 | 1.92 | 1.92 | 0 | 5 | [13 empty] | 0 | 1.92
[6 empty] | 1.93 | 1.93 | 5 | 363.1301 | 66.6793 | -13.5756 | 5 | 11.2653 | 0 | 0.43261 | 1.93 | 1.93 | 0 | 5 | [13 empty] | 0 | 1.93
[6 empty] | 1.94 | 1.94 | 5 | 360.8675 | 64.4393 | -13.3041 | 5 | 11.2331 | 0 | 0.43152 | 1.94 | 1.94 | 0 | 5 | [13 empty] | 0 | 1.94
[6 empty] | 1.95 | 1.95 | 5 | 358.6502 | 62.2441 | -13.038 | 5 | 11.2013 | 0 | 0.43046 | 1.95 | 1.95 | 0 | 5 | [13 empty] | 0 | 1.95
[6 empty] | 1.96 | 1.96 | 5 | 356.4772 | 60.0929 | -12.7773 | 5 | 11.1699 | 0 | 0.42941 | 1.96 | 1.96 | 0 | 5 | [13 empty] | 0 | 1.96
[6 empty] | 1.97 | 1.97 | 5 | 354.3476 | 57.9846 | -12.5217 | 5 | 11.1387 | 0 | 0.42839 | 1.97 | 1.97 | 0 | 5 | [13 empty] | 0 | 1.97
[6 empty] | 1.98 | 1.98 | 5 | 352.2607 | 55.9185 | -12.2713 | 5 | 11.1079 | 0 | 0.42738 | 1.98 | 1.98 | 0 | 5 | [13 empty] | 0 | 1.98
[6 empty] | 1.99 | 1.99 | 5 | 350.2155 | 53.8938 | -12.0259 | 5 | 11.0773 | 0 | 0.4264 | 1.99 | 1.99 | 0 | 5 | [13 empty] | 0 | 1.99
[6 empty] | 2 | 2 | 5 | 348.2112 | 51.9095 | -11.7853 | 5 | 11.0471 | 0 | 0.42544 | 2 | 2 | 0 | 5 | [13 empty] | 0 | 2
[6 empty] | 2.01 | 2.01 | 5 | 346.2469 | 49.9649 | -11.5496 | 5 | 11.0171 | 0 | 0.42449 | 2.01 | 2.01 | 0 | 5 | [13 empty] | 0 | 2.01
[6 empty] | 2.02 | 2.02 | 5 | 344.322 | 48.0592 | -11.3186 | 5 | 10.9875 | 0 | 0.42357 | 2.02 | 2.02 | 0 | 5 | [13 empty] | 0 | 2.02
[6 empty] | 2.03 | 2.03 | 5 | 342.4356 | 46.1917 | -11.0923 | 5 | 10.9582 | 0 | 0.42266 | 2.03 | 2.03 | 0 | 5 | [13 empty] | 0 | 2.03
[6 empty] | 2.04 | 2.04 | 5 | 340.5869 | 44.3614 | -10.8704 | 5 | 10.9291 | 0 | 0.42177 | 2.04 | 2.04 | 0 | 5 | [13 empty] | 0 | 2.04
[6 empty] | 2.05 | 2.05 | 5 | 338.7751 | 42.5678 | -10.653 | 5 | 10.9003 | 0 | 0.4209 | 2.05 | 2.05 | 0 | 5 | [13 empty] | 0 | 2.05
[6 empty] | 2.06 | 2.06 | 5 | 336.9996 | 40.8101 | -10.44 | 5 | 10.8718 | 0 | 0.42005 | 2.06 | 2.06 | 0 | 5 | [13 empty] | 0 | 2.06
[6 empty] | 2.07 | 2.07 | 5 | 335.2596 | 39.0875 | -10.2312 | 5 | 10.8436 | 0 | 0.41921 | 2.07 | 2.07 | 0 | 5 | [13 empty] | 0 | 2.07
[6 empty] | 2.08 | 2.08 | 5 | 333.5544 | 37.3993 | -10.0265 | 5 | 10.8156 | 0 | 0.41839 | 2.08 | 2.08 | 0 | 5 | [13 empty] | 0 | 2.08
[6 empty] | 2.09 | 2.09 | 5 | 331.8833 | 35.745 | -9.826 | 5 | 10.7879 | 0 | 0.41759 | 2.09 | 2.09 | 0 | 5 | [13 empty] | 0 | 2.09
[6 empty] | 2.1 | 2.1 | 5 | 330.2457 | 34.1237 | -9.6295 | 5 | 10.7605 | 0 | 0.4168 | 2.1 | 2.1 | 0 | 5 | [13 empty] | 0 | 2.1
[6 empty] | 2.11 | 2.11 | 5 | 328.6408 | 32.5348 | -9.4369 | 5 | 10.7333 | 0 | 0.41603 | 2.11 | 2.11 | 0 | 5 | [13 empty] | 0 | 2.11
[6 empty] | 2.12 | 2.12 | 5 | 327.0679 | 30.9777 | -9.2482 | 5 | 10.7064 | 0 | 0.41527 | 2.12 | 2.12 | 0 | 5 | [13 empty] | 0 | 2.12
[6 empty] | 2.13 | 2.13 | 5 | 325.5266 | 29.4518 | -9.0632 | 5 | 10.6797 | 0 | 0.41453 | 2.13 | 2.13 | 0 | 5 | [13 empty] | 0 | 2.13
[6 empty] | 2.14 | 2.14 | 5 | 324.0161 | 27.9563 | -8.8819 | 5 | 10.6533 | 0 | 0.41381 | 2.14 | 2.14 | 0 | 5 | [13 empty] | 0 | 2.14
[6 empty] | 2.15 | 2.15 | 5 | 322.5357 | 26.4908 | -8.7043 | 5 | 10.6271 | 0 | 0.4131 | 2.15 | 2.15 | 0 | 5 | [13 empty] | 0 | 2.15
[6 empty] | 2.16 | 2.16 | 5 | 321.085 | 25.0546 | -8.5302 | 5 | 10.6012 | 0 | 0.4124 | 2.16 | 2.16 | 0 | 5 | [13 empty] | 0 | 2.16
[6 empty] | 2.17 | 2.17 | 5 | 319.6633 | 23.6471 | -8.3596 | 5 | 10.5755 | 0 | 0.41172 | 2.17 | 2.17 | 0 | 5 | [13 empty] | 0 | 2.17
[6 empty] | 2.18 | 2.18 | 5 | 318.2701 | 22.2678 | -8.1924 | 5 | 10.5501 | 0 | 0.41105 | 2.18 | 2.18 | 0 | 5 | [13 empty] | 0 | 2.18
[6 empty] | 2.19 | 2.19 | 5 | 316.9047 | 20.9161 | -8.0286 | 5 | 10.5248 | 0 | 0.41039 | 2.19 | 2.19 | 0 | 5 | [13 empty] | 0 | 2.19
[6 empty] | 2.2 | 2.2 | 5 | 315.5666 | 19.5913 | -7.868 | 5 | 10.4998 | 0 | 0.40975 | 2.2 | 2.2 | 0 | 5 | [13 empty] | 0 | 2.2
[6 empty] | 2.21 | 2.21 | 5 | 314.2552 | 18.2931 | -7.7106 | 5 | 10.4751 | 0 | 0.40912 | 2.21 | 2.21 | 0 | 5 | [13 empty] | 0 | 2.21
[6 empty] | 2.22 | 2.22 | 5 | 312.9701 | 17.0209 | -7.5564 | 5 | 10.4505 | 0 | 0.4085 | 2.22 | 2.22 | 0 | 5 | [13 empty] | 0 | 2.22
[6 empty] | 2.23 | 2.23 | 5 | 311.7107 | 15.7741 | -7.4053 | 5 | 10.4262 | 0 | 0.40789 | 2.23 | 2.23 | 0 | 5 | [13 empty] | 0 | 2.23
[6 empty] | 2.24 | 2.24 | 5 | 310.4765 | 14.5522 | -7.2572 | 5 | 10.4021 | 0 | 0.4073 | 2.24 | 2.24 | 0 | 5 | [13 empty] | 0 | 2.24
[6 empty] | 2.25 | 2.25 | 5 | 309.267 | 13.3548 | -7.112 | 5 | 10.3782 | 0 | 0.40672 | 2.25 | 2.25 | 0 | 5 | [13 empty] | 0 | 2.25
[6 empty] | 2.26 | 2.26 | 5 | 308.0816 | 12.1813 | -6.9698 | 5 | 10.3545 | 0 | 0.40615 | 2.26 | 2.26 | 0 | 5 | [13 empty] | 0 | 2.26
[6 empty] | 2.27 | 2.27 | 5 | 306.92 | 11.0313 | -6.8304 | 5 | 10.331 | 0 | 0.40559 | 2.27 | 2.27 | 0 | 5 | [13 empty] | 0 | 2.27
[6 empty] | 2.28 | 2.28 | 5 | 305.7816 | 9.9042 | -6.6938 | 5 | 10.3077 | 0 | 0.40504 | 2.28 | 2.28 | 0 | 5 | [13 empty] | 0 | 2.28
[6 empty] | 2.29 | 2.29 | 5 | 304.666 | 8.7998 | -6.5599 | 5 | 10.2846 | 0 | 0.40451 | 2.29 | 2.29 | 0 | 5 | [13 empty] | 0 | 2.29
[6 empty] | 2.3 | 2.3 | 5 | 303.5727 | 7.7174 | -6.4287 | 5 | 10.2617 | 0 | 0.40398 | 2.3 | 2.3 | 0 | 5 | [13 empty] | 0 | 2.3
[6 empty] | 2.31 | 2.31 | 5 | 302.5012 | 6.6566 | -6.3001 | 5 | 10.2391 | 0 | 0.40347 | 2.31 | 2.31 | 0 | 5 | [13 empty] | 0 | 2.31
[6 empty] | 2.32 | 2.32 | 5 | 301.4512 | 5.6171 | -6.1741 | 5 | 10.2166 | 0 | 0.40296 | 2.32 | 2.32 | 0 | 5 | [13 empty] | 0 | 2.32
[6 empty] | 2.33 | 2.33 | 5 | 300.4222 | 4.5984 | -6.0507 | 5 | 10.1943 | 0 | 0.40247 | 2.33 | 2.33 | 0 | 5 | [13 empty] | 0 | 2.33
[6 empty] | 2.34 | 2.34 | 5 | 299.4137 | 3.6 | -5.9296 | 5 | 10.1722 | 0 | 0.40198 | 2.34 | 2.34 | 0 | 5 | [13 empty] | 0 | 2.34
[6 empty] | 2.35 | 2.35 | 5 | 298.4254 | 2.6216 | -5.8111 | 5 | 10.1503 | 0 | 0.40151 | 2.35 | 2.35 | 0 | 5 | [13 empty] | 0 | 2.35
[6 empty] | 2.36 | 2.36 | 5 | 297.4569 | 1.6628 | -5.6948 | 5 | 10.1285 | 0 | 0.40104 | 2.36 | 2.36 | 0 | 5 | [13 empty] | 0 | 2.36
[6 empty] | 2.37 | 2.37 | 5 | 296.5078 | 0.72316 | -5.5809 | 5 | 10.107 | 0 | 0.40058 | 2.37 | 2.37 | 0 | 5 | [13 empty] | 0 | 2.37
[6 empty] | 2.38 | 2.38 | 5 | 295.5776 | -0.19769 | -5.4693 | 5 | 10.0856 | 0 | 0.40014 | 2.38 | 2.38 | 0 | 5 | [13 empty] | 0 | 2.38
[6 empty] | 2.39 | 2.39 | 5 | 294.6661 | -1.1001 | -5.3599 | 5 | 10.0644 | 0 | 0.3997 | 2.39 | 2.39 | 0 | 5 | [13 empty] | 0 | 2.39
[6 empty] | 0 | 2.4 | 12.7088 | 293.7728 | 1.8699 | 41 | 12.7088 | 10.0754 | 0 | 0.40822 | 2.4 | 2.4 | 0 | 12.7088 | [13 empty] | 0 | 2.4
[6 empty] | 0.01 | 2.41 | 12.7899 | 300.6061 | 8.6754 | 40.6665 | 12.7899 | 10.0866 | 0 | 0.41672 | 2.41 | 2.41 | 0 | 12.7899 | [13 empty] | 0 | 2.41
[6 empty] | 0.02 | 2.42 | 12.8698 | 307.3838 | 15.4254 | 40.3329 | 12.8698 | 10.0981 | 0 | 0.42521 | 2.42 | 2.42 | 0 | 12.8698 | [13 empty] | 0 | 2.42
[6 empty] | 0.03 | 2.43 | 12.9487 | 314.106 | 22.1197 | 39.9994 | 12.9487 | 10.1098 | 0 | 0.43368 | 2.43 | 2.43 | 0 | 12.9487 | [13 empty] | 0 | 2.43
[6 empty] | 0.04 | 2.44 | 13.0264 | 320.7726 | 28.7585 | 39.6659 | 13.0264 | 10.1217 | 0 | 0.44213 | 2.44 | 2.44 | 0 | 13.0264 | [13 empty] | 0 | 2.44
[6 empty] | 0.05 | 2.45 | 13.1031 | 327.3835 | 35.3417 | 39.3323 | 13.1031 | 10.1338 | 0 | 0.45055 | 2.45 | 2.45 | 0 | 13.1031 | [13 empty] | 0 | 2.45
[6 empty] | 0.06 | 2.46 | 13.1786 | 333.9389 | 41.8693 | 38.9988 | 13.1786 | 10.1461 | 0 | 0.45896 | 2.46 | 2.46 | 0 | 13.1786 | [13 empty] | 0 | 2.46
[6 empty] | 0.07 | 2.47 | 13.253 | 340.4387 | 48.3413 | 38.6653 | 13.253 | 10.1586 | 0 | 0.46733 | 2.47 | 2.47 | 0 | 13.253 | [13 empty] | 0 | 2.47
[6 empty] | 0.08 | 2.48 | 13.3263 | 346.8829 | 54.7577 | 38.3317 | 13.3263 | 10.1714 | 0 | 0.47569 | 2.48 | 2.48 | 0 | 13.3263 | [13 empty] | 0 | 2.48
[6 empty] | 0.09 | 2.49 | 13.3985 | 353.2716 | 61.1185 | 37.9982 | 13.3985 | 10.1843 | 0 | 0.48401 | 2.49 | 2.49 | 0 | 13.3985 | [13 empty] | 0 | 2.49
[6 empty] | 0.1 | 2.5 | 13.4695 | 359.6046 | 67.4238 | 37.6647 | 13.4695 | 10.1974 | 0 | 0.4923 | 2.5 | 2.5 | 0 | 13.4695 | [13 empty] | 0 | 2.5
[6 empty] | 0.11 | 2.51 | 13.5395 | 365.8821 | 73.6734 | 37.3312 | 13.5395 | 10.2106 | 0 | 0.50057 | 2.51 | 2.51 | 0 | 13.5395 | [13 empty] | 0 | 2.51
[6 empty] | 0.12 | 2.52 | 13.6083 | 372.1039 | 79.8675 | 36.9976 | 13.6083 | 10.224 | 0 | 0.5088 | 2.52 | 2.52 | 0 | 13.6083 | [13 empty] | 0 | 2.52
[6 empty] | 0.13 | 2.53 | 13.6761 | 378.2702 | 86.006 | 36.6641 | 13.6761 | 10.2376 | 0 | 0.517 | 2.53 | 2.53 | 0 | 13.6761 | [13 empty] | 0 | 2.53
[6 empty] | 0.14 | 2.54 | 13.7427 | 384.3809 | 92.0889 | 36.3306 | 13.7427 | 10.2514 | 0 | 0.52516 | 2.54 | 2.54 | 0 | 13.7427 | [13 empty] | 0 | 2.54
[6 empty] | 0.15 | 2.55 | 13.8082 | 390.436 | 98.1162 | 35.997 | 13.8082 | 10.2653 | 0 | 0.53329 | 2.55 | 2.55 | 0 | 13.8082 | [13 empty] | 0 | 2.55
[6 empty] | 0.16 | 2.56 | 13.8726 | 396.4355 | 104.0879 | 35.6635 | 13.8726 | 10.2793 | 0 | 0.54138 | 2.56 | 2.56 | 0 | 13.8726 | [13 empty] | 0 | 2.56
[6 empty] | 0.17 | 2.57 | 13.9359 | 402.3794 | 110.004 | 35.33 | 13.9359 | 10.2935 | 0 | 0.54943 | 2.57 | 2.57 | 0 | 13.9359 | [13 empty] | 0 | 2.57
[6 empty] | 0.18 | 2.58 | 13.9981 | 408.2677 | 115.8645 | 34.9964 | 13.9981 | 10.3078 | 0 | 0.55744 | 2.58 | 2.58 | 0 | 13.9981 | [13 empty] | 0 | 2.58
[6 empty] | 0.19 | 2.59 | 14.0592 | 414.1004 | 121.6695 | 34.6629 | 14.0592 | 10.3222 | 0 | 0.56541 | 2.59 | 2.59 | 0 | 14.0592 | [13 empty] | 0 | 2.59
[6 empty] | 0.2 | 2.6 | 14.1191 | 419.8776 | 127.4188 | 34.3294 | 14.1191 | 10.3368 | 0 | 0.57333 | 2.6 | 2.6 | 0 | 14.1191 | [13 empty] | 0 | 2.6
[6 empty] | 0.21 | 2.61 | 14.178 | 425.5991 | 133.1126 | 33.9958 | 14.178 | 10.3514 | 0 | 0.58121 | 2.61 | 2.61 | 0 | 14.178 | [13 empty] | 0 | 2.61
[6 empty] | 0.22 | 2.62 | 14.2357 | 431.2651 | 138.7508 | 33.6623 | 14.2357 | 10.3662 | 0 | 0.58905 | 2.62 | 2.62 | 0 | 14.2357 | [13 empty] | 0 | 2.62
[6 empty] | 0.23 | 2.63 | 14.2923 | 436.8755 | 144.3333 | 33.3288 | 14.2923 | 10.3811 | 0 | 0.59683 | 2.63 | 2.63 | 0 | 14.2923 | [13 empty] | 0 | 2.63
[6 empty] | 0.24 | 2.64 | 14.3478 | 442.4303 | 149.8604 | 32.9952 | 14.3478 | 10.396 | 0 | 0.60457 | 2.64 | 2.64 | 0 | 14.3478 | [13 empty] | 0 | 2.64
[6 empty] | 0.25 | 2.65 | 14.4022 | 447.9295 | 155.3318 | 32.6617 | 14.4022 | 10.4111 | 0 | 0.61227 | 2.65 | 2.65 | 0 | 14.4022 | [13 empty] | 0 | 2.65
[6 empty] | 0.26 | 2.66 | 14.4555 | 453.3731 | 160.7476 | 32.3282 | 14.4555 | 10.4262 | 0 | 0.61991 | 2.66 | 2.66 | 0 | 14.4555 | [13 empty] | 0 | 2.66
[6 empty] | 0.27 | 2.67 | 14.5077 | 458.7612 | 166.1078 | 31.9946 | 14.5077 | 10.4415 | 0 | 0.62749 | 2.67 | 2.67 | 0 | 14.5077 | [13 empty] | 0 | 2.67
[6 empty] | 0.28 | 2.68 | 14.5587 | 464.0936 | 171.4125 | 31.6611 | 14.5587 | 10.4568 | 0 | 0.63503 | 2.68 | 2.68 | 0 | 14.5587 | [13 empty] | 0 | 2.68
[6 empty] | 0.29 | 2.69 | 14.6087 | 469.3704 | 176.6615 | 31.3276 | 14.6087 | 10.4722 | 0 | 0.64251 | 2.69 | 2.69 | 0 | 14.6087 | [13 empty] | 0 | 2.69
[6 empty] | 0.3 | 2.7 | 14.6575 | 474.5917 | 181.855 | 30.994 | 14.6575 | 10.4876 | 0 | 0.64994 | 2.7 | 2.7 | 0 | 14.6575 | [13 empty] | 0 | 2.7
[6 empty] | 0.31 | 2.71 | 14.7052 | 479.7574 | 186.9929 | 30.6605 | 14.7052 | 10.5031 | 0 | 0.65731 | 2.71 | 2.71 | 0 | 14.7052 | [13 empty] | 0 | 2.71
[6 empty] | 0.32 | 2.72 | 14.7518 | 484.8675 | 192.0752 | 30.327 | 14.7518 | 10.5187 | 0 | 0.66463 | 2.72 | 2.72 | 0 | 14.7518 | [13 empty] | 0 | 2.72
[6 empty] | 0.33 | 2.73 | 14.7973 | 489.922 | 197.1019 | 29.9935 | 14.7973 | 10.5343 | 0 | 0.67188 | 2.73 | 2.73 | 0 | 14.7973 | [13 empty] | 0 | 2.73
[6 empty] | 0.34 | 2.74 | 14.8417 | 494.9209 | 202.073 | 29.6599 | 14.8417 | 10.55 | 0 | 0.67908 | 2.74 | 2.74 | 0 | 14.8417 | [13 empty] | 0 | 2.74
[6 empty] | 0.35 | 2.75 | 14.885 | 499.8642 | 206.9885 | 29.3264 | 14.885 | 10.5657 | 0 | 0.68622 | 2.75 | 2.75 | 0 | 14.885 | [13 empty] | 0 | 2.75
[6 empty] | 0.36 | 2.76 | 14.9272 | 504.7519 | 211.8484 | 28.9929 | 14.9272 | 10.5814 | 0 | 0.6933 | 2.76 | 2.76 | 0 | 14.9272 | [13 empty] | 0 | 2.76
[6 empty] | 0.37 | 2.77 | 14.9682 | 509.5841 | 216.6528 | 28.6593 | 14.9682 | 10.5972 | 0 | 0.70031 | 2.77 | 2.77 | 0 | 14.9682 | [13 empty] | 0 | 2.77
[6 empty] | 0.38 | 2.78 | 15.0082 | 514.3606 | 221.4016 | 28.3258 | 15.0082 | 10.613 | 0 | 0.70726 | 2.78 | 2.78 | 0 | 15.0082 | [13 empty] | 0 | 2.78
[6 empty] | 0.39 | 2.79 | 15.047 | 519.0816 | 226.0947 | 27.9923 | 15.047 | 10.6288 | 0 | 0.71415 | 2.79 | 2.79 | 0 | 15.047 | [13 empty] | 0 | 2.79
[6 empty] | 0.4 | 2.8 | 15.0847 | 523.747 | 230.7323 | 27.6587 | 15.0847 | 10.6447 | 0 | 0.72097 | 2.8 | 2.8 | 0 | 15.0847 | [13 empty] | 0 | 2.8
[6 empty] | 0.41 | 2.81 | 15.1213 | 528.3568 | 235.3143 | 27.3252 | 15.1213 | 10.6606 | 0 | 0.72772 | 2.81 | 2.81 | 0 | 15.1213 | [13 empty] | 0 | 2.81
[6 empty] | 0.42 | 2.82 | 15.1568 | 532.911 | 239.8407 | 26.9917 | 15.1568 | 10.6764 | 0 | 0.73441 | 2.82 | 2.82 | 0 | 15.1568 | [13 empty] | 0 | 2.82
[6 empty] | 0.43 | 2.83 | 15.1912 | 537.4096 | 244.3115 | 26.6581 | 15.1912 | 10.6923 | 0 | 0.74103 | 2.83 | 2.83 | 0 | 15.1912 | [13 empty] | 0 | 2.83
[6 empty] | 0.44 | 2.84 | 15.2245 | 541.8526 | 248.7268 | 26.3246 | 15.2245 | 10.7082 | 0 | 0.74759 | 2.84 | 2.84 | 0 | 15.2245 | [13 empty] | 0 | 2.84
[6 empty] | 0.45 | 2.85 | 15.2566 | 546.24 | 253.0864 | 25.9911 | 15.2566 | 10.7242 | 0 | 0.75407 | 2.85 | 2.85 | 0 | 15.2566 | [13 empty] | 0 | 2.85
[6 empty] | 0.46 | 2.86 | 15.2877 | 550.5719 | 257.3904 | 25.6575 | 15.2877 | 10.7401 | 0 | 0.76048 | 2.86 | 2.86 | 0 | 15.2877 | [13 empty] | 0 | 2.86
[6 empty] | 0.47 | 2.87 | 15.3176 | 554.8481 | 261.6389 | 25.324 | 15.3176 | 10.7559 | 0 | 0.76683 | 2.87 | 2.87 | 0 | 15.3176 | [13 empty] | 0 | 2.87
[6 empty] | 0.48 | 2.88 | 15.3465 | 559.0688 | 265.8318 | 24.9905 | 15.3465 | 10.7718 | 0 | 0.7731 | 2.88 | 2.88 | 0 | 15.3465 | [13 empty] | 0 | 2.88
[6 empty] | 0.49 | 2.89 | 15.3742 | 563.2339 | 269.9691 | 24.6569 | 15.3742 | 10.7877 | 0 | 0.77929 | 2.89 | 2.89 | 0 | 15.3742 | [13 empty] | 0 | 2.89
[6 empty] | 0.5 | 2.9 | 15.4008 | 567.3434 | 274.0508 | 24.3234 | 15.4008 | 10.8036 | 0 | 0.78542 | 2.9 | 2.9 | 0 | 15.4008 | [13 empty] | 0 | 2.9
[6 empty] | 0.51 | 2.91 | 15.4263 | 571.3973 | 278.0769 | 23.9899 | 15.4263 | 10.8194 | 0 | 0.79147 | 2.91 | 2.91 | 0 | 15.4263 | [13 empty] | 0 | 2.91
[6 empty] | 0.52 | 2.92 | 15.4506 | 575.3956 | 282.0474 | 23.6563 | 15.4506 | 10.8352 | 0 | 0.79745 | 2.92 | 2.92 | 0 | 15.4506 | [13 empty] | 0 | 2.92
[6 empty] | 0.53 | 2.93 | 15.4739 | 579.3383 | 285.9623 | 23.3228 | 15.4739 | 10.851 | 0 | 0.80335 | 2.93 | 2.93 | 0 | 15.4739 | [13 empty] | 0 | 2.93
[6 empty] | 0.54 | 2.94 | 15.4961 | 583.2254 | 289.8217 | 22.9893 | 15.4961 | 10.8667 | 0 | 0.80917 | 2.94 | 2.94 | 0 | 15.4961 | [13 empty] | 0 | 2.94
[6 empty] | 0.55 | 2.95 | 15.5171 | 587.057 | 293.6254 | 22.6558 | 15.5171 | 10.8824 | 0 | 0.81492 | 2.95 | 2.95 | 0 | 15.5171 | [13 empty] | 0 | 2.95
[6 empty] | 0.56 | 2.96 | 15.537 | 590.8329 | 297.3736 | 22.3222 | 15.537 | 10.8981 | 0 | 0.82059 | 2.96 | 2.96 | 0 | 15.537 | [13 empty] | 0 | 2.96
[6 empty] | 0.57 | 2.97 | 15.5558 | 594.5533 | 301.0662 | 21.9887 | 15.5558 | 10.9137 | 0 | 0.82618 | 2.97 | 2.97 | 0 | 15.5558 | [13 empty] | 0 | 2.97
[6 empty] | 0.58 | 2.98 | 15.5736 | 598.2181 | 304.7031 | 21.6552 | 15.5736 | 10.9293 | 0 | 0.8317 | 2.98 | 2.98 | 0 | 15.5736 | [13 empty] | 0 | 2.98
[6 empty] | 0.59 | 2.99 | 15.5902 | 601.8273 | 308.2845 | 21.3216 | 15.5902 | 10.9448 | 0 | 0.83713 | 2.99 | 2.99 | 0 | 15.5902 | [13 empty] | 0 | 2.99
[6 empty] | 0.6 | 3 | 15.6056 | 605.3809 | 311.8104 | 20.9881 | 15.6056 | 10.9603 | 0 | 0.84249 | 3 | 3 | 0 | 15.6056 | [13 empty] | 0 | 3
[6 empty] | 0.61 | 3.01 | 15.62 | 608.8789 | 315.2806 | 20.6546 | 15.62 | 10.9758 | 0 | 0.84776 | 3.01 | 3.01 | 0 | 15.62 | [14 empty] | 3.01
[6 empty] | 0.62 | 3.02 | 15.6333 | 612.3213 | 318.6952 | 20.321 | 15.6333 | 10.9911 | 0 | 0.85295 | 3.02 | 3.02 | 0 | 15.6333 | [14 empty] | 3.02
[6 empty] | 0.63 | 3.03 | 15.6454 | 615.7082 | 322.0542 | 19.9875 | 15.6454 | 11.0064 | 0 | 0.85807 | 3.03 | 3.03 | 0 | 15.6454 | [14 empty] | 3.03
[6 empty] | 0.64 | 3.04 | 15.6564 | 619.0394 | 325.3577 | 19.654 | 15.6564 | 11.0217 | 0 | 0.8631 | 3.04 | 3.04 | 0 | 15.6564 | [14 empty] | 3.04
[6 empty] | 0.65 | 3.05 | 15.6664 | 622.3151 | 328.6056 | 19.3204 | 15.6664 | 11.0369 | 0 | 0.86805 | 3.05 | 3.05 | 0 | 15.6664 | [14 empty] | 3.05
[6 empty] | 0.66 | 3.06 | 15.6752 | 625.5352 | 331.7979 | 18.9869 | 15.6752 | 11.052 | 0 | 0.87291 | 3.06 | 3.06 | 0 | 15.6752 | [14 empty] | 3.06
[6 empty] | 0.67 | 3.07 | 15.6829 | 628.6996 | 334.9345 | 18.6534 | 15.6829 | 11.067 | 0 | 0.8777 | 3.07 | 3.07 | 0 | 15.6829 | [14 empty] | 3.07
[6 empty] | 0.68 | 3.08 | 15.6895 | 631.8085 | 338.0156 | 18.3198 | 15.6895 | 11.082 | 0 | 0.8824 | 3.08 | 3.08 | 0 | 15.6895 | [14 empty] | 3.08
[6 empty] | 0.69 | 3.09 | 15.695 | 634.8618 | 341.0412 | 17.9863 | 15.695 | 11.0969 | 0 | 0.88701 | 3.09 | 3.09 | 0 | 15.695 | [14 empty] | 3.09
[6 empty] | 0.7 | 3.1 | 15.6993 | 637.8596 | 344.0111 | 17.6528 | 15.6993 | 11.1117 | 0 | 0.89154 | 3.1 | 3.1 | 0 | 15.6993 | [14 empty] | 3.1
[6 empty] | 0.71 | 3.11 | 15.7026 | 640.8017 | 346.9254 | 17.3192 | 15.7026 | 11.1264 | 0 | 0.89599 | 3.11 | 3.11 | 0 | 15.7026 | [14 empty] | 3.11
[6 empty] | 0.72 | 3.12 | 15.7047 | 643.6882 | 349.7842 | 16.9857 | 15.7047 | 11.141 | 0 | 0.90035 | 3.12 | 3.12 | 0 | 15.7047 | [14 empty] | 3.12
[6 empty] | 0.73 | 3.13 | 15.7057 | 646.5192 | 352.5873 | 16.6522 | 15.7057 | 11.1555 | 0 | 0.90463 | 3.13 | 3.13 | 0 | 15.7057 | [14 empty] | 3.13
[6 empty] | 0.74 | 3.14 | 15.7057 | 649.2946 | 355.3349 | 16.3187 | 15.7057 | 11.17 | 0 | 0.90882 | 3.14 | 3.14 | 0 | 15.7057 | [14 empty] | 3.14
[6 empty] | 0.75 | 3.15 | 15.7045 | 652.0143 | 358.0269 | 15.9851 | 15.7045 | 11.1843 | 0 | 0.91292 | 3.15 | 3.15 | 0 | 15.7045 | [14 empty] | 3.15
[6 empty] | 0.76 | 3.16 | 15.7022 | 654.6785 | 360.6633 | 15.6516 | 15.7022 | 11.1986 | 0 | 0.91694 | 3.16 | 3.16 | 0 | 15.7022 | [14 empty] | 3.16
[6 empty] | 0.77 | 3.17 | 15.6988 | 657.2871 | 363.2441 | 15.3181 | 15.6988 | 11.2127 | 0 | 0.92087 | 3.17 | 3.17 | 0 | 15.6988 | [14 empty] | 3.17
[6 empty] | 0.78 | 3.18 | 15.6942 | 659.8401 | 365.7693 | 14.9845 | 15.6942 | 11.2268 | 0 | 0.92472 | 3.18 | 3.18 | 0 | 15.6942 | [14 empty] | 3.18
[6 empty] | 0.79 | 3.19 | 15.6886 | 662.3375 | 368.2389 | 14.651 | 15.6886 | 11.2407 | 0 | 0.92847 | 3.19 | 3.19 | 0 | 15.6886 | [14 empty] | 3.19
[6 empty] | 0.8 | 3.2 | 15.6818 | 664.7794 | 370.6529 | 14.3175 | 15.6818 | 11.2546 | 0 | 0.93215 | 3.2 | 3.2 | 0 | 15.6818 | [14 empty] | 3.2
[6 empty] | 0.81 | 3.21 | 15.674 | 667.1656 | 373.0114 | 13.9839 | 15.674 | 11.2683 | 0 | 0.93573 | 3.21 | 3.21 | 0 | 15.674 | [14 empty] | 3.21
[6 empty] | 0.82 | 3.22 | 15.665 | 669.4963 | 375.3143 | 13.6504 | 15.665 | 11.2819 | 0 | 0.93923 | 3.22 | 3.22 | 0 | 15.665 | [14 empty] | 3.22
[6 empty] | 0.83 | 3.23 | 15.6549 | 671.7713 | 377.5615 | 13.3169 | 15.6549 | 11.2954 | 0 | 0.94263 | 3.23 | 3.23 | 0 | 15.6549 | [14 empty] | 3.23
[6 empty] | 0.84 | 3.24 | 15.6437 | 673.9908 | 379.7532 | 12.9833 | 15.6437 | 11.3088 | 0 | 0.94595 | 3.24 | 3.24 | 0 | 15.6437 | [14 empty] | 3.24
[6 empty] | 0.85 | 3.25 | 15.6314 | 676.1547 | 381.8893 | 12.6498 | 15.6314 | 11.322 | 0 | 0.94919 | 3.25 | 3.25 | 0 | 15.6314 | [14 empty] | 3.25
[6 empty] | 0.86 | 3.26 | 15.618 | 678.263 | 383.9698 | 12.3163 | 15.618 | 11.3352 | 0 | 0.95233 | 3.26 | 3.26 | 0 | 15.618 | [14 empty] | 3.26
[6 empty] | 0.87 | 3.27 | 15.6034 | 680.3157 | 385.9947 | 11.9827 | 15.6034 | 11.3482 | 0 | 0.95539 | 3.27 | 3.27 | 0 | 15.6034 | [14 empty] | 3.27
[6 empty] | 0.88 | 3.28 | 15.5878 | 682.3128 | 387.9641 | 11.6492 | 15.5878 | 11.3611 | 0 | 0.95836 | 3.28 | 3.28 | 0 | 15.5878 | [14 empty] | 3.28
[6 empty] | 0.89 | 3.29 | 15.571 | 684.2544 | 389.8778 | 11.3157 | 15.571 | 11.3738 | 0 | 0.96124 | 3.29 | 3.29 | 0 | 15.571 | [14 empty] | 3.29
[6 empty] | 0.9 | 3.3 | 15.5532 | 686.1403 | 391.7359 | 10.9821 | 15.5532 | 11.3864 | 0 | 0.96403 | 3.3 | 3.3 | 0 | 15.5532 | [14 empty] | 3.3
[6 empty] | 0.91 | 3.31 | 15.5342 | 687.9707 | 393.5385 | 10.6486 | 15.5342 | 11.3989 | 0 | 0.96674 | 3.31 | 3.31 | 0 | 15.5342 | [14 empty] | 3.31
[6 empty] | 0.92 | 3.32 | 15.5141 | 689.7454 | 395.2855 | 10.3151 | 15.5141 | 11.4113 | 0 | 0.96935 | 3.32 | 3.32 | 0 | 15.5141 | [14 empty] | 3.32
[6 empty] | 0.93 | 3.33 | 15.4929 | 691.4646 | 396.9769 | 9.9815 | 15.4929 | 11.4235 | 0 | 0.97188 | 3.33 | 3.33 | 0 | 15.4929 | [14 empty] | 3.33
[6 empty] | 0.94 | 3.34 | 15.4706 | 693.1282 | 398.6127 | 9.648 | 15.4706 | 11.4356 | 0 | 0.97432 | 3.34 | 3.34 | 0 | 15.4706 | [14 empty] | 3.34
[6 empty] | 0.95 | 3.35 | 15.4471 | 694.7362 | 400.1929 | 9.3145 | 15.4471 | 11.4475 | 0 | 0.97667 | 3.35 | 3.35 | 0 | 15.4471 | [14 empty] | 3.35
[6 empty] | 0.96 | 3.36 | 15.4226 | 696.2886 | 401.7175 | 8.981 | 15.4226 | 11.4593 | 0 | 0.97894 | 3.36 | 3.36 | 0 | 15.4226 | [14 empty] | 3.36
[6 empty] | 0.97 | 3.37 | 15.3969 | 697.7855 | 403.1865 | 8.6474 | 15.3969 | 11.471 | 0 | 0.98112 | 3.37 | 3.37 | 0 | 15.3969 | [14 empty] | 3.37
[6 empty] | 0.98 | 3.38 | 15.3702 | 699.2267 | 404.6 | 8.3139 | 15.3702 | 11.4825 | 0 | 0.98321 | 3.38 | 3.38 | 0 | 15.3702 | [14 empty] | 3.38
[6 empty] | 0.99 | 3.39 | 15.3423 | 700.6123 | 405.9578 | 7.9804 | 15.3423 | 11.4938 | 0 | 0.98521 | 3.39 | 3.39 | 0 | 15.3423 | [14 empty] | 3.39
[6 empty] | 1 | 3.4 | 15.3133 | 701.9424 | 407.2601 | 7.6468 | 15.3133 | 11.505 | 0 | 0.98712 | 3.4 | 3.4 | 0 | 15.3133 | [14 empty] | 3.4
[6 empty] | 1.01 | 3.41 | 15.2832 | 703.2169 | 408.5068 | 7.3133 | 15.2832 | 11.5161 | 0 | 0.98895 | 3.41 | 3.41 | 0 | 15.2832 | [14 empty] | 3.41
[6 empty] | 1.02 | 3.42 | 15.252 | 704.4358 | 409.6979 | 6.9798 | 15.252 | 11.527 | 0 | 0.99069 | 3.42 | 3.42 | 0 | 15.252 | [14 empty] | 3.42
[6 empty] | 1.03 | 3.43 | 15.2197 | 705.599 | 410.8333 | 6.6462 | 15.2197 | 11.5377 | 0 | 0.99234 | 3.43 | 3.43 | 0 | 15.2197 | [14 empty] | 3.43
[6 empty] | 1.04 | 3.44 | 15.1863 | 706.7068 | 411.9133 | 6.3127 | 15.1863 | 11.5483 | 0 | 0.99391 | 3.44 | 3.44 | 0 | 15.1863 | [14 empty] | 3.44
[6 empty] | 1.05 | 3.45 | 15.1517 | 707.7589 | 412.9376 | 5.9792 | 15.1517 | 11.5587 | 0 | 0.99539 | 3.45 | 3.45 | 0 | 15.1517 | [14 empty] | 3.45
[6 empty] | 1.06 | 3.46 | 15.116 | 708.7554 | 413.9063 | 5.6456 | 15.116 | 11.569 | 0 | 0.99679 | 3.46 | 3.46 | 0 | 15.116 | [14 empty] | 3.46
[6 empty] | 1.07 | 3.47 | 15.0793 | 709.6963 | 414.8195 | 5.3121 | 15.0793 | 11.579 | 0 | 0.9981 | 3.47 | 3.47 | 0 | 15.0793 | [14 empty] | 3.47
[6 empty] | 1.08 | 3.48 | 15.0414 | 710.5817 | 415.677 | 4.9786 | 15.0414 | 11.589 | 0 | 0.99932 | 3.48 | 3.48 | 0 | 15.0414 | [14 empty] | 3.48
[6 empty] | 1.09 | 3.49 | 15.0024 | 711.4114 | 416.479 | 4.645 | 15.0024 | 11.5987 | 0 | 1.0005 | 3.49 | 3.49 | 0 | 15.0024 | [14 empty] | 3.49
[6 empty] | 1.1 | 3.5 | 14.9623 | 712.1856 | 417.2254 | 4.3115 | 14.9623 | 11.6083 | 0 | 1.0015 | 3.5 | 3.5 | 0 | 14.9623 | [14 empty] | 3.5
[6 empty] | 1.11 | 3.51 | 14.9211 | 712.9042 | 417.9162 | 3.978 | 14.9211 | 11.6177 | 0 | 1.0025 | 3.51 | 3.51 | 0 | 14.9211 | [14 empty] | 3.51
[6 empty] | 1.12 | 3.52 | 14.8788 | 713.5672 | 418.5514 | 3.6444 | 14.8788 | 11.6269 | 0 | 1.0034 | 3.52 | 3.52 | 0 | 14.8788 | [14 empty] | 3.52
[6 empty] | 1.13 | 3.53 | 14.8353 | 714.1746 | 419.131 | 3.3109 | 14.8353 | 11.636 | 0 | 1.0042 | 3.53 | 3.53 | 0 | 14.8353 | [14 empty] | 3.53
[6 empty] | 1.14 | 3.54 | 14.7908 | 714.7264 | 419.655 | 2.9774 | 14.7908 | 11.6449 | 0 | 1.0049 | 3.54 | 3.54 | 0 | 14.7908 | [14 empty] | 3.54
[6 empty] | 1.15 | 3.55 | 14.7451 | 715.2227 | 420.1234 | 2.6438 | 14.7451 | 11.6536 | 0 | 1.0055 | 3.55 | 3.55 | 0 | 14.7451 | [14 empty] | 3.55
[6 empty] | 1.16 | 3.56 | 14.6983 | 715.6633 | 420.5363 | 2.3103 | 14.6983 | 11.6621 | 0 | 1.0061 | 3.56 | 3.56 | 0 | 14.6983 | [14 empty] | 3.56
[6 empty] | 1.17 | 3.57 | 14.6504 | 716.0484 | 420.8935 | 1.9768 | 14.6504 | 11.6705 | 0 | 1.0066 | 3.57 | 3.57 | 0 | 14.6504 | [14 empty] | 3.57
[6 empty] | 1.18 | 3.58 | 14.6014 | 716.3778 | 421.1952 | 1.6433 | 14.6014 | 11.6786 | 0 | 1.007 | 3.58 | 3.58 | 0 | 14.6014 | [14 empty] | 3.58
[6 empty] | 1.19 | 3.59 | 14.5513 | 716.6517 | 421.4413 | 1.3097 | 14.5513 | 11.6866 | 0 | 1.0073 | 3.59 | 3.59 | 0 | 14.5513 | [14 empty] | 3.59
[6 empty] | 1.2 | 3.6 | 14.5001 | 716.87 | 421.6318 | 0.97619 | 14.5001 | 11.6944 | 0 | 1.0075 | 3.6 | 3.6 | 0 | 14.5001 | [14 empty] | 3.6
[6 empty] | 1.21 | 3.61 | 14.4478 | 717.0327 | 421.7667 | 0.64266 | 14.4478 | 11.702 | 0 | 1.0076 | 3.61 | 3.61 | 0 | 14.4478 | [14 empty] | 3.61
[6 empty] | 1.22 | 3.62 | 14.3943 | 717.1398 | 421.846 | 0.30913 | 14.3943 | 11.7094 | 0 | 1.0077 | 3.62 | 3.62 | 0 | 14.3943 | [14 empty] | 3.62
[6 empty] | 1.23 | 3.63 | 5 | 717.1913 | 417.1998 | -56.063 | 5 | 11.691 | 0 | 1.0031 | 3.63 | 3.63 | 0 | 5 | [14 empty] | 3.63
[6 empty] | 1.24 | 3.64 | 5 | 707.8475 | 407.9495 | -54.9417 | 5 | 11.6727 | 0 | 0.99865 | 3.64 | 3.64 | 0 | 5 | [14 empty] | 3.64
[6 empty] | 1.25 | 3.65 | 5 | 698.6905 | 398.8841 | -53.8429 | 5 | 11.6544 | 0 | 0.99425 | 3.65 | 3.65 | 0 | 5 | [14 empty] | 3.65
[6 empty] | 1.26 | 3.66 | 5 | 689.7167 | 390 | -52.766 | 5 | 11.6363 | 0 | 0.98993 | 3.66 | 3.66 | 0 | 5 | [14 empty] | 3.66
[6 empty] | 1.27 | 3.67 | 5 | 680.9224 | 381.2936 | -51.7107 | 5 | 11.6183 | 0 | 0.9857 | 3.67 | 3.67 | 0 | 5 | [14 empty] | 3.67
[6 empty] | 1.28 | 3.68 | 5 | 672.3039 | 372.7614 | -50.6765 | 5 | 11.6003 | 0 | 0.98156 | 3.68 | 3.68 | 0 | 5 | [14 empty] | 3.68
[6 empty] | 1.29 | 3.69 | 5 | 663.8579 | 364.3997 | -49.6629 | 5 | 11.5825 | 0 | 0.9775 | 3.69 | 3.69 | 0 | 5 | [14 empty] | 3.69
[6 empty] | 1.3 | 3.7 | 5 | 655.5807 | 356.2053 | -48.6697 | 5 | 11.5648 | 0 | 0.97352 | 3.7 | 3.7 | 0 | 5 | [14 empty] | 3.7
[6 empty] | 1.31 | 3.71 | 5 | 647.4691 | 348.1748 | -47.6963 | 5 | 11.5471 | 0 | 0.96962 | 3.71 | 3.71 | 0 | 5 | [14 empty] | 3.71
[6 empty] | 1.32 | 3.72 | 5 | 639.5197 | 340.305 | -46.7424 | 5 | 11.5296 | 0 | 0.9658 | 3.72 | 3.72 | 0 | 5 | [14 empty] | 3.72
[6 empty] | 1.33 | 3.73 | 5 | 631.7293 | 332.5925 | -45.8075 | 5 | 11.5121 | 0 | 0.96206 | 3.73 | 3.73 | 0 | 5 | [14 empty] | 3.73
[6 empty] | 1.34 | 3.74 | 5 | 624.0947 | 325.0342 | -44.8914 | 5 | 11.4947 | 0 | 0.95839 | 3.74 | 3.74 | 0 | 5 | [14 empty] | 3.74
[6 empty] | 1.35 | 3.75 | 5 | 616.6128 | 317.6272 | -43.9935 | 5 | 11.4775 | 0 | 0.95479 | 3.75 | 3.75 | 0 | 5 | [14 empty] | 3.75
[6 empty] | 1.36 | 3.76 | 5 | 609.2806 | 310.3682 | -43.1137 | 5 | 11.4603 | 0 | 0.95127 | 3.76 | 3.76 | 0 | 5 | [14 empty] | 3.76
[6 empty] | 1.37 | 3.77 | 5 | 602.095 | 303.2545 | -42.2514 | 5 | 11.4432 | 0 | 0.94781 | 3.77 | 3.77 | 0 | 5 | [14 empty] | 3.77
[6 empty] | 1.38 | 3.78 | 5 | 595.0531 | 296.283 | -41.4064 | 5 | 11.4262 | 0 | 0.94443 | 3.78 | 3.78 | 0 | 5 | [14 empty] | 3.78
[6 empty] | 1.39 | 3.79 | 5 | 588.152 | 289.4509 | -40.5782 | 5 | 11.4093 | 0 | 0.94111 | 3.79 | 3.79 | 0 | 5 | [14 empty] | 3.79
[6 empty] | 1.4 | 3.8 | 5 | 581.389 | 282.7555 | -39.7667 | 5 | 11.3925 | 0 | 0.93786 | 3.8 | 3.8 | 0 | 5 | [14 empty] | 3.8
[6 empty] | 1.41 | 3.81 | 5 | 574.7612 | 276.194 | -38.9713 | 5 | 11.3757 | 0 | 0.93467 | 3.81 | 3.81 | 0 | 5 | [14 empty] | 3.81
[6 empty] | 1.42 | 3.82 | 5 | 568.266 | 269.7638 | -38.1919 | 5 | 11.3591 | 0 | 0.93155 | 3.82 | 3.82 | 0 | 5 | [14 empty] | 3.82
[6 empty] | 1.43 | 3.83 | 5 | 561.9006 | 263.4621 | -37.4281 | 5 | 11.3425 | 0 | 0.92849 | 3.83 | 3.83 | 0 | 5 | [14 empty] | 3.83
[6 empty] | 1.44 | 3.84 | 5 | 555.6626 | 257.2865 | -36.6795 | 5 | 11.326 | 0 | 0.92549 | 3.84 | 3.84 | 0 | 5 | [14 empty] | 3.84
[6 empty] | 1.45 | 3.85 | 5 | 549.5494 | 251.2343 | -35.9459 | 5 | 11.3097 | 0 | 0.92255 | 3.85 | 3.85 | 0 | 5 | [14 empty] | 3.85
[6 empty] | 1.46 | 3.86 | 5 | 543.5584 | 245.3033 | -35.227 | 5 | 11.2934 | 0 | 0.91968 | 3.86 | 3.86 | 0 | 5 | [14 empty] | 3.86
[6 empty] | 1.47 | 3.87 | 5 | 537.6872 | 239.4908 | -34.5225 | 5 | 11.2771 | 0 | 0.91685 | 3.87 | 3.87 | 0 | 5 | [14 empty] | 3.87
[6 empty] | 1.48 | 3.88 | 5 | 531.9335 | 233.7946 | -33.832 | 5 | 11.261 | 0 | 0.91409 | 3.88 | 3.88 | 0 | 5 | [14 empty] | 3.88
[6 empty] | 1.49 | 3.89 | 5 | 526.2948 | 228.2123 | -33.1554 | 5 | 11.2449 | 0 | 0.91138 | 3.89 | 3.89 | 0 | 5 | [14 empty] | 3.89
[6 empty] | 1.5 | 3.9 | 5 | 520.7689 | 222.7417 | -32.4923 | 5 | 11.229 | 0 | 0.90872 | 3.9 | 3.9 | 0 | 5 | [14 empty] | 3.9
[6 empty] | 1.51 | 3.91 | 5 | 515.3535 | 217.3804 | -31.8424 | 5 | 11.2131 | 0 | 0.90612 | 3.91 | 3.91 | 0 | 5 | [14 empty] | 3.91
[6 empty] | 1.52 | 3.92 | 5 | 510.0465 | 212.1264 | -31.2056 | 5 | 11.1973 | 0 | 0.90357 | 3.92 | 3.92 | 0 | 5 | [14 empty] | 3.92
[6 empty] | 1.53 | 3.93 | 5 | 504.8455 | 206.9775 | -30.5815 | 5 | 11.1815 | 0 | 0.90107 | 3.93 | 3.93 | 0 | 5 | [14 empty] | 3.93
[6 empty] | 1.54 | 3.94 | 5 | 499.7486 | 201.9316 | -29.9698 | 5 | 11.1659 | 0 | 0.89862 | 3.94 | 3.94 | 0 | 5 | [14 empty] | 3.94
[6 empty] | 1.55 | 3.95 | 5 | 494.7537 | 196.9866 | -29.3704 | 5 | 11.1503 | 0 | 0.89622 | 3.95 | 3.95 | 0 | 5 | [14 empty] | 3.95
[6 empty] | 1.56 | 3.96 | 5 | 489.8586 | 192.1404 | -28.783 | 5 | 11.1348 | 0 | 0.89386 | 3.96 | 3.96 | 0 | 5 | [14 empty] | 3.96
[6 empty] | 1.57 | 3.97 | 5 | 485.0614 | 187.3912 | -28.2074 | 5 | 11.1194 | 0 | 0.89156 | 3.97 | 3.97 | 0 | 5 | [14 empty] | 3.97
[6 empty] | 1.58 | 3.98 | 5 | 480.3602 | 182.737 | -27.6432 | 5 | 11.1041 | 0 | 0.8893 | 3.98 | 3.98 | 0 | 5 | [14 empty] | 3.98
[6 empty] | 1.59 | 3.99 | 5 | 475.753 | 178.1759 | -27.0904 | 5 | 11.0888 | 0 | 0.88708 | 3.99 | 3.99 | 0 | 5 | [14 empty] | 3.99
[6 empty] | 1.6 | 4 | 5 | 471.2379 | 173.706 | -26.5485 | 5 | 11.0736 | 0 | 0.88491 | 4 | 4 | 0 | 5 | [14 empty] | 4
[6 empty] | 1.61 | 4.01 | 5 | 466.8132 | 169.3255 | -26.0176 | 5 | 11.0585 | 0 | 0.88278 | 4.01 | 4.01 | 0 | 5 | [14 empty] | 4.01
[6 empty] | 1.62 | 4.02 | 5 | 462.4769 | 165.0326 | -25.4972 | 5 | 11.0435 | 0 | 0.8807 | 4.02 | 4.02 | 0 | 5 | [14 empty] | 4.02
[6 empty] | 1.63 | 4.03 | 5 | 458.2274 | 160.8255 | -24.9873 | 5 | 11.0285 | 0 | 0.87866 | 4.03 | 4.03 | 0 | 5 | [14 empty] | 4.03
[6 empty] | 1.64 | 4.04 | 5 | 454.0628 | 156.7026 | -24.4875 | 5 | 11.0136 | 0 | 0.87666 | 4.04 | 4.04 | 0 | 5 | [14 empty] | 4.04
[6 empty] | 1.65 | 4.05 | 5 | 449.9816 | 152.6622 | -23.9978 | 5 | 10.9988 | 0 | 0.87469 | 4.05 | 4.05 | 0 | 5 | [14 empty] | 4.05
[6 empty] | 1.66 | 4.06 | 5 | 445.9819 | 148.7026 | -23.5178 | 5 | 10.9841 | 0 | 0.87277 | 4.06 | 4.06 | 0 | 5 | [14 empty] | 4.06
[6 empty] | 1.67 | 4.07 | 5 | 442.0623 | 144.8221 | -23.0475 | 5 | 10.9694 | 0 | 0.87089 | 4.07 | 4.07 | 0 | 5 | [14 empty] | 4.07
[6 empty] | 1.68 | 4.08 | 5 | 438.221 | 141.0193 | -22.5865 | 5 | 10.9548 | 0 | 0.86904 | 4.08 | 4.08 | 0 | 5 | [14 empty] | 4.08
[6 empty] | 1.69 | 4.09 | 5 | 434.4566 | 137.2925 | -22.1348 | 5 | 10.9403 | 0 | 0.86723 | 4.09 | 4.09 | 0 | 5 | [14 empty] | 4.09
[6 empty] | 1.7 | 4.1 | 5 | 430.7675 | 133.6403 | -21.6921 | 5 | 10.9259 | 0 | 0.86546 | 4.1 | 4.1 | 0 | 5 | [14 empty] | 4.1
[6 empty] | 1.71 | 4.11 | 5 | 427.1521 | 130.0611 | -21.2583 | 5 | 10.9115 | 0 | 0.86372 | 4.11 | 4.11 | 0 | 5 | [14 empty] | 4.11
[6 empty] | 1.72 | 4.12 | 5 | 423.6091 | 126.5534 | -20.8331 | 5 | 10.8972 | 0 | 0.86202 | 4.12 | 4.12 | 0 | 5 | [14 empty] | 4.12
[6 empty] | 1.73 | 4.13 | 5 | 420.1369 | 123.116 | -20.4164 | 5 | 10.8829 | 0 | 0.86035 | 4.13 | 4.13 | 0 | 5 | [14 empty] | 4.13
[6 empty] | 1.74 | 4.14 | 5 | 416.7342 | 119.7473 | -20.0081 | 5 | 10.8687 | 0 | 0.85871 | 4.14 | 4.14 | 0 | 5 | [14 empty] | 4.14
[6 empty] | 1.75 | 4.15 | 5 | 413.3995 | 116.4459 | -19.6079 | 5 | 10.8546 | 0 | 0.85711 | 4.15 | 4.15 | 0 | 5 | [14 empty] | 4.15
[6 empty] | 1.76 | 4.16 | 5 | 410.1315 | 113.2106 | -19.2158 | 5 | 10.8406 | 0 | 0.85554 | 4.16 | 4.16 | 0 | 5 | [14 empty] | 4.16
[6 empty] | 1.77 | 4.17 | 5 | 406.9289 | 110.04 | -18.8315 | 5 | 10.8266 | 0 | 0.854 | 4.17 | 4.17 | 0 | 5 | [14 empty] | 4.17
[6 empty] | 1.78 | 4.18 | 5 | 403.7903 | 106.9328 | -18.4548 | 5 | 10.8127 | 0 | 0.85249 | 4.18 | 4.18 | 0 | 5 | [14 empty] | 4.18
[6 empty] | 1.79 | 4.19 | 5 | 400.7145 | 103.8878 | -18.0857 | 5 | 10.7989 | 0 | 0.85101 | 4.19 | 4.19 | 0 | 5 | [14 empty] | 4.19
[6 empty] | 1.8 | 4.2 | 5 | 397.7002 | 100.9036 | -17.724 | 5 | 10.7851 | 0 | 0.84956 | 4.2 | 4.2 | 0 | 5 | [14 empty] | 4.2
[6 empty] | 1.81 | 4.21 | 5 | 394.7462 | 97.9792 | -17.3695 | 5 | 10.7714 | 0 | 0.84814 | 4.21 | 4.21 | 0 | 5 | [14 empty] | 4.21
[6 empty] | 1.82 | 4.22 | 5 | 391.8513 | 95.1132 | -17.0222 | 5 | 10.7577 | 0 | 0.84675 | 4.22 | 4.22 | 0 | 5 | [14 empty] | 4.22
[6 empty] | 1.83 | 4.23 | 5 | 389.0142 | 92.3045 | -16.6817 | 5 | 10.7442 | 0 | 0.84539 | 4.23 | 4.23 | 0 | 5 | [14 empty] | 4.23
[6 empty] | 1.84 | 4.24 | 5 | 386.234 | 89.5521 | -16.3481 | 5 | 10.7307 | 0 | 0.84405 | 4.24 | 4.24 | 0 | 5 | [14 empty] | 4.24
[6 empty] | 1.85 | 4.25 | 5 | 383.5093 | 86.8546 | -16.0211 | 5 | 10.7172 | 0 | 0.84274 | 4.25 | 4.25 | 0 | 5 | [14 empty] | 4.25
[6 empty] | 1.86 | 4.26 | 5 | 380.8391 | 84.2112 | -15.7007 | 5 | 10.7038 | 0 | 0.84146 | 4.26 | 4.26 | 0 | 5 | [14 empty] | 4.26
[6 empty] | 1.87 | 4.27 | 5 | 378.2223 | 81.6205 | -15.3867 | 5 | 10.6905 | 0 | 0.8402 | 4.27 | 4.27 | 0 | 5 | [14 empty] | 4.27
[6 empty] | 1.88 | 4.28 | 5 | 375.6579 | 79.0817 | -15.0789 | 5 | 10.6772 | 0 | 0.83897 | 4.28 | 4.28 | 0 | 5 | [14 empty] | 4.28
[6 empty] | 1.89 | 4.29 | 5 | 373.1447 | 76.5937 | -14.7774 | 5 | 10.664 | 0 | 0.83776 | 4.29 | 4.29 | 0 | 5 | [14 empty] | 4.29
[6 empty] | 1.9 | 4.3 | 5 | 370.6818 | 74.1554 | -14.4818 | 5 | 10.6509 | 0 | 0.83658 | 4.3 | 4.3 | 0 | 5 | [14 empty] | 4.3
[6 empty] | 1.91 | 4.31 | 5 | 368.2682 | 71.7659 | -14.1922 | 5 | 10.6378 | 0 | 0.83541 | 4.31 | 4.31 | 0 | 5 | [14 empty] | 4.31
[6 empty] | 1.92 | 4.32 | 5 | 365.9028 | 69.4242 | -13.9083 | 5 | 10.6248 | 0 | 0.83428 | 4.32 | 4.32 | 0 | 5 | [14 empty] | 4.32
[6 empty] | 1.93 | 4.33 | 5 | 363.5848 | 67.1294 | -13.6302 | 5 | 10.6118 | 0 | 0.83316 | 4.33 | 4.33 | 0 | 5 | [14 empty] | 4.33
[6 empty] | 1.94 | 4.34 | 5 | 361.3131 | 64.8804 | -13.3576 | 5 | 10.5989 | 0 | 0.83207 | 4.34 | 4.34 | 0 | 5 | [14 empty] | 4.34
[6 empty] | 1.95 | 4.35 | 5 | 359.0868 | 62.6764 | -13.0904 | 5 | 10.5861 | 0 | 0.831 | 4.35 | 4.35 | 0 | 5 | [14 empty] | 4.35
[6 empty] | 1.96 | 4.36 | 5 | 356.9051 | 60.5165 | -12.8286 | 5 | 10.5733 | 0 | 0.82995 | 4.36 | 4.36 | 0 | 5 | [14 empty] | 4.36
[6 empty] | 1.97 | 4.37 | 5 | 354.767 | 58.3997 | -12.572 | 5 | 10.5606 | 0 | 0.82892 | 4.37 | 4.37 | 0 | 5 | [14 empty] | 4.37
[6 empty] | 1.98 | 4.38 | 5 | 352.6716 | 56.3254 | -12.3206 | 5 | 10.5479 | 0 | 0.82792 | 4.38 | 4.38 | 0 | 5 | [14 empty] | 4.38
[6 empty] | 1.99 | 4.39 | 5 | 350.6182 | 54.2925 | -12.0742 | 5 | 10.5353 | 0 | 0.82693 | 4.39 | 4.39 | 0 | 5 | [14 empty] | 4.39
[6 empty] | 2 | 4.4 | 5 | 348.6058 | 52.3002 | -11.8327 | 5 | 10.5227 | 0 | 0.82596 | 4.4 | 4.4 | 0 | 5 | [14 empty] | 4.4
[6 empty] | 2.01 | 4.41 | 5 | 346.6337 | 50.3478 | -11.596 | 5 | 10.5102 | 0 | 0.82502 | 4.41 | 4.41 | 0 | 5 | [14 empty] | 4.41
[6 empty] | 2.02 | 4.42 | 5 | 344.701 | 48.4345 | -11.3641 | 5 | 10.4978 | 0 | 0.82409 | 4.42 | 4.42 | 0 | 5 | [14 empty] | 4.42
[6 empty] | 2.03 | 4.43 | 5 | 342.807 | 46.5594 | -11.1368 | 5 | 10.4854 | 0 | 0.82318 | 4.43 | 4.43 | 0 | 5 | [14 empty] | 4.43
[6 empty] | 2.04 | 4.44 | 5 | 340.9509 | 44.7218 | -10.9141 | 5 | 10.4731 | 0 | 0.82228 | 4.44 | 4.44 | 0 | 5 | [14 empty] | 4.44
[6 empty] | 2.05 | 4.45 | 5 | 339.1319 | 42.921 | -10.6958 | 5 | 10.4608 | 0 | 0.82141 | 4.45 | 4.45 | 0 | 5 | [14 empty] | 4.45
[6 empty] | 2.06 | 4.46 | 5 | 337.3492 | 41.1562 | -10.4819 | 5 | 10.4486 | 0 | 0.82055 | 4.46 | 4.46 | 0 | 5 | [14 empty] | 4.46
[6 empty] | 2.07 | 4.47 | 5 | 335.6022 | 39.4267 | -10.2723 | 5 | 10.4364 | 0 | 0.81971 | 4.47 | 4.47 | 0 | 5 | [14 empty] | 4.47
[6 empty] | 2.08 | 4.48 | 5 | 333.8902 | 37.7317 | -10.0668 | 5 | 10.4243 | 0 | 0.81889 | 4.48 | 4.48 | 0 | 5 | [14 empty] | 4.48
[6 empty] | 2.09 | 4.49 | 5 | 332.2124 | 36.0707 | -9.8655 | 5 | 10.4123 | 0 | 0.81808 | 4.49 | 4.49 | 0 | 5 | [14 empty] | 4.49
[6 empty] | 2.1 | 4.5 | 5 | 330.5681 | 34.4429 | -9.6682 | 5 | 10.4003 | 0 | 0.81729 | 4.5 | 4.5 | 0 | 5 | [14 empty] | 4.5
[6 empty] | 2.11 | 4.51 | 5 | 328.9568 | 32.8477 | -9.4748 | 5 | 10.3883 | 0 | 0.81652 | 4.51 | 4.51 | 0 | 5 | [14 empty] | 4.51
[6 empty] | 2.12 | 4.52 | 5 | 327.3776 | 31.2843 | -9.2853 | 5 | 10.3764 | 0 | 0.81576 | 4.52 | 4.52 | 0 | 5 | [14 empty] | 4.52
[6 empty] | 2.13 | 4.53 | 5 | 325.8301 | 29.7522 | -9.0996 | 5 | 10.3646 | 0 | 0.81502 | 4.53 | 4.53 | 0 | 5 | [14 empty] | 4.53
[6 empty] | 2.14 | 4.54 | 5 | 324.3135 | 28.2508 | -8.9176 | 5 | 10.3528 | 0 | 0.81429 | 4.54 | 4.54 | 0 | 5 | [14 empty] | 4.54
[6 empty] | 2.15 | 4.55 | 5 | 322.8272 | 26.7794 | -8.7393 | 5 | 10.3411 | 0 | 0.81357 | 4.55 | 4.55 | 0 | 5 | [14 empty] | 4.55
[6 empty] | 2.16 | 4.56 | 5 | 321.3707 | 25.3374 | -8.5645 | 5 | 10.3294 | 0 | 0.81287 | 4.56 | 4.56 | 0 | 5 | [14 empty] | 4.56
[6 empty] | 2.17 | 4.57 | 5 | 319.9433 | 23.9243 | -8.3932 | 5 | 10.3177 | 0 | 0.81219 | 4.57 | 4.57 | 0 | 5 | [14 empty] | 4.57
[6 empty] | 2.18 | 4.58 | 5 | 318.5444 | 22.5394 | -8.2253 | 5 | 10.3062 | 0 | 0.81151 | 4.58 | 4.58 | 0 | 5 | [14 empty] | 4.58
[6 empty] | 2.19 | 4.59 | 5 | 317.1735 | 21.1822 | -8.0608 | 5 | 10.2946 | 0 | 0.81085 | 4.59 | 4.59 | 0 | 5 | [14 empty] | 4.59
[6 empty] | 2.2 | 4.6 | 5 | 315.83 | 19.8522 | -7.8996 | 5 | 10.2831 | 0 | 0.81021 | 4.6 | 4.6 | 0 | 5 | [14 empty] | 4.6
[6 empty] | 2.21 | 4.61 | 5 | 314.5134 | 18.5488 | -7.7416 | 5 | 10.2717 | 0 | 0.80958 | 4.61 | 4.61 | 0 | 5 | [14 empty] | 4.61
[6 empty] | 2.22 | 4.62 | 5 | 313.2232 | 17.2714 | -7.5868 | 5 | 10.2603 | 0 | 0.80896 | 4.62 | 4.62 | 0 | 5 | [14 empty] | 4.62
[6 empty] | 2.23 | 4.63 | 5 | 311.9587 | 16.0196 | -7.435 | 5 | 10.249 | 0 | 0.80835 | 4.63 | 4.63 | 0 | 5 | [14 empty] | 4.63
[6 empty] | 2.24 | 4.64 | 5 | 310.7195 | 14.7928 | -7.2863 | 5 | 10.2377 | 0 | 0.80775 | 4.64 | 4.64 | 0 | 5 | [14 empty] | 4.64
[6 empty] | 2.25 | 4.65 | 5 | 309.5051 | 13.5905 | -7.1406 | 5 | 10.2265 | 0 | 0.80717 | 4.65 | 4.65 | 0 | 5 | [14 empty] | 4.65
[6 empty] | 2.26 | 4.66 | 5 | 308.315 | 12.4123 | -6.9978 | 5 | 10.2153 | 0 | 0.8066 | 4.66 | 4.66 | 0 | 5 | [14 empty] | 4.66
[6 empty] | 2.27 | 4.67 | 5 | 307.1487 | 11.2577 | -6.8578 | 5 | 10.2041 | 0 | 0.80604 | 4.67 | 4.67 | 0 | 5 | [14 empty] | 4.67
[6 empty] | 2.28 | 4.68 | 5 | 306.0058 | 10.1262 | -6.7207 | 5 | 10.193 | 0 | 0.80549 | 4.68 | 4.68 | 0 | 5 | [14 empty] | 4.68
[6 empty] | 2.29 | 4.69 | 5 | 304.8856 | 9.0172 | -6.5863 | 5 | 10.182 | 0 | 0.80495 | 4.69 | 4.69 | 0 | 5 | [14 empty] | 4.69
[6 empty] | 2.3 | 4.7 | 5 | 303.7879 | 7.9305 | -6.4546 | 5 | 10.171 | 0 | 0.80442 | 4.7 | 4.7 | 0 | 5 | [14 empty] | 4.7
[6 empty] | 2.31 | 4.71 | 5 | 302.7122 | 6.8655 | -6.3255 | 5 | 10.16 | 0 | 0.8039 | 4.71 | 4.71 | 0 | 5 | [14 empty] | 4.71
[6 empty] | 2.32 | 4.72 | 5 | 301.6579 | 5.8218 | -6.199 | 5 | 10.1491 | 0 | 0.8034 | 4.72 | 4.72 | 0 | 5 | [14 empty] | 4.72
[6 empty] | 2.33 | 4.73 | 5 | 300.6248 | 4.799 | -6.075 | 5 | 10.1382 | 0 | 0.8029 | 4.73 | 4.73 | 0 | 5 | [14 empty] | 4.73
[6 empty] | 2.34 | 4.74 | 5 | 299.6123 | 3.7966 | -5.9535 | 5 | 10.1274 | 0 | 0.80241 | 4.74 | 4.74 | 0 | 5 | [14 empty] | 4.74
[6 empty] | 2.35 | 4.75 | 5 | 298.62 | 2.8143 | -5.8344 | 5 | 10.1167 | 0 | 0.80194 | 4.75 | 4.75 | 0 | 5 | [14 empty] | 4.75
[6 empty] | 2.36 | 4.76 | 5 | 297.6476 | 1.8516 | -5.7177 | 5 | 10.1059 | 0 | 0.80147 | 4.76 | 4.76 | 0 | 5 | [14 empty] | 4.76
[6 empty] | 2.37 | 4.77 | 5 | 296.6947 | 0.90818 | -5.6034 | 5 | 10.0952 | 0 | 0.80101 | 4.77 | 4.77 | 0 | 5 | [14 empty] | 4.77
[6 empty] | 2.38 | 4.78 | 5 | 295.7608 | -0.016374 | -5.4913 | 5 | 10.0846 | 0 | 0.80056 | 4.78 | 4.78 | 0 | 5 | [14 empty] | 4.78
[6 empty] | 2.39 | 4.79 | 5 | 294.8456 | -0.92244 | -5.3815 | 5 | 10.074 | 0 | 0.80012 | 4.79 | 4.79 | 0 | 5 | [14 empty] | 4.79
[6 empty] | 3.9968e-14 | 4.8 | 12.7123 | 293.9487 | 2.0458 | 41 | 12.7123 | 10.0795 | 0 | 0.80864 | 4.8 | 4.8 | 0 | 12.7123 | [14 empty] | 4.8
[6 empty] | 0.01 | 4.81 | 12.7934 | 300.782 | 8.8513 | 40.6665 | 12.7934 | 10.0851 | 0 | 0.81715 | 4.81 | 4.81 | 0 | 12.7934 | [14 empty] | 4.81
[6 empty] | 0.02 | 4.82 | 12.8734 | 307.5597 | 15.6013 | 40.3329 | 12.8734 | 10.0909 | 0 | 0.82564 | 4.82 | 4.82 | 0 | 12.8734 | [14 empty] | 4.82
[6 empty] | 0.03 | 4.83 | 12.9522 | 314.2819 | 22.2956 | 39.9994 | 12.9522 | 10.0968 | 0 | 0.83411 | 4.83 | 4.83 | 0 | 12.9522 | [14 empty] | 4.83
[6 empty] | 0.04 | 4.84 | 13.0299 | 320.9485 | 28.9344 | 39.6659 | 13.0299 | 10.1029 | 0 | 0.84256 | 4.84 | 4.84 | 0 | 13.0299 | [14 empty] | 4.84
[6 empty] | 0.05 | 4.85 | 13.1066 | 327.5594 | 35.5176 | 39.3323 | 13.1066 | 10.109 | 0 | 0.85099 | 4.85 | 4.85 | 0 | 13.1066 | [14 empty] | 4.85
[6 empty] | 0.06 | 4.86 | 13.1821 | 334.1148 | 42.0452 | 38.9988 | 13.1821 | 10.1154 | 0 | 0.85939 | 4.86 | 4.86 | 0 | 13.1821 | [14 empty] | 4.86
[6 empty] | 0.07 | 4.87 | 13.2565 | 340.6146 | 48.5172 | 38.6653 | 13.2565 | 10.1218 | 0 | 0.86777 | 4.87 | 4.87 | 0 | 13.2565 | [14 empty] | 4.87
[6 empty] | 0.08 | 4.88 | 13.3298 | 347.0588 | 54.9336 | 38.3317 | 13.3298 | 10.1284 | 0 | 0.87613 | 4.88 | 4.88 | 0 | 13.3298 | [14 empty] | 4.88
[6 empty] | 0.09 | 4.89 | 13.402 | 353.4475 | 61.2944 | 37.9982 | 13.402 | 10.135 | 0 | 0.88445 | 4.89 | 4.89 | 0 | 13.402 | [14 empty] | 4.89
[6 empty] | 0.1 | 4.9 | 13.4731 | 359.7805 | 67.5997 | 37.6647 | 13.4731 | 10.1418 | 0 | 0.89275 | 4.9 | 4.9 | 0 | 13.4731 | [14 empty] | 4.9
[6 empty] | 0.11 | 4.91 | 13.543 | 366.0579 | 73.8493 | 37.3312 | 13.543 | 10.1487 | 0 | 0.90102 | 4.91 | 4.91 | 0 | 13.543 | [14 empty] | 4.91
[6 empty] | 0.12 | 4.92 | 13.6119 | 372.2798 | 80.0434 | 36.9976 | 13.6119 | 10.1558 | 0 | 0.90925 | 4.92 | 4.92 | 0 | 13.6119 | [14 empty] | 4.92
[6 empty] | 0.13 | 4.93 | 13.6796 | 378.4461 | 86.1819 | 36.6641 | 13.6796 | 10.1629 | 0 | 0.91745 | 4.93 | 4.93 | 0 | 13.6796 | [14 empty] | 4.93
[6 empty] | 0.14 | 4.94 | 13.7462 | 384.5568 | 92.2648 | 36.3306 | 13.7462 | 10.1701 | 0 | 0.92562 | 4.94 | 4.94 | 0 | 13.7462 | [14 empty] | 4.94
[6 empty] | 0.15 | 4.95 | 13.8117 | 390.6119 | 98.2921 | 35.997 | 13.8117 | 10.1775 | 0 | 0.93375 | 4.95 | 4.95 | 0 | 13.8117 | [14 empty] | 4.95
[6 empty] | 0.16 | 4.96 | 13.8761 | 396.6114 | 104.2638 | 35.6635 | 13.8761 | 10.1849 | 0 | 0.94184 | 4.96 | 4.96 | 0 | 13.8761 | [14 empty] | 4.96
[6 empty] | 0.17 | 4.97 | 13.9394 | 402.5553 | 110.1799 | 35.33 | 13.9394 | 10.1925 | 0 | 0.94989 | 4.97 | 4.97 | 0 | 13.9394 | [14 empty] | 4.97
[6 empty] | 0.18 | 4.98 | 14.0016 | 408.4436 | 116.0404 | 34.9964 | 14.0016 | 10.2001 | 0 | 0.9579 | 4.98 | 4.98 | 0 | 14.0016 | [14 empty] | 4.98
[6 empty] | 0.19 | 4.99 | 14.0627 | 414.2763 | 121.8454 | 34.6629 | 14.0627 | 10.2078 | 0 | 0.96587 | 4.99 | 4.99 | 0 | 14.0627 | [14 empty] | 4.99
[6 empty] | 0.2 | 5 | 14.1226 | 420.0535 | 127.5947 | 34.3294 | 14.1226 | 10.2156 | 0 | 0.9738 | 5 | 5 | 0 | 14.1226 | [14 empty] | 5
[6 empty] | 0.21 | 5.01 | 14.1815 | 425.775 | 133.2885 | 33.9958 | 14.1815 | 10.2235 | 0 | 0.98168 | 5.01 | 5.01 | 0 | 14.1815 | [14 empty] | 5.01
[6 empty] | 0.22 | 5.02 | 14.2392 | 431.441 | 138.9267 | 33.6623 | 14.2392 | 10.2315 | 0 | 0.98952 | 5.02 | 5.02 | 0 | 14.2392 | [14 empty] | 5.02
[6 empty] | 0.23 | 5.03 | 14.2958 | 437.0514 | 144.5092 | 33.3288 | 14.2958 | 10.2396 | 0 | 0.99731 | 5.03 | 5.03 | 0 | 14.2958 | [14 empty] | 5.03
[6 empty] | 0.24 | 5.04 | 14.3513 | 442.6062 | 150.0362 | 32.9952 | 14.3513 | 10.2477 | 0 | 1.005 | 5.04 | 5.04 | 0 | 14.3513 | [14 empty] | 5.04
[6 empty] | 0.25 | 5.05 | 14.4057 | 448.1054 | 155.5077 | 32.6617 | 14.4057 | 10.2559 | 0 | 1.0127 | 5.05 | 5.05 | 0 | 14.4057 | [14 empty] | 5.05
[6 empty] | 0.26 | 5.06 | 14.459 | 453.549 | 160.9235 | 32.3282 | 14.459 | 10.2642 | 0 | 1.0204 | 5.06 | 5.06 | 0 | 14.459 | [14 empty] | 5.06
[6 empty] | 0.27 | 5.07 | 14.5112 | 458.9371 | 166.2837 | 31.9946 | 14.5112 | 10.2726 | 0 | 1.028 | 5.07 | 5.07 | 0 | 14.5112 | [14 empty] | 5.07
[6 empty] | 0.28 | 5.08 | 14.5622 | 464.2695 | 171.5884 | 31.6611 | 14.5622 | 10.281 | 0 | 1.0355 | 5.08 | 5.08 | 0 | 14.5622 | [14 empty] | 5.08
[6 empty] | 0.29 | 5.09 | 14.6122 | 469.5463 | 176.8374 | 31.3276 | 14.6122 | 10.2895 | 0 | 1.043 | 5.09 | 5.09 | 0 | 14.6122 | [14 empty] | 5.09
[6 empty] | 0.3 | 5.1 | 14.661 | 474.7676 | 182.0309 | 30.994 | 14.661 | 10.2981 | 0 | 1.0504 | 5.1 | 5.1 | 0 | 14.661 | [14 empty] | 5.1
[6 empty] | 0.31 | 5.11 | 14.7088 | 479.9333 | 187.1688 | 30.6605 | 14.7088 | 10.3067 | 0 | 1.0578 | 5.11 | 5.11 | 0 | 14.7088 | [14 empty] | 5.11
[6 empty] | 0.32 | 5.12 | 14.7554 | 485.0434 | 192.2511 | 30.327 | 14.7554 | 10.3154 | 0 | 1.0651 | 5.12 | 5.12 | 0 | 14.7554 | [14 empty] | 5.12
[6 empty] | 0.33 | 5.13 | 14.8009 | 490.0979 | 197.2778 | 29.9935 | 14.8009 | 10.3241 | 0 | 1.0724 | 5.13 | 5.13 | 0 | 14.8009 | [14 empty] | 5.13
[6 empty] | 0.34 | 5.14 | 14.8453 | 495.0968 | 202.2489 | 29.6599 | 14.8453 | 10.3329 | 0 | 1.0796 | 5.14 | 5.14 | 0 | 14.8453 | [14 empty] | 5.14
[6 empty] | 0.35 | 5.15 | 14.8885 | 500.0401 | 207.1644 | 29.3264 | 14.8885 | 10.3417 | 0 | 1.0867 | 5.15 | 5.15 | 0 | 14.8885 | [14 empty] | 5.15
[6 empty] | 0.36 | 5.16 | 14.9307 | 504.9278 | 212.0243 | 28.9929 | 14.9307 | 10.3506 | 0 | 1.0938 | 5.16 | 5.16 | 0 | 14.9307 | [14 empty] | 5.16
[6 empty] | 0.37 | 5.17 | 14.9718 | 509.76 | 216.8287 | 28.6593 | 14.9718 | 10.3595 | 0 | 1.1008 | 5.17 | 5.17 | 0 | 14.9718 | [14 empty] | 5.17
[6 empty] | 0.38 | 5.18 | 15.0117 | 514.5365 | 221.5775 | 28.3258 | 15.0117 | 10.3685 | 0 | 1.1078 | 5.18 | 5.18 | 0 | 15.0117 | [14 empty] | 5.18
[6 empty] | 0.39 | 5.19 | 15.0505 | 519.2575 | 226.2706 | 27.9923 | 15.0505 | 10.3775 | 0 | 1.1146 | 5.19 | 5.19 | 0 | 15.0505 | [14 empty] | 5.19
[6 empty] | 0.4 | 5.2 | 15.0882 | 523.9229 | 230.9082 | 27.6587 | 15.0882 | 10.3865 | 0 | 1.1215 | 5.2 | 5.2 | 0 | 15.0882 | [14 empty] | 5.2
[6 empty] | 0.41 | 5.21 | 15.1249 | 528.5327 | 235.4902 | 27.3252 | 15.1249 | 10.3956 | 0 | 1.1282 | 5.21 | 5.21 | 0 | 15.1249 | [14 empty] | 5.21
[6 empty] | 0.42 | 5.22 | 15.1603 | 533.0869 | 240.0166 | 26.9917 | 15.1603 | 10.4047 | 0 | 1.1349 | 5.22 | 5.22 | 0 | 15.1603 | [14 empty] | 5.22
[6 empty] | 0.43 | 5.23 | 15.1947 | 537.5855 | 244.4874 | 26.6581 | 15.1947 | 10.4138 | 0 | 1.1415 | 5.23 | 5.23 | 0 | 15.1947 | [14 empty] | 5.23
[6 empty] | 0.44 | 5.24 | 15.228 | 542.0285 | 248.9027 | 26.3246 | 15.228 | 10.423 | 0 | 1.1481 | 5.24 | 5.24 | 0 | 15.228 | [14 empty] | 5.24
[6 empty] | 0.45 | 5.25 | 15.2602 | 546.4159 | 253.2623 | 25.9911 | 15.2602 | 10.4322 | 0 | 1.1546 | 5.25 | 5.25 | 0 | 15.2602 | [14 empty] | 5.25
[6 empty] | 0.46 | 5.26 | 15.2912 | 550.7478 | 257.5663 | 25.6575 | 15.2912 | 10.4414 | 0 | 1.161 | 5.26 | 5.26 | 0 | 15.2912 | [14 empty] | 5.26
[6 empty] | 0.47 | 5.27 | 15.3211 | 555.024 | 261.8148 | 25.324 | 15.3211 | 10.4507 | 0 | 1.1673 | 5.27 | 5.27 | 0 | 15.3211 | [14 empty] | 5.27
[6 empty] | 0.48 | 5.28 | 15.35 | 559.2447 | 266.0077 | 24.9905 | 15.35 | 10.4599 | 0 | 1.1736 | 5.28 | 5.28 | 0 | 15.35 | [14 empty] | 5.28
[6 empty] | 0.49 | 5.29 | 15.3777 | 563.4098 | 270.145 | 24.6569 | 15.3777 | 10.4692 | 0 | 1.1798 | 5.29 | 5.29 | 0 | 15.3777 | [14 empty] | 5.29
[6 empty] | 0.5 | 5.3 | 15.4043 | 567.5193 | 274.2267 | 24.3234 | 15.4043 | 10.4785 | 0 | 1.1859 | 5.3 | 5.3 | 0 | 15.4043 | [14 empty] | 5.3
[6 empty] | 0.51 | 5.31 | 15.4298 | 571.5732 | 278.2528 | 23.9899 | 15.4298 | 10.4878 | 0 | 1.192 | 5.31 | 5.31 | 0 | 15.4298 | [14 empty] | 5.31
[6 empty] | 0.52 | 5.32 | 15.4542 | 575.5715 | 282.2233 | 23.6563 | 15.4542 | 10.4971 | 0 | 1.198 | 5.32 | 5.32 | 0 | 15.4542 | [14 empty] | 5.32
[6 empty] | 0.53 | 5.33 | 15.4774 | 579.5142 | 286.1382 | 23.3228 | 15.4774 | 10.5064 | 0 | 1.2039 | 5.33 | 5.33 | 0 | 15.4774 | [14 empty] | 5.33
[6 empty] | 0.54 | 5.34 | 15.4996 | 583.4013 | 289.9976 | 22.9893 | 15.4996 | 10.5158 | 0 | 1.2097 | 5.34 | 5.34 | 0 | 15.4996 | [14 empty] | 5.34
[6 empty] | 0.55 | 5.35 | 15.5206 | 587.2329 | 293.8013 | 22.6558 | 15.5206 | 10.5251 | 0 | 1.2154 | 5.35 | 5.35 | 0 | 15.5206 | [14 empty] | 5.35
[6 empty] | 0.56 | 5.36 | 15.5405 | 591.0088 | 297.5495 | 22.3222 | 15.5405 | 10.5344 | 0 | 1.2211 | 5.36 | 5.36 | 0 | 15.5405 | [14 empty] | 5.36
[6 empty] | 0.57 | 5.37 | 15.5594 | 594.7292 | 301.2421 | 21.9887 | 15.5594 | 10.5438 | 0 | 1.2267 | 5.37 | 5.37 | 0 | 15.5594 | [14 empty] | 5.37
[6 empty] | 0.58 | 5.38 | 15.5771 | 598.394 | 304.879 | 21.6552 | 15.5771 | 10.5531 | 0 | 1.2322 | 5.38 | 5.38 | 0 | 15.5771 | [14 empty] | 5.38
[6 empty] | 0.59 | 5.39 | 15.5937 | 602.0032 | 308.4604 | 21.3216 | 15.5937 | 10.5625 | 0 | 1.2377 | 5.39 | 5.39 | 0 | 15.5937 | [14 empty] | 5.39
[6 empty] | 0.6 | 5.4 | 15.6092 | 605.5568 | 311.9862 | 20.9881 | 15.6092 | 10.5718 | 0 | 1.243 | 5.4 | 5.4 | 0 | 15.6092 | [14 empty] | 5.4
[6 empty] | 0.61 | 5.41 | 15.6235 | 609.0548 | 315.4565 | 20.6546 | 15.6235 | 10.5811 | 0 | 1.2483 | 5.41 | 5.41 | 0 | 15.6235 | [14 empty] | 5.41
[6 empty] | 0.62 | 5.42 | 15.6368 | 612.4972 | 318.8711 | 20.321 | 15.6368 | 10.5904 | 0 | 1.2535 | 5.42 | 5.42 | 0 | 15.6368 | [14 empty] | 5.42
[6 empty] | 0.63 | 5.43 | 15.6489 | 615.8841 | 322.2301 | 19.9875 | 15.6489 | 10.5997 | 0 | 1.2586 | 5.43 | 5.43 | 0 | 15.6489 | [14 empty] | 5.43
[6 empty] | 0.64 | 5.44 | 15.66 | 619.2153 | 325.5336 | 19.654 | 15.66 | 10.609 | 0 | 1.2636 | 5.44 | 5.44 | 0 | 15.66 | [14 empty] | 5.44
[6 empty] | 0.65 | 5.45 | 15.6699 | 622.491 | 328.7815 | 19.3204 | 15.6699 | 10.6183 | 0 | 1.2686 | 5.45 | 5.45 | 0 | 15.6699 | [14 empty] | 5.45
[6 empty] | 0.66 | 5.46 | 15.6787 | 625.7111 | 331.9737 | 18.9869 | 15.6787 | 10.6275 | 0 | 1.2735 | 5.46 | 5.46 | 0 | 15.6787 | [14 empty] | 5.46
[6 empty] | 0.67 | 5.47 | 15.6864 | 628.8755 | 335.1104 | 18.6534 | 15.6864 | 10.6368 | 0 | 1.2782 | 5.47 | 5.47 | 0 | 15.6864 | [14 empty] | 5.47
[6 empty] | 0.68 | 5.48 | 15.693 | 631.9844 | 338.1915 | 18.3198 | 15.693 | 10.646 | 0 | 1.2829 | 5.48 | 5.48 | 0 | 15.693 | [14 empty] | 5.48
[6 empty] | 0.69 | 5.49 | 15.6985 | 635.0377 | 341.2171 | 17.9863 | 15.6985 | 10.6552 | 0 | 1.2876 | 5.49 | 5.49 | 0 | 15.6985 | [14 empty] | 5.49
[6 empty] | 0.7 | 5.5 | 15.7028 | 638.0355 | 344.187 | 17.6528 | 15.7028 | 10.6643 | 0 | 1.2921 | 5.5 | 5.5 | 0 | 15.7028 | [14 empty] | 5.5
[6 empty] | 0.71 | 5.51 | 15.7061 | 640.9776 | 347.1013 | 17.3192 | 15.7061 | 10.6734 | 0 | 1.2965 | 5.51 | 5.51 | 0 | 15.7061 | [14 empty] | 5.51
[6 empty] | 0.72 | 5.52 | 15.7082 | 643.8641 | 349.9601 | 16.9857 | 15.7082 | 10.6826 | 0 | 1.3009 | 5.52 | 5.52 | 0 | 15.7082 | [14 empty] | 5.52
[6 empty] | 0.73 | 5.53 | 15.7093 | 646.6951 | 352.7632 | 16.6522 | 15.7093 | 10.6916 | 0 | 1.3052 | 5.53 | 5.53 | 0 | 15.7093 | [14 empty] | 5.53
[6 empty] | 0.74 | 5.54 | 15.7092 | 649.4704 | 355.5108 | 16.3187 | 15.7092 | 10.7007 | 0 | 1.3094 | 5.54 | 5.54 | 0 | 15.7092 | [14 empty] | 5.54
[6 empty] | 0.75 | 5.55 | 15.708 | 652.1902 | 358.2028 | 15.9851 | 15.708 | 10.7097 | 0 | 1.3135 | 5.55 | 5.55 | 0 | 15.708 | [14 empty] | 5.55
[6 empty] | 0.76 | 5.56 | 15.7057 | 654.8544 | 360.8392 | 15.6516 | 15.7057 | 10.7186 | 0 | 1.3175 | 5.56 | 5.56 | 0 | 15.7057 | [14 empty] | 5.56
[6 empty] | 0.77 | 5.57 | 15.7023 | 657.463 | 363.42 | 15.3181 | 15.7023 | 10.7276 | 0 | 1.3214 | 5.57 | 5.57 | 0 | 15.7023 | [14 empty] | 5.57
[6 empty] | 0.78 | 5.58 | 15.6977 | 660.016 | 365.9452 | 14.9845 | 15.6977 | 10.7365 | 0 | 1.3253 | 5.58 | 5.58 | 0 | 15.6977 | [14 empty] | 5.58
[6 empty] | 0.79 | 5.59 | 15.6921 | 662.5134 | 368.4148 | 14.651 | 15.6921 | 10.7453 | 0 | 1.329 | 5.59 | 5.59 | 0 | 15.6921 | [14 empty] | 5.59
[6 empty] | 0.8 | 5.6 | 15.6853 | 664.9553 | 370.8288 | 14.3175 | 15.6853 | 10.7541 | 0 | 1.3327 | 5.6 | 5.6 | 0 | 15.6853 | [14 empty] | 5.6
[6 empty] | 0.81 | 5.61 | 15.6775 | 667.3415 | 373.1873 | 13.9839 | 15.6775 | 10.7629 | 0 | 1.3363 | 5.61 | 5.61 | 0 | 15.6775 | [14 empty] | 5.61
[6 empty] | 0.82 | 5.62 | 15.6685 | 669.6722 | 375.4902 | 13.6504 | 15.6685 | 10.7716 | 0 | 1.3398 | 5.62 | 5.62 | 0 | 15.6685 | [14 empty] | 5.62
[6 empty] | 0.83 | 5.63 | 15.6584 | 671.9472 | 377.7374 | 13.3169 | 15.6584 | 10.7803 | 0 | 1.3432 | 5.63 | 5.63 | 0 | 15.6584 | [14 empty] | 5.63
[6 empty] | 0.84 | 5.64 | 15.6472 | 674.1667 | 379.9291 | 12.9833 | 15.6472 | 10.7889 | 0 | 1.3465 | 5.64 | 5.64 | 0 | 15.6472 | [14 empty] | 5.64
[6 empty] | 0.85 | 5.65 | 15.6349 | 676.3306 | 382.0652 | 12.6498 | 15.6349 | 10.7974 | 0 | 1.3497 | 5.65 | 5.65 | 0 | 15.6349 | [14 empty] | 5.65
[6 empty] | 0.86 | 5.66 | 15.6215 | 678.4389 | 384.1457 | 12.3163 | 15.6215 | 10.8059 | 0 | 1.3529 | 5.66 | 5.66 | 0 | 15.6215 | [14 empty] | 5.66
[6 empty] | 0.87 | 5.67 | 15.607 | 680.4916 | 386.1706 | 11.9827 | 15.607 | 10.8144 | 0 | 1.3559 | 5.67 | 5.67 | 0 | 15.607 | [14 empty] | 5.67
[6 empty] | 0.88 | 5.68 | 15.5913 | 682.4887 | 388.14 | 11.6492 | 15.5913 | 10.8228 | 0 | 1.3589 | 5.68 | 5.68 | 0 | 15.5913 | [14 empty] | 5.68
[6 empty] | 0.89 | 5.69 | 15.5746 | 684.4303 | 390.0537 | 11.3157 | 15.5746 | 10.8311 | 0 | 1.3618 | 5.69 | 5.69 | 0 | 15.5746 | [14 empty] | 5.69
[6 empty] | 0.9 | 5.7 | 15.5567 | 686.3162 | 391.9118 | 10.9821 | 15.5567 | 10.8394 | 0 | 1.3646 | 5.7 | 5.7 | 0 | 15.5567 | [14 empty] | 5.7
[6 empty] | 0.91 | 5.71 | 15.5377 | 688.1466 | 393.7144 | 10.6486 | 15.5377 | 10.8476 | 0 | 1.3673 | 5.71 | 5.71 | 0 | 15.5377 | [14 empty] | 5.71
[6 empty] | 0.92 | 5.72 | 15.5176 | 689.9213 | 395.4614 | 10.3151 | 15.5176 | 10.8558 | 0 | 1.3699 | 5.72 | 5.72 | 0 | 15.5176 | [14 empty] | 5.72
[6 empty] | 0.93 | 5.73 | 15.4964 | 691.6405 | 397.1528 | 9.9815 | 15.4964 | 10.8639 | 0 | 1.3724 | 5.73 | 5.73 | 0 | 15.4964 | [14 empty] | 5.73
[6 empty] | 0.94 | 5.74 | 15.4741 | 693.3041 | 398.7886 | 9.648 | 15.4741 | 10.8719 | 0 | 1.3749 | 5.74 | 5.74 | 0 | 15.4741 | [14 empty] | 5.74
[6 empty] | 0.95 | 5.75 | 15.4507 | 694.9121 | 400.3688 | 9.3145 | 15.4507 | 10.8798 | 0 | 1.3772 | 5.75 | 5.75 | 0 | 15.4507 | [14 empty] | 5.75
[6 empty] | 0.96 | 5.76 | 15.4261 | 696.4645 | 401.8934 | 8.981 | 15.4261 | 10.8877 | 0 | 1.3795 | 5.76 | 5.76 | 0 | 15.4261 | [14 empty] | 5.76
[6 empty] | 0.97 | 5.77 | 15.4005 | 697.9614 | 403.3624 | 8.6474 | 15.4005 | 10.8955 | 0 | 1.3817 | 5.77 | 5.77 | 0 | 15.4005 | [14 empty] | 5.77
[6 empty] | 0.98 | 5.78 | 15.3737 | 699.4026 | 404.7759 | 8.3139 | 15.3737 | 10.9032 | 0 | 1.3838 | 5.78 | 5.78 | 0 | 15.3737 | [14 empty] | 5.78
[6 empty] | 0.99 | 5.79 | 15.3458 | 700.7882 | 406.1337 | 7.9804 | 15.3458 | 10.9109 | 0 | 1.3858 | 5.79 | 5.79 | 0 | 15.3458 | [14 empty] | 5.79
[6 empty] | 1 | 5.8 | 15.3168 | 702.1183 | 407.436 | 7.6468 | 15.3168 | 10.9185 | 0 | 1.3877 | 5.8 | 5.8 | 0 | 15.3168 | [14 empty] | 5.8
[6 empty] | 1.01 | 5.81 | 15.2867 | 703.3928 | 408.6827 | 7.3133 | 15.2867 | 10.926 | 0 | 1.3895 | 5.81 | 5.81 | 0 | 15.2867 | [14 empty] | 5.81
[6 empty] | 1.02 | 5.82 | 15.2555 | 704.6117 | 409.8737 | 6.9798 | 15.2555 | 10.9334 | 0 | 1.3913 | 5.82 | 5.82 | 0 | 15.2555 | [14 empty] | 5.82
[6 empty] | 1.03 | 5.83 | 15.2232 | 705.7749 | 411.0092 | 6.6462 | 15.2232 | 10.9408 | 0 | 1.3929 | 5.83 | 5.83 | 0 | 15.2232 | [14 empty] | 5.83
[6 empty] | 1.04 | 5.84 | 15.1898 | 706.8827 | 412.0892 | 6.3127 | 15.1898 | 10.948 | 0 | 1.3945 | 5.84 | 5.84 | 0 | 15.1898 | [14 empty] | 5.84
[6 empty] | 1.05 | 5.85 | 15.1552 | 707.9348 | 413.1135 | 5.9792 | 15.1552 | 10.9552 | 0 | 1.396 | 5.85 | 5.85 | 0 | 15.1552 | [14 empty] | 5.85
[6 empty] | 1.06 | 5.86 | 15.1196 | 708.9313 | 414.0822 | 5.6456 | 15.1196 | 10.9623 | 0 | 1.3974 | 5.86 | 5.86 | 0 | 15.1196 | [14 empty] | 5.86
[6 empty] | 1.07 | 5.87 | 15.0828 | 709.8722 | 414.9954 | 5.3121 | 15.0828 | 10.9693 | 0 | 1.3987 | 5.87 | 5.87 | 0 | 15.0828 | [14 empty] | 5.87
[6 empty] | 1.08 | 5.88 | 15.0449 | 710.7576 | 415.8529 | 4.9786 | 15.0449 | 10.9762 | 0 | 1.3999 | 5.88 | 5.88 | 0 | 15.0449 | [14 empty] | 5.88
[6 empty] | 1.09 | 5.89 | 15.0059 | 711.5873 | 416.6549 | 4.645 | 15.0059 | 10.9831 | 0 | 1.401 | 5.89 | 5.89 | 0 | 15.0059 | [14 empty] | 5.89
[6 empty] | 1.1 | 5.9 | 14.9658 | 712.3615 | 417.4013 | 4.3115 | 14.9658 | 10.9898 | 0 | 1.4021 | 5.9 | 5.9 | 0 | 14.9658 | [14 empty] | 5.9
[6 empty] | 1.11 | 5.91 | 14.9246 | 713.0801 | 418.0921 | 3.978 | 14.9246 | 10.9964 | 0 | 1.4031 | 5.91 | 5.91 | 0 | 14.9246 | [14 empty] | 5.91
[6 empty] | 1.12 | 5.92 | 14.8823 | 713.7431 | 418.7273 | 3.6444 | 14.8823 | 11.003 | 0 | 1.4039 | 5.92 | 5.92 | 0 | 14.8823 | [14 empty] | 5.92
[6 empty] | 1.13 | 5.93 | 14.8388 | 714.3505 | 419.3069 | 3.3109 | 14.8388 | 11.0094 | 0 | 1.4047 | 5.93 | 5.93 | 0 | 14.8388 | [14 empty] | 5.93
[6 empty] | 1.14 | 5.94 | 14.7943 | 714.9023 | 419.8309 | 2.9774 | 14.7943 | 11.0158 | 0 | 1.4055 | 5.94 | 5.94 | 0 | 14.7943 | [14 empty] | 5.94
[6 empty] | 1.15 | 5.95 | 14.7486 | 715.3986 | 420.2993 | 2.6438 | 14.7486 | 11.0221 | 0 | 1.4061 | 5.95 | 5.95 | 0 | 14.7486 | [14 empty] | 5.95
[6 empty] | 1.16 | 5.96 | 14.7018 | 715.8392 | 420.7122 | 2.3103 | 14.7018 | 11.0282 | 0 | 1.4067 | 5.96 | 5.96 | 0 | 14.7018 | [14 empty] | 5.96
[6 empty] | 1.17 | 5.97 | 14.6539 | 716.2243 | 421.0694 | 1.9768 | 14.6539 | 11.0343 | 0 | 1.4071 | 5.97 | 5.97 | 0 | 14.6539 | [14 empty] | 5.97
[6 empty] | 1.18 | 5.98 | 14.6049 | 716.5537 | 421.3711 | 1.6433 | 14.6049 | 11.0403 | 0 | 1.4075 | 5.98 | 5.98 | 0 | 14.6049 | [14 empty] | 5.98
[6 empty] | 1.19 | 5.99 | 14.5548 | 716.8276 | 421.6172 | 1.3097 | 14.5548 | 11.0461 | 0 | 1.4078 | 5.99 | 5.99 | 0 | 14.5548 | [14 empty] | 5.99
[6 empty] | 1.2 | 6 | 14.5036 | 717.0459 | 421.8077 | 0.97619 | 14.5036 | 11.0519 | 0 | 1.4081 | 6 | 6 | 0 | 14.5036 | [14 empty] | 6
Menu
square
ramp
Passive Pressure -Volume Curve
Blue = Static Green = Dynamic

## PC (PIP)
Pressure Control
Lung Model
|  | Elastance (cm H2O/L) | 20
|  | Compliance (L/cm H2O) | 0.05 |  | 50
|  | Resistance (cm H2O/L/s) | 10 |  | 10
|  | Time Constant (s) | 0.5
|  | Tidal Volume (mL) | 499.6122
|  | Potential display error (%) | 0.10139 |  | Ref.
Effort Model
|  | Delay Before Effort (s) | 0.12 |  | 12
|  | Duration of Effort (s) | 0.8 |  | 16
|  | Freq (hz) | 0.3125
|  | Period (s) | 1.6
|  | max Pmus (cm H2O) | 0 | 49 | 0
Ventilator Settings
|  | Inspiratory Pressure Target (cm H2O) | 11
|  | Peak Inspiratory Pressure (cm H2O) | 16 |  | 16
|  | Frequency (breaths/min) | 20 |  | 20
|  | Inflation Pressure Time (s) | 1.5 |  | 15
|  | PEEP (cm H2O) | 5 |  | 5
|  | autoPEEP (cm H2O) | 0.51763
|  | Pause Time (s) | 0 |  | 0
|  | Peak Airway Pressure (cm H2O) | 16
|  | Total Cycle Time (s) | 3
|  | Inspirtory Time (s) | 1.5
|  | Expiratory Time (s) | 1.5
|  | I:E | 1:1
|  | Duty Cycle (%) | 50
|  | Plateau Pressure (cm H2O) | 15.4934
|  | Tidal (Driving) Pressure (cm H2O) | 9.9922
|  | Peak Inspiratory Flow (L/min) | 62.932
|  | Peak Expiratory Flow (L/min) | -63.0214
|  | Mean Inspiratory Pressure | 16
|  | Mean Airway Pressure | 10.5184
|  | End Inspiratory Volume (mL) | 499.1056
|  | Inspiratory Flow Time (s) | 1.5
Passive |  | Ventilator Elastic Work (J) | 0.51056
|  | Ventilator Resistive Work (J) | 0.29545
|  | Ventilator Total Work (J) | 0.80601
Active |  | Ventilator Work (J) | 0.78419
|  | Patient Work (J) | 0
|  | Total Work (J) | 0.78419
|  | Work Shift Index (%) | 0
|  | Ventilator Inspiratory Power (W) | 0.2614
|  | Ventilator Inspiratory Power (J/min) | 15.6838
pause start time
1.5
[11 empty] | pause stop time | [8 empty] | Mean Arwy Pressure
[11 empty] | 1.5 |  |  | Peak Insp Flow |  |  |  | Elastance |  | 10.5184 |  |  | …………………….Pressure Volume Curve Data……………….
[9 empty] | Vmus equiv to constant mean Pmus |  | Initial | Passive VT |  | 62.932 |  |  |  | 0.02 |  | Mean Insp Pressure | [6 empty] | Seg |  | Seg |  | Seg |  | Seg
[9 empty] | 0 |  | 0.27557 | 513.5133 |  | Peak Exp Flow | [5 empty] | 16 | Pinsp | 11 | Pstat |  | Pdyn |  | L1 |  | L2 |  | U1 |  | U2
[11 empty] | includes est autoPEEP |  |  | -63.0214 |  |  |  |  | Integral | Mean | R | 10 | X | Y | X | Y | x | y | x | y | x | y | x | y
[8 empty] | Vent | Muscle | Lung | Lung | Tidal | Tidal | Tidal | Muscle | Muscle | Total | Total | Total | Total | C | 0.05 | 5 | 0 | 16 | 0 | 0 | 0 | 5 | 0 | 0 | 513.5133 | 15.2703 | 513.5133
[7 empty] | Relative | Pressure | Pressure | Pressure | Volume | Volume | Flow | Flow | Power | Work | Power | Work | Pressure | Pressure | RC | 0.5 | 15.2703 | 513.5133 | 16 | 513.5133 | 5 | 0 | 16 | 0 | 15.2703 | 513.5133 | 16 | 513.5133
[6 empty] | Time | Time | (cm H2O) | (cm H2O) | (cm H2O) | (L) | (mL) | (L/s) | (L/min) | (W) | (J) | (W) | (J) | (cm H2O) | (cm H2O) | Inspiratory Pressure Time | 1.5
[6 empty] | 0 | 0 | 16 | 0 | 5.5113 | 0.27557 | 0 | 1.0489 | 62.932 | 0 | 0 | 1.6463 | 0.016463 | 0 | 0 | Total Cycle Time | 3
[6 empty] | 0.01 | 0 | 16 | 0 | 5.7211 | 0.28605 | 10.4887 | 1.0279 | 61.6734 | 0 | 0 | 1.6134 | 0.032597 | 0.16 | 16 | D | 0.5
[6 empty] | 0.02 | 0 | 16 | 0 | 5.9267 | 0.29633 | 20.7676 | 1.0073 | 60.4399 | 0 | 0 | 1.5811 | 0.048408 | 0.32 | 16
[6 empty] | 0.03 | 0 | 16 | 0 | 6.1281 | 0.30641 | 30.8409 | 0.98719 | 59.2311 | 0 | 0 | 1.5495 | 0.063903 | 0.48 | 16 | f | 0.33333
[6 empty] | 0.04 | 0 | 16 | 0 | 6.3256 | 0.31628 | 40.7128 | 0.96744 | 58.0465 | 0 | 0 | 1.5185 | 0.079088 | 0.64 | 16 | fRC | 0.16667
[6 empty] | 0.05 | 0 | 16 | 0 | 6.5191 | 0.32595 | 50.3872 | 0.94809 | 56.8856 | 0 | 0 | 1.4881 | 0.093969 | 0.8 | 16 | -D/fRC | -3
[6 empty] | 0.06 | 0 | 16 | 0 | 6.7087 | 0.33543 | 59.8681 | 0.92913 | 55.7479 | 0 | 0 | 1.4584 | 0.10855 | 0.96 | 16 | -1-D/fRC | -4
[6 empty] | 0.07 | 0 | 16 | 0 | 6.8945 | 0.34473 | 69.1594 | 0.91055 | 54.6329 | 0 | 0 | 1.4292 | 0.12284 | 1.12 | 16 | exp(-D/fRC) | 0.049787
[6 empty] | 0.08 | 0 | 16 | 0 | 7.0766 | 0.35383 | 78.2649 | 0.89234 | 53.5403 | 0 | 0 | 1.4006 | 0.13685 | 1.28 | 16 | exp(-1-D/fRC) | 0.018316
[6 empty] | 0.09 | 0 | 16 | 0 | 7.2551 | 0.36275 | 87.1883 | 0.87449 | 52.4694 | 0 | 0 | 1.3726 | 0.15058 | 1.44 | 16
[6 empty] | 0.1 | 0 | 16 | 0 | 7.43 | 0.3715 | 95.9332 | 0.857 | 51.4201 | 0 | 0 | 1.3451 | 0.16403 | 1.6 | 16 | 1-exp(-D/fRC) | 0.95021
[6 empty] | 0.11 | 0 | 16 | 0 | 7.6014 | 0.38007 | 104.5032 | 0.83986 | 50.3917 | 0 | 0 | 1.3182 | 0.17721 | 1.76 | 16 | 1-exp(-1-D/fRC) | 0.98168
[6 empty] | 0.12 | 0.01 | 16 | 0 | 7.7694 | 0.38847 | 112.9018 | 0.82306 | 49.3838 | 0 | 0 | 1.2919 | 0.19013 | 1.92 | 16 | exp(-1-D/fRC) | 0.018316
[6 empty] | 0.13 | 0.02 | 16 | 0 | 7.934 | 0.3967 | 121.1324 | 0.8066 | 48.3961 | 0 | 0 | 1.266 | 0.20279 | 2.08 | 16
[6 empty] | 0.14 | 0.03 | 16 | 0 | 8.0953 | 0.40476 | 129.1985 | 0.79047 | 47.4282 | 0 | 0 | 1.2407 | 0.2152 | 2.24 | 16 | 1-exp(-D/fRC)*exp(-1-D/fRC) | 0.93281
[6 empty] | 0.15 | 0.04 | 16 | 0 | 8.2534 | 0.41267 | 137.1032 | 0.77466 | 46.4797 | 0 | 0 | 1.2159 | 0.22736 | 2.4 | 16 | (1-exp(-D/fRC))*(exp(-1-D/fRC)) | 0.99909
[6 empty] | 0.16 | 0.05 | 16 | 0 | 8.4083 | 0.42042 | 144.8498 | 0.75917 | 45.5501 | 0 | 0 | 1.1916 | 0.23927 | 2.56 | 16
[6 empty] | 0.17 | 0.06 | 16 | 0 | 8.5602 | 0.42801 | 152.4415 | 0.74398 | 44.6391 | 0 | 0 | 1.1678 | 0.25095 | 2.72 | 16 | VT | 0.51351
[6 empty] | 0.18 | 0.07 | 16 | 0 | 8.709 | 0.43545 | 159.8813 | 0.7291 | 43.7463 | 0 | 0 | 1.1444 | 0.26239 | 2.88 | 16
[6 empty] | 0.19 | 0.08 | 16 | 0 | 8.8548 | 0.44274 | 167.1723 | 0.71452 | 42.8714 | 0 | 0 | 1.1215 | 0.27361 | 3.04 | 16
[6 empty] | 0.2 | 0.09 | 16 | 0 | 8.9977 | 0.44988 | 174.3176 | 0.70023 | 42.0139 | 0 | 0 | 1.0991 | 0.2846 | 3.2 | 16
[6 empty] | 0.21 | 0.1 | 16 | 0 | 9.1377 | 0.45689 | 181.3199 | 0.68623 | 41.1737 | 0 | 0 | 1.0771 | 0.29537 | 3.36 | 16
[6 empty] | 0.22 | 0.11 | 16 | 0 | 9.275 | 0.46375 | 188.1822 | 0.6725 | 40.3502 | 0 | 0 | 1.0556 | 0.30593 | 3.52 | 16
[6 empty] | 0.23 | 0.12 | 16 | 0 | 9.4095 | 0.47047 | 194.9072 | 0.65905 | 39.5432 | 0 | 0 | 1.0344 | 0.31627 | 3.68 | 16
[6 empty] | 0.24 | 0.13 | 16 | 0 | 9.5413 | 0.47706 | 201.4977 | 0.64587 | 38.7523 | 0 | 0 | 1.0138 | 0.32641 | 3.84 | 16
[6 empty] | 0.25 | 0.14 | 16 | 0 | 9.6705 | 0.48352 | 207.9565 | 0.63295 | 37.9773 | 0 | 0 | 0.99349 | 0.33634 | 4 | 16
[6 empty] | 0.26 | 0.15 | 16 | 0 | 9.797 | 0.48985 | 214.286 | 0.6203 | 37.2177 | 0 | 0 | 0.97362 | 0.34608 | 4.16 | 16
[6 empty] | 0.27 | 0.16 | 16 | 0 | 9.9211 | 0.49606 | 220.4889 | 0.60789 | 36.4734 | 0 | 0 | 0.95414 | 0.35562 | 4.32 | 16
[6 empty] | 0.28 | 0.17 | 16 | 0 | 10.0427 | 0.50213 | 226.5678 | 0.59573 | 35.7439 | 0 | 0 | 0.93506 | 0.36497 | 4.48 | 16
[6 empty] | 0.29 | 0.18 | 16 | 0 | 10.1618 | 0.50809 | 232.5252 | 0.58382 | 35.029 | 0 | 0 | 0.91636 | 0.37414 | 4.64 | 16
[6 empty] | 0.3 | 0.19 | 16 | 0 | 10.2786 | 0.51393 | 238.3633 | 0.57214 | 34.3284 | 0 | 0 | 0.89803 | 0.38312 | 4.8 | 16
[6 empty] | 0.31 | 0.2 | 16 | 0 | 10.393 | 0.51965 | 244.0847 | 0.5607 | 33.6419 | 0 | 0 | 0.88007 | 0.39192 | 4.96 | 16
[6 empty] | 0.32 | 0.21 | 16 | 0 | 10.5052 | 0.52526 | 249.6917 | 0.54948 | 32.969 | 0 | 0 | 0.86247 | 0.40054 | 5.12 | 16
[6 empty] | 0.33 | 0.22 | 16 | 0 | 10.6151 | 0.53075 | 255.1866 | 0.53849 | 32.3097 | 0 | 0 | 0.84522 | 0.40899 | 5.28 | 16
[6 empty] | 0.34 | 0.23 | 16 | 0 | 10.7228 | 0.53614 | 260.5715 | 0.52772 | 31.6635 | 0 | 0 | 0.82832 | 0.41728 | 5.44 | 16
[6 empty] | 0.35 | 0.24 | 16 | 0 | 10.8283 | 0.54142 | 265.8487 | 0.51717 | 31.0302 | 0 | 0 | 0.81175 | 0.42539 | 5.6 | 16
[6 empty] | 0.36 | 0.25 | 16 | 0 | 10.9317 | 0.54659 | 271.0204 | 0.50683 | 30.4096 | 0 | 0 | 0.79551 | 0.43335 | 5.76 | 16
[6 empty] | 0.37 | 0.26 | 16 | 0 | 11.0331 | 0.55166 | 276.0887 | 0.49669 | 29.8014 | 0 | 0 | 0.7796 | 0.44114 | 5.92 | 16
[6 empty] | 0.38 | 0.27 | 16 | 0 | 11.1324 | 0.55662 | 281.0556 | 0.48676 | 29.2054 | 0 | 0 | 0.76401 | 0.44879 | 6.08 | 16
[6 empty] | 0.39 | 0.28 | 16 | 0 | 11.2298 | 0.56149 | 285.9232 | 0.47702 | 28.6213 | 0 | 0 | 0.74873 | 0.45627 | 6.24 | 16
[6 empty] | 0.4 | 0.29 | 16 | 0 | 11.3252 | 0.56626 | 290.6934 | 0.46748 | 28.0488 | 0 | 0 | 0.73376 | 0.46361 | 6.4 | 16
[6 empty] | 0.41 | 0.3 | 16 | 0 | 11.4187 | 0.57093 | 295.3682 | 0.45813 | 27.4879 | 0 | 0 | 0.71908 | 0.4708 | 6.56 | 16
[6 empty] | 0.42 | 0.31 | 16 | 0 | 11.5103 | 0.57552 | 299.9495 | 0.44897 | 26.9381 | 0 | 0 | 0.7047 | 0.47785 | 6.72 | 16
[6 empty] | 0.43 | 0.32 | 16 | 0 | 11.6001 | 0.58001 | 304.4392 | 0.43999 | 26.3993 | 0 | 0 | 0.69061 | 0.48475 | 6.88 | 16
[6 empty] | 0.44 | 0.33 | 16 | 0 | 11.6881 | 0.58441 | 308.8391 | 0.43119 | 25.8714 | 0 | 0 | 0.67679 | 0.49152 | 7.04 | 16
[6 empty] | 0.45 | 0.34 | 16 | 0 | 11.7743 | 0.58872 | 313.151 | 0.42257 | 25.3539 | 0 | 0 | 0.66326 | 0.49815 | 7.2 | 16
[6 empty] | 0.46 | 0.35 | 16 | 0 | 11.8589 | 0.59294 | 317.3766 | 0.41411 | 24.8468 | 0 | 0 | 0.64999 | 0.50465 | 7.36 | 16
[6 empty] | 0.47 | 0.36 | 16 | 0 | 11.9417 | 0.59708 | 321.5178 | 0.40583 | 24.3499 | 0 | 0 | 0.63699 | 0.51102 | 7.52 | 16
[6 empty] | 0.48 | 0.37 | 16 | 0 | 12.0228 | 0.60114 | 325.5761 | 0.39772 | 23.8629 | 0 | 0 | 0.62425 | 0.51727 | 7.68 | 16
[6 empty] | 0.49 | 0.38 | 16 | 0 | 12.1024 | 0.60512 | 329.5532 | 0.38976 | 23.3857 | 0 | 0 | 0.61177 | 0.52338 | 7.84 | 16
[6 empty] | 0.5 | 0.39 | 16 | 0 | 12.1803 | 0.60902 | 333.4508 | 0.38197 | 22.9179 | 0 | 0 | 0.59953 | 0.52938 | 8 | 16
[6 empty] | 0.51 | 0.4 | 16 | 0 | 12.2567 | 0.61284 | 337.2705 | 0.37433 | 22.4596 | 0 | 0 | 0.58754 | 0.53526 | 8.16 | 16
[6 empty] | 0.52 | 0.41 | 16 | 0 | 12.3316 | 0.61658 | 341.0138 | 0.36684 | 22.0104 | 0 | 0 | 0.57579 | 0.54101 | 8.32 | 16
[6 empty] | 0.53 | 0.42 | 16 | 0 | 12.405 | 0.62025 | 344.6822 | 0.3595 | 21.5702 | 0 | 0 | 0.56428 | 0.54666 | 8.48 | 16
[6 empty] | 0.54 | 0.43 | 16 | 0 | 12.4769 | 0.62384 | 348.2772 | 0.35231 | 21.1388 | 0 | 0 | 0.55299 | 0.55219 | 8.64 | 16
[6 empty] | 0.55 | 0.44 | 16 | 0 | 12.5473 | 0.62737 | 351.8003 | 0.34527 | 20.716 | 0 | 0 | 0.54193 | 0.55761 | 8.8 | 16
[6 empty] | 0.56 | 0.45 | 16 | 0 | 12.6164 | 0.63082 | 355.253 | 0.33836 | 20.3017 | 0 | 0 | 0.53109 | 0.56292 | 8.96 | 16
[6 empty] | 0.57 | 0.46 | 16 | 0 | 12.6841 | 0.6342 | 358.6366 | 0.33159 | 19.8957 | 0 | 0 | 0.52047 | 0.56812 | 9.12 | 16
[6 empty] | 0.58 | 0.47 | 16 | 0 | 12.7504 | 0.63752 | 361.9525 | 0.32496 | 19.4977 | 0 | 0 | 0.51006 | 0.57322 | 9.28 | 16
[6 empty] | 0.59 | 0.48 | 16 | 0 | 12.8154 | 0.64077 | 365.2022 | 0.31846 | 19.1078 | 0 | 0 | 0.49986 | 0.57822 | 9.44 | 16
[6 empty] | 0.6 | 0.49 | 16 | 0 | 12.8791 | 0.64395 | 368.3868 | 0.31209 | 18.7256 | 0 | 0 | 0.48986 | 0.58312 | 9.6 | 16
[6 empty] | 0.61 | 0.5 | 16 | 0 | 12.9415 | 0.64707 | 371.5077 | 0.30585 | 18.3511 | 0 | 0 | 0.48007 | 0.58792 | 9.76 | 16
[6 empty] | 0.62 | 0.51 | 16 | 0 | 13.0027 | 0.65013 | 374.5662 | 0.29973 | 17.9841 | 0 | 0 | 0.47046 | 0.59262 | 9.92 | 16
[6 empty] | 0.63 | 0.52 | 16 | 0 | 13.0626 | 0.65313 | 377.5636 | 0.29374 | 17.6244 | 0 | 0 | 0.46105 | 0.59723 | 10.08 | 16
[6 empty] | 0.64 | 0.53 | 16 | 0 | 13.1213 | 0.65607 | 380.501 | 0.28787 | 17.2719 | 0 | 0 | 0.45183 | 0.60175 | 10.24 | 16
[6 empty] | 0.65 | 0.54 | 16 | 0 | 13.1789 | 0.65895 | 383.3796 | 0.28211 | 16.9265 | 0 | 0 | 0.4428 | 0.60618 | 10.4 | 16
[6 empty] | 0.66 | 0.55 | 16 | 0 | 13.2353 | 0.66177 | 386.2007 | 0.27647 | 16.588 | 0 | 0 | 0.43394 | 0.61052 | 10.56 | 16
[6 empty] | 0.67 | 0.56 | 16 | 0 | 13.2906 | 0.66453 | 388.9654 | 0.27094 | 16.2562 | 0 | 0 | 0.42526 | 0.61477 | 10.72 | 16
[6 empty] | 0.68 | 0.57 | 16 | 0 | 13.3448 | 0.66724 | 391.6748 | 0.26552 | 15.9311 | 0 | 0 | 0.41676 | 0.61894 | 10.88 | 16
[6 empty] | 0.69 | 0.58 | 16 | 0 | 13.3979 | 0.6699 | 394.3299 | 0.26021 | 15.6124 | 0 | 0 | 0.40842 | 0.62302 | 11.04 | 16
[6 empty] | 0.7 | 0.59 | 16 | 0 | 13.45 | 0.6725 | 396.932 | 0.255 | 15.3002 | 0 | 0 | 0.40025 | 0.62703 | 11.2 | 16
[6 empty] | 0.71 | 0.6 | 16 | 0 | 13.501 | 0.67505 | 399.482 | 0.2499 | 14.9942 | 0 | 0 | 0.39225 | 0.63095 | 11.36 | 16
[6 empty] | 0.72 | 0.61 | 16 | 0 | 13.5509 | 0.67755 | 401.9811 | 0.24491 | 14.6943 | 0 | 0 | 0.3844 | 0.63479 | 11.52 | 16
[6 empty] | 0.73 | 0.62 | 16 | 0 | 13.5999 | 0.68 | 404.4301 | 0.24001 | 14.4004 | 0 | 0 | 0.37672 | 0.63856 | 11.68 | 16
[6 empty] | 0.74 | 0.63 | 16 | 0 | 13.6479 | 0.6824 | 406.8302 | 0.23521 | 14.1124 | 0 | 0 | 0.36918 | 0.64225 | 11.84 | 16
[6 empty] | 0.75 | 0.64 | 16 | 0 | 13.695 | 0.68475 | 409.1823 | 0.2305 | 13.8302 | 0 | 0 | 0.3618 | 0.64587 | 12 | 16
[6 empty] | 0.76 | 0.65 | 16 | 0 | 13.7411 | 0.68705 | 411.4873 | 0.22589 | 13.5536 | 0 | 0 | 0.35456 | 0.64942 | 12.16 | 16
[6 empty] | 0.77 | 0.66 | 16 | 0 | 13.7863 | 0.68931 | 413.7462 | 0.22137 | 13.2825 | 0 | 0 | 0.34747 | 0.65289 | 12.32 | 16
[6 empty] | 0.78 | 0.67 | 16 | 0 | 13.8305 | 0.69153 | 415.96 | 0.21695 | 13.0168 | 0 | 0 | 0.34052 | 0.6563 | 12.48 | 16
[6 empty] | 0.79 | 0.68 | 16 | 0 | 13.8739 | 0.6937 | 418.1294 | 0.21261 | 12.7565 | 0 | 0 | 0.33371 | 0.65963 | 12.64 | 16
[6 empty] | 0.8 | 0.69 | 16 | 0 | 13.9164 | 0.69582 | 420.2555 | 0.20836 | 12.5014 | 0 | 0 | 0.32704 | 0.6629 | 12.8 | 16
[6 empty] | 0.81 | 0.7 | 16 | 0 | 13.9581 | 0.69791 | 422.3391 | 0.20419 | 12.2514 | 0 | 0 | 0.3205 | 0.66611 | 12.96 | 16
[6 empty] | 0.82 | 0.71 | 16 | 0 | 13.9989 | 0.69995 | 424.381 | 0.20011 | 12.0063 | 0 | 0 | 0.31409 | 0.66925 | 13.12 | 16
[6 empty] | 0.83 | 0.72 | 16 | 0 | 14.039 | 0.70195 | 426.382 | 0.1961 | 11.7662 | 0 | 0 | 0.3078 | 0.67233 | 13.28 | 16
[6 empty] | 0.84 | 0.73 | 16 | 0 | 14.0782 | 0.70391 | 428.3431 | 0.19218 | 11.5309 | 0 | 0 | 0.30165 | 0.67534 | 13.44 | 16
[6 empty] | 0.85 | 0.74 | 16 | 0 | 14.1166 | 0.70583 | 430.2649 | 0.18834 | 11.3003 | 0 | 0 | 0.29561 | 0.6783 | 13.6 | 16
[6 empty] | 0.86 | 0.75 | 16 | 0 | 14.1543 | 0.70771 | 432.1483 | 0.18457 | 11.0743 | 0 | 0 | 0.2897 | 0.6812 | 13.76 | 16
[6 empty] | 0.87 | 0.76 | 16 | 0 | 14.1912 | 0.70956 | 433.994 | 0.18088 | 10.8528 | 0 | 0 | 0.28391 | 0.68404 | 13.92 | 16
[6 empty] | 0.88 | 0.77 | 16 | 0 | 14.2274 | 0.71137 | 435.8028 | 0.17726 | 10.6357 | 0 | 0 | 0.27823 | 0.68682 | 14.08 | 16
[6 empty] | 0.89 | 0.78 | 16 | 0 | 14.2628 | 0.71314 | 437.5754 | 0.17372 | 10.423 | 0 | 0 | 0.27267 | 0.68954 | 14.24 | 16
[6 empty] | 0.9 | 0.79 | 16 | 0 | 14.2976 | 0.71488 | 439.3125 | 0.17024 | 10.2145 | 0 | 0 | 0.26721 | 0.69222 | 14.4 | 16
[6 empty] | 0.91 | 0.8 | 16 | 0 | 14.3316 | 0.71658 | 441.015 | 0.16684 | 10.0102 | 0 | 0 | 0.26187 | 0.69484 | 14.56 | 16
[6 empty] | 0.92 | 0.81 | 16 | 0 | 14.365 | 0.71825 | 442.6833 | 0.1635 | 9.81 | 0 | 0 | 0.25663 | 0.6974 | 14.72 | 16
[6 empty] | 0.93 | 0.82 | 16 | 0 | 14.3977 | 0.71988 | 444.3184 | 0.16023 | 9.6138 | 0 | 0 | 0.2515 | 0.69992 | 14.88 | 16
[6 empty] | 0.94 | 0.83 | 16 | 0 | 14.4297 | 0.72149 | 445.9207 | 0.15703 | 9.4216 | 0 | 0 | 0.24647 | 0.70238 | 15.04 | 16
[6 empty] | 0.95 | 0.84 | 16 | 0 | 14.4611 | 0.72306 | 447.4909 | 0.15389 | 9.2331 | 0 | 0 | 0.24154 | 0.7048 | 15.2 | 16
[6 empty] | 0.96 | 0.85 | 16 | 0 | 14.4919 | 0.7246 | 449.0298 | 0.15081 | 9.0485 | 0 | 0 | 0.23671 | 0.70716 | 15.36 | 16
[6 empty] | 0.97 | 0.86 | 16 | 0 | 14.5221 | 0.7261 | 450.5379 | 0.14779 | 8.8675 | 0 | 0 | 0.23197 | 0.70948 | 15.52 | 16
[6 empty] | 0.98 | 0.87 | 16 | 0 | 14.5516 | 0.72758 | 452.0158 | 0.14484 | 8.6901 | 0 | 0 | 0.22733 | 0.71176 | 15.68 | 16
[6 empty] | 0.99 | 0.88 | 16 | 0 | 14.5806 | 0.72903 | 453.4641 | 0.14194 | 8.5163 | 0 | 0 | 0.22279 | 0.71399 | 15.84 | 16
[6 empty] | 1 | 0.89 | 16 | 0 | 14.609 | 0.73045 | 454.8835 | 0.1391 | 8.346 | 0 | 0 | 0.21833 | 0.71617 | 16 | 16
[6 empty] | 1.01 | 0.9 | 16 | 0 | 14.6368 | 0.73184 | 456.2745 | 0.13632 | 8.1791 | 0 | 0 | 0.21397 | 0.71831 | 16.16 | 16
[6 empty] | 1.02 | 0.91 | 16 | 0 | 14.6641 | 0.7332 | 457.6377 | 0.13359 | 8.0155 | 0 | 0 | 0.20969 | 0.7204 | 16.32 | 16
[6 empty] | 1.03 | 0.92 | 16 | 0 | 14.6908 | 0.73454 | 458.9736 | 0.13092 | 7.8552 | 0 | 0 | 0.20549 | 0.72246 | 16.48 | 16
[6 empty] | 1.04 | 0.93 | 16 | 0 | 14.717 | 0.73585 | 460.2828 | 0.1283 | 7.6981 | 0 | 0 | 0.20138 | 0.72447 | 16.64 | 16
[6 empty] | 1.05 | 0.94 | 16 | 0 | 14.7426 | 0.73713 | 461.5658 | 0.12574 | 7.5441 | 0 | 0 | 0.19735 | 0.72645 | 16.8 | 16
[6 empty] | 1.06 | 0.95 | 16 | 0 | 14.7678 | 0.73839 | 462.8232 | 0.12322 | 7.3933 | 0 | 0 | 0.19341 | 0.72838 | 16.96 | 16
[6 empty] | 1.07 | 0.96 | 16 | 0 | 14.7924 | 0.73962 | 464.0554 | 0.12076 | 7.2454 | 0 | 0 | 0.18954 | 0.73028 | 17.12 | 16
[6 empty] | 1.08 | 0.97 | 16 | 0 | 14.8166 | 0.74083 | 465.263 | 0.11834 | 7.1005 | 0 | 0 | 0.18575 | 0.73213 | 17.28 | 16
[6 empty] | 1.09 | 0.98 | 16 | 0 | 14.8403 | 0.74201 | 466.4464 | 0.11597 | 6.9585 | 0 | 0 | 0.18203 | 0.73395 | 17.44 | 16
[6 empty] | 1.1 | 0.99 | 16 | 0 | 14.8634 | 0.74317 | 467.6061 | 0.11366 | 6.8193 | 0 | 0 | 0.17839 | 0.73574 | 17.6 | 16
[6 empty] | 1.11 | 1 | 16 | 0 | 14.8862 | 0.74431 | 468.7427 | 0.11138 | 6.6829 | 0 | 0 | 0.17483 | 0.73749 | 17.76 | 16
[6 empty] | 1.12 | 1.01 | 16 | 0 | 14.9085 | 0.74542 | 469.8565 | 0.10915 | 6.5493 | 0 | 0 | 0.17133 | 0.7392 | 17.92 | 16
[6 empty] | 1.13 | 1.02 | 16 | 0 | 14.9303 | 0.74651 | 470.948 | 0.10697 | 6.4183 | 0 | 0 | 0.1679 | 0.74088 | 18.08 | 16
[6 empty] | 1.14 | 1.03 | 16 | 0 | 14.9517 | 0.74758 | 472.0178 | 0.10483 | 6.2899 | 0 | 0 | 0.16454 | 0.74252 | 18.24 | 16
[6 empty] | 1.15 | 1.04 | 16 | 0 | 14.9726 | 0.74863 | 473.0661 | 0.10274 | 6.1641 | 0 | 0 | 0.16125 | 0.74414 | 18.4 | 16
[6 empty] | 1.16 | 1.05 | 16 | 0 | 14.9932 | 0.74966 | 474.0934 | 0.10068 | 6.0408 | 0 | 0 | 0.15803 | 0.74572 | 18.56 | 16
[6 empty] | 1.17 | 1.06 | 16 | 0 | 15.0133 | 0.75067 | 475.1002 | 0.098667 | 5.92 | 0 | 0 | 0.15487 | 0.74727 | 18.72 | 16
[6 empty] | 1.18 | 1.07 | 16 | 0 | 15.0331 | 0.75165 | 476.0869 | 0.096694 | 5.8016 | 0 | 0 | 0.15177 | 0.74878 | 18.88 | 16
[6 empty] | 1.19 | 1.08 | 16 | 0 | 15.0524 | 0.75262 | 477.0538 | 0.09476 | 5.6856 | 0 | 0 | 0.14873 | 0.75027 | 19.04 | 16
[6 empty] | 1.2 | 1.09 | 16 | 0 | 15.0714 | 0.75357 | 478.0014 | 0.092864 | 5.5719 | 0 | 0 | 0.14576 | 0.75173 | 19.2 | 16
[6 empty] | 1.21 | 1.1 | 16 | 0 | 15.0899 | 0.7545 | 478.9301 | 0.091007 | 5.4604 | 0 | 0 | 0.14284 | 0.75316 | 19.36 | 16
[6 empty] | 1.22 | 1.11 | 16 | 0 | 15.1081 | 0.75541 | 479.8402 | 0.089187 | 5.3512 | 0 | 0 | 0.13999 | 0.75456 | 19.52 | 16
[6 empty] | 1.23 | 1.12 | 16 | 0 | 15.126 | 0.7563 | 480.732 | 0.087403 | 5.2442 | 0 | 0 | 0.13719 | 0.75593 | 19.68 | 16
[6 empty] | 1.24 | 1.13 | 16 | 0 | 15.1434 | 0.75717 | 481.6061 | 0.085655 | 5.1393 | 0 | 0 | 0.13444 | 0.75727 | 19.84 | 16
[6 empty] | 1.25 | 1.14 | 16 | 0 | 15.1606 | 0.75803 | 482.4626 | 0.083942 | 5.0365 | 0 | 0 | 0.13176 | 0.75859 | 20 | 16
[6 empty] | 1.26 | 1.15 | 16 | 0 | 15.1774 | 0.75887 | 483.302 | 0.082263 | 4.9358 | 0 | 0 | 0.12912 | 0.75988 | 20.16 | 16
[6 empty] | 1.27 | 1.16 | 16 | 0 | 15.1938 | 0.75969 | 484.1247 | 0.080618 | 4.8371 | 0 | 0 | 0.12654 | 0.76115 | 20.32 | 16
[6 empty] | 1.28 | 1.17 | 16 | 0 | 15.2099 | 0.7605 | 484.9308 | 0.079006 | 4.7403 | 0 | 0 | 0.12401 | 0.76239 | 20.48 | 16
[6 empty] | 1.29 | 1.18 | 16 | 0 | 15.2257 | 0.76129 | 485.7209 | 0.077426 | 4.6455 | 0 | 0 | 0.12153 | 0.7636 | 20.64 | 16
[6 empty] | 1.3 | 1.19 | 16 | 0 | 15.2412 | 0.76206 | 486.4952 | 0.075877 | 4.5526 | 0 | 0 | 0.1191 | 0.76479 | 20.8 | 16
[6 empty] | 1.31 | 1.2 | 16 | 0 | 15.2564 | 0.76282 | 487.2539 | 0.07436 | 4.4616 | 0 | 0 | 0.11671 | 0.76596 | 20.96 | 16
[6 empty] | 1.32 | 1.21 | 16 | 0 | 15.2713 | 0.76356 | 487.9975 | 0.072872 | 4.3723 | 0 | 0 | 0.11438 | 0.7671 | 21.12 | 16
[6 empty] | 1.33 | 1.22 | 16 | 0 | 15.2859 | 0.76429 | 488.7262 | 0.071415 | 4.2849 | 0 | 0 | 0.11209 | 0.76823 | 21.28 | 16
[6 empty] | 1.34 | 1.23 | 16 | 0 | 15.3001 | 0.76501 | 489.4404 | 0.069987 | 4.1992 | 0 | 0 | 0.10985 | 0.76932 | 21.44 | 16
[6 empty] | 1.35 | 1.24 | 16 | 0 | 15.3141 | 0.76571 | 490.1403 | 0.068587 | 4.1152 | 0 | 0 | 0.10765 | 0.7704 | 21.6 | 16
[6 empty] | 1.36 | 1.25 | 16 | 0 | 15.3278 | 0.76639 | 490.8261 | 0.067215 | 4.0329 | 0 | 0 | 0.1055 | 0.77146 | 21.76 | 16
[6 empty] | 1.37 | 1.26 | 16 | 0 | 15.3413 | 0.76706 | 491.4983 | 0.065871 | 3.9522 | 0 | 0 | 0.10339 | 0.77249 | 21.92 | 16
[6 empty] | 1.38 | 1.27 | 16 | 0 | 15.3545 | 0.76772 | 492.157 | 0.064553 | 3.8732 | 0 | 0 | 0.10132 | 0.7735 | 22.08 | 16
[6 empty] | 1.39 | 1.28 | 16 | 0 | 15.3674 | 0.76837 | 492.8025 | 0.063262 | 3.7957 | 0 | 0 | 0.099297 | 0.7745 | 22.24 | 16
[6 empty] | 1.4 | 1.29 | 16 | 0 | 15.38 | 0.769 | 493.4351 | 0.061997 | 3.7198 | 0 | 0 | 0.097311 | 0.77547 | 22.4 | 16
[6 empty] | 1.41 | 1.3 | 16 | 0 | 15.3924 | 0.76962 | 494.0551 | 0.060757 | 3.6454 | 0 | 0 | 0.095364 | 0.77642 | 22.56 | 16
[6 empty] | 1.42 | 1.31 | 16 | 0 | 15.4046 | 0.77023 | 494.6627 | 0.059542 | 3.5725 | 0 | 0 | 0.093457 | 0.77736 | 22.72 | 16
[6 empty] | 1.43 | 1.32 | 16 | 0 | 15.4165 | 0.77082 | 495.2581 | 0.058351 | 3.5011 | 0 | 0 | 0.091588 | 0.77827 | 22.88 | 16
[6 empty] | 1.44 | 1.33 | 16 | 0 | 15.4282 | 0.77141 | 495.8416 | 0.057184 | 3.431 | 0 | 0 | 0.089756 | 0.77917 | 23.04 | 16
[6 empty] | 1.45 | 1.34 | 16 | 0 | 15.4396 | 0.77198 | 496.4135 | 0.05604 | 3.3624 | 0 | 0 | 0.087961 | 0.78005 | 23.2 | 16
[6 empty] | 1.46 | 1.35 | 16 | 0 | 15.4508 | 0.77254 | 496.9739 | 0.05492 | 3.2952 | 0 | 0 | 0.086202 | 0.78091 | 23.36 | 16
[6 empty] | 1.47 | 1.36 | 16 | 0 | 15.4618 | 0.77309 | 497.5231 | 0.053821 | 3.2293 | 0 | 0 | 0.084478 | 0.78176 | 23.52 | 16
[6 empty] | 1.48 | 1.37 | 16 | 0 | 15.4726 | 0.77363 | 498.0613 | 0.052745 | 3.1647 | 0 | 0 | 0.082788 | 0.78258 | 23.68 | 16
[6 empty] | 1.49 | 1.38 | 16 | 0 | 15.4831 | 0.77416 | 498.5887 | 0.05169 | 3.1014 | 0 | 0 | 0.081132 | 0.7834 | 23.84 | 16
[6 empty] | 1.5 | 1.39 | 16 | 0 | 15.4934 | 0.77467 | 499.1056 | 0.050656 | 3.0394 | 0 | 0 | 0.07951 | 0.78419 | 24 | 16
[6 empty] | 1.51 | 1.4 | 5 | 0 | 15.5036 | 0.77518 | 499.6122 | -1.0504 | -63.0214 | 0 | 0 | -0.5152 | 0.77904 | 24.05 | 15.9272
[6 empty] | 1.52 | 1.41 | 5 | 0 | 15.2935 | 0.76467 | 489.1086 | -1.0293 | -61.761 | 0 | 0 | -0.5049 | 0.77399 | 24.1 | 15.8553
[6 empty] | 1.53 | 1.42 | 5 | 0 | 15.0876 | 0.75438 | 478.8151 | -1.0088 | -60.5258 | 0 | 0 | -0.4948 | 0.76904 | 24.15 | 15.7843
[6 empty] | 1.54 | 1.43 | 5 | 0 | 14.8859 | 0.74429 | 468.7275 | -0.98859 | -59.3153 | 0 | 0 | -0.4849 | 0.76419 | 24.2 | 15.7143
[6 empty] | 1.55 | 1.44 | 5 | 0 | 14.6882 | 0.73441 | 458.8416 | -0.96882 | -58.129 | 0 | 0 | -0.4752 | 0.75944 | 24.25 | 15.6452
[6 empty] | 1.56 | 1.45 | 5 | 0 | 14.4944 | 0.72472 | 449.1534 | -0.94944 | -56.9664 | 0 | 0 | -0.4657 | 0.75478 | 24.3 | 15.5769
[6 empty] | 1.57 | 1.46 | 5 | 0 | 14.3045 | 0.71523 | 439.6591 | -0.93045 | -55.827 | 0 | 0 | -0.45639 | 0.75022 | 24.35 | 15.5096
[6 empty] | 1.58 | 1.47 | 5 | 0 | 14.1184 | 0.70592 | 430.3545 | -0.91184 | -54.7105 | 0 | 0 | -0.44726 | 0.74575 | 24.4 | 15.443
[6 empty] | 1.59 | 1.48 | 5 | 0 | 13.936 | 0.6968 | 421.2361 | -0.8936 | -53.6163 | 0 | 0 | -0.43831 | 0.74136 | 24.45 | 15.3774
[6 empty] | 1.6 | 1.49 | 5 | 0 | 13.7573 | 0.68787 | 412.3001 | -0.87573 | -52.544 | 0 | 0 | -0.42955 | 0.73707 | 24.5 | 15.3125
[6 empty] | 1.61 | 1.5 | 5 | 0 | 13.5822 | 0.67911 | 403.5427 | -0.85822 | -51.4931 | 0 | 0 | -0.42096 | 0.73286 | 24.55 | 15.2484
[6 empty] | 1.62 | 1.51 | 5 | 0 | 13.4105 | 0.67053 | 394.9606 | -0.84105 | -50.4632 | 0 | 0 | -0.41254 | 0.72873 | 24.6 | 15.1852
[6 empty] | 1.63 | 1.52 | 5 | 0 | 13.2423 | 0.66212 | 386.55 | -0.82423 | -49.454 | 0 | 0 | -0.40429 | 0.72469 | 24.65 | 15.1227
[6 empty] | 1.64 | 1.53 | 5 | 0 | 13.0775 | 0.65387 | 378.3077 | -0.80775 | -48.4649 | 0 | 0 | -0.3962 | 0.72073 | 24.7 | 15.061
[6 empty] | 1.65 | 1.54 | 5 | 0 | 12.9159 | 0.6458 | 370.2302 | -0.79159 | -47.4956 | 0 | 0 | -0.38828 | 0.71685 | 24.75 | 15
[6 empty] | 1.66 | 1.55 | 5 | 0 | 12.7576 | 0.63788 | 362.3143 | -0.77576 | -46.5457 | 0 | 0 | -0.38051 | 0.71304 | 24.8 | 14.9398
[6 empty] | 1.67 | 1.56 | 5 | 0 | 12.6025 | 0.63012 | 354.5567 | -0.76025 | -45.6148 | 0 | 0 | -0.3729 | 0.70931 | 24.85 | 14.8802
[6 empty] | 1.68 | 1.57 | 5 | 0 | 12.4504 | 0.62252 | 346.9542 | -0.74504 | -44.7025 | 0 | 0 | -0.36544 | 0.70566 | 24.9 | 14.8214
[6 empty] | 1.69 | 1.58 | 5 | 0 | 12.3014 | 0.61507 | 339.5038 | -0.73014 | -43.8084 | 0 | 0 | -0.35813 | 0.70208 | 24.95 | 14.7633
[6 empty] | 1.7 | 1.59 | 5 | 0 | 12.1554 | 0.60777 | 332.2024 | -0.71554 | -42.9322 | 0 | 0 | -0.35097 | 0.69857 | 25 | 14.7059
[6 empty] | 1.71 | 1.6 | 5 | 0 | 12.0123 | 0.60061 | 325.047 | -0.70123 | -42.0736 | 0 | 0 | -0.34395 | 0.69513 | 25.05 | 14.6491
[6 empty] | 1.72 | 1.61 | 5 | 0 | 11.872 | 0.5936 | 318.0348 | -0.6872 | -41.2321 | 0 | 0 | -0.33707 | 0.69176 | 25.1 | 14.593
[6 empty] | 1.73 | 1.62 | 5 | 0 | 11.7346 | 0.58673 | 311.1627 | -0.67346 | -40.4075 | 0 | 0 | -0.33033 | 0.68845 | 25.15 | 14.5376
[6 empty] | 1.74 | 1.63 | 5 | 0 | 11.5999 | 0.57999 | 304.4282 | -0.65999 | -39.5993 | 0 | 0 | -0.32372 | 0.68522 | 25.2 | 14.4828
[6 empty] | 1.75 | 1.64 | 5 | 0 | 11.4679 | 0.57339 | 297.8283 | -0.64679 | -38.8074 | 0 | 0 | -0.31725 | 0.68204 | 25.25 | 14.4286
[6 empty] | 1.76 | 1.65 | 5 | 0 | 11.3385 | 0.56693 | 291.3604 | -0.63385 | -38.0312 | 0 | 0 | -0.31091 | 0.67893 | 25.3 | 14.375
[6 empty] | 1.77 | 1.66 | 5 | 0 | 11.2118 | 0.56059 | 285.0218 | -0.62118 | -37.2706 | 0 | 0 | -0.30469 | 0.67589 | 25.35 | 14.322
[6 empty] | 1.78 | 1.67 | 5 | 0 | 11.0875 | 0.55438 | 278.8101 | -0.60875 | -36.5252 | 0 | 0 | -0.29859 | 0.6729 | 25.4 | 14.2697
[6 empty] | 1.79 | 1.68 | 5 | 0 | 10.9658 | 0.54829 | 272.7226 | -0.59658 | -35.7947 | 0 | 0 | -0.29262 | 0.66998 | 25.45 | 14.2179
[6 empty] | 1.8 | 1.69 | 5 | 0 | 10.8465 | 0.54232 | 266.7568 | -0.58465 | -35.0788 | 0 | 0 | -0.28677 | 0.66711 | 25.5 | 14.1667
[6 empty] | 1.81 | 1.7 | 5 | 0 | 10.7295 | 0.53648 | 260.9103 | -0.57295 | -34.3772 | 0 | 0 | -0.28103 | 0.6643 | 25.55 | 14.116
[6 empty] | 1.82 | 1.71 | 5 | 0 | 10.6149 | 0.53075 | 255.1808 | -0.56149 | -33.6897 | 0 | 0 | -0.27541 | 0.66154 | 25.6 | 14.0659
[6 empty] | 1.83 | 1.72 | 5 | 0 | 10.5026 | 0.52513 | 249.5658 | -0.55026 | -33.0159 | 0 | 0 | -0.2699 | 0.65884 | 25.65 | 14.0164
[6 empty] | 1.84 | 1.73 | 5 | 0 | 10.3926 | 0.51963 | 244.0632 | -0.53926 | -32.3555 | 0 | 0 | -0.26451 | 0.6562 | 25.7 | 13.9674
[6 empty] | 1.85 | 1.74 | 5 | 0 | 10.2847 | 0.51424 | 238.6706 | -0.52847 | -31.7084 | 0 | 0 | -0.25922 | 0.65361 | 25.75 | 13.9189
[6 empty] | 1.86 | 1.75 | 5 | 0 | 10.179 | 0.50895 | 233.3859 | -0.5179 | -31.0743 | 0 | 0 | -0.25403 | 0.65107 | 25.8 | 13.871
[6 empty] | 1.87 | 1.76 | 5 | 0 | 10.0755 | 0.50377 | 228.2068 | -0.50755 | -30.4528 | 0 | 0 | -0.24895 | 0.64858 | 25.85 | 13.8235
[6 empty] | 1.88 | 1.77 | 5 | 0 | 9.974 | 0.4987 | 223.1314 | -0.4974 | -29.8437 | 0 | 0 | -0.24397 | 0.64614 | 25.9 | 13.7766
[6 empty] | 1.89 | 1.78 | 5 | 0 | 9.8745 | 0.49372 | 218.1574 | -0.48745 | -29.2468 | 0 | 0 | -0.23909 | 0.64375 | 25.95 | 13.7302
[6 empty] | 1.9 | 1.79 | 5 | 0 | 9.777 | 0.48885 | 213.2829 | -0.4777 | -28.6619 | 0 | 0 | -0.23431 | 0.6414 | 26 | 13.6842
[6 empty] | 1.91 | 1.8 | 5 | 0 | 9.6814 | 0.48407 | 208.5059 | -0.46814 | -28.0887 | 0 | 0 | -0.22962 | 0.63911 | 26.05 | 13.6387
[6 empty] | 1.92 | 1.81 | 5 | 0 | 9.5878 | 0.47939 | 203.8245 | -0.45878 | -27.5269 | 0 | 0 | -0.22503 | 0.63686 | 26.1 | 13.5938
[6 empty] | 1.93 | 1.82 | 5 | 0 | 9.4961 | 0.4748 | 199.2367 | -0.44961 | -26.9764 | 0 | 0 | -0.22053 | 0.63465 | 26.15 | 13.5492
[6 empty] | 1.94 | 1.83 | 5 | 0 | 9.4061 | 0.47031 | 194.7406 | -0.44061 | -26.4368 | 0 | 0 | -0.21612 | 0.63249 | 26.2 | 13.5052
[6 empty] | 1.95 | 1.84 | 5 | 0 | 9.318 | 0.4659 | 190.3345 | -0.4318 | -25.9081 | 0 | 0 | -0.2118 | 0.63037 | 26.25 | 13.4615
[6 empty] | 1.96 | 1.85 | 5 | 0 | 9.2317 | 0.46158 | 186.0165 | -0.42317 | -25.3899 | 0 | 0 | -0.20756 | 0.6283 | 26.3 | 13.4184
[6 empty] | 1.97 | 1.86 | 5 | 0 | 9.147 | 0.45735 | 181.7848 | -0.4147 | -24.8821 | 0 | 0 | -0.20341 | 0.62626 | 26.35 | 13.3756
[6 empty] | 1.98 | 1.87 | 5 | 0 | 9.0641 | 0.4532 | 177.6378 | -0.40641 | -24.3845 | 0 | 0 | -0.19934 | 0.62427 | 26.4 | 13.3333
[6 empty] | 1.99 | 1.88 | 5 | 0 | 8.9828 | 0.44914 | 173.5737 | -0.39828 | -23.8968 | 0 | 0 | -0.19536 | 0.62232 | 26.45 | 13.2915
[6 empty] | 2 | 1.89 | 5 | 0 | 8.9031 | 0.44516 | 169.5909 | -0.39031 | -23.4189 | 0 | 0 | -0.19145 | 0.6204 | 26.5 | 13.25
[6 empty] | 2.01 | 1.9 | 5 | 0 | 8.8251 | 0.44125 | 165.6878 | -0.38251 | -22.9505 | 0 | 0 | -0.18762 | 0.61853 | 26.55 | 13.209
[6 empty] | 2.02 | 1.91 | 5 | 0 | 8.7486 | 0.43743 | 161.8627 | -0.37486 | -22.4915 | 0 | 0 | -0.18387 | 0.61669 | 26.6 | 13.1683
[6 empty] | 2.03 | 1.92 | 5 | 0 | 8.6736 | 0.43368 | 158.1141 | -0.36736 | -22.0417 | 0 | 0 | -0.18019 | 0.61488 | 26.65 | 13.1281
[6 empty] | 2.04 | 1.93 | 5 | 0 | 8.6001 | 0.43001 | 154.4405 | -0.36001 | -21.6008 | 0 | 0 | -0.17659 | 0.61312 | 26.7 | 13.0882
[6 empty] | 2.05 | 1.94 | 5 | 0 | 8.5281 | 0.42641 | 150.8404 | -0.35281 | -21.1688 | 0 | 0 | -0.17305 | 0.61139 | 26.75 | 13.0488
[6 empty] | 2.06 | 1.95 | 5 | 0 | 8.4576 | 0.42288 | 147.3122 | -0.34576 | -20.7454 | 0 | 0 | -0.16959 | 0.60969 | 26.8 | 13.0097
[6 empty] | 2.07 | 1.96 | 5 | 0 | 8.3884 | 0.41942 | 143.8547 | -0.33884 | -20.3305 | 0 | 0 | -0.1662 | 0.60803 | 26.85 | 12.971
[6 empty] | 2.08 | 1.97 | 5 | 0 | 8.3207 | 0.41603 | 140.4662 | -0.33207 | -19.9239 | 0 | 0 | -0.16288 | 0.6064 | 26.9 | 12.9327
[6 empty] | 2.09 | 1.98 | 5 | 0 | 8.2542 | 0.41271 | 137.1456 | -0.32542 | -19.5254 | 0 | 0 | -0.15962 | 0.60481 | 26.95 | 12.8947
[6 empty] | 2.1 | 1.99 | 5 | 0 | 8.1892 | 0.40946 | 133.8913 | -0.31892 | -19.1349 | 0 | 0 | -0.15643 | 0.60324 | 27 | 12.8571
[6 empty] | 2.11 | 2 | 5 | 0 | 8.1254 | 0.40627 | 130.7022 | -0.31254 | -18.7522 | 0 | 0 | -0.1533 | 0.60171 | 27.05 | 12.8199
[6 empty] | 2.12 | 2.01 | 5 | 0 | 8.0629 | 0.40314 | 127.5768 | -0.30629 | -18.3772 | 0 | 0 | -0.15023 | 0.60021 | 27.1 | 12.783
[6 empty] | 2.13 | 2.02 | 5 | 0 | 8.0016 | 0.40008 | 124.514 | -0.30016 | -18.0096 | 0 | 0 | -0.14723 | 0.59873 | 27.15 | 12.7465
[6 empty] | 2.14 | 2.03 | 5 | 0 | 7.9416 | 0.39708 | 121.5124 | -0.29416 | -17.6494 | 0 | 0 | -0.14428 | 0.59729 | 27.2 | 12.7103
[6 empty] | 2.15 | 2.04 | 5 | 0 | 7.8827 | 0.39414 | 118.5708 | -0.28827 | -17.2965 | 0 | 0 | -0.1414 | 0.59588 | 27.25 | 12.6744
[6 empty] | 2.16 | 2.05 | 5 | 0 | 7.8251 | 0.39125 | 115.688 | -0.28251 | -16.9505 | 0 | 0 | -0.13857 | 0.59449 | 27.3 | 12.6389
[6 empty] | 2.17 | 2.06 | 5 | 0 | 7.7686 | 0.38843 | 112.8629 | -0.27686 | -16.6115 | 0 | 0 | -0.1358 | 0.59313 | 27.35 | 12.6037
[6 empty] | 2.18 | 2.07 | 5 | 0 | 7.7132 | 0.38566 | 110.0944 | -0.27132 | -16.2793 | 0 | 0 | -0.13308 | 0.5918 | 27.4 | 12.5688
[6 empty] | 2.19 | 2.08 | 5 | 0 | 7.6589 | 0.38295 | 107.3812 | -0.26589 | -15.9537 | 0 | 0 | -0.13042 | 0.5905 | 27.45 | 12.5342
[6 empty] | 2.2 | 2.09 | 5 | 0 | 7.6058 | 0.38029 | 104.7222 | -0.26058 | -15.6346 | 0 | 0 | -0.12781 | 0.58922 | 27.5 | 12.5
[6 empty] | 2.21 | 2.1 | 5 | 0 | 7.5537 | 0.37768 | 102.1164 | -0.25537 | -15.3219 | 0 | 0 | -0.12526 | 0.58797 | 27.55 | 12.4661
[6 empty] | 2.22 | 2.11 | 5 | 0 | 7.5026 | 0.37513 | 99.5628 | -0.25026 | -15.0155 | 0 | 0 | -0.12275 | 0.58674 | 27.6 | 12.4324
[6 empty] | 2.23 | 2.12 | 5 | 0 | 7.4525 | 0.37263 | 97.0602 | -0.24525 | -14.7152 | 0 | 0 | -0.1203 | 0.58554 | 27.65 | 12.3991
[6 empty] | 2.24 | 2.13 | 5 | 0 | 7.4035 | 0.37017 | 94.6077 | -0.24035 | -14.4209 | 0 | 0 | -0.11789 | 0.58436 | 27.7 | 12.3661
[6 empty] | 2.25 | 2.14 | 5 | 0 | 7.3554 | 0.36777 | 92.2042 | -0.23554 | -14.1325 | 0 | 0 | -0.11553 | 0.5832 | 27.75 | 12.3333
[6 empty] | 2.26 | 2.15 | 5 | 0 | 7.3083 | 0.36542 | 89.8488 | -0.23083 | -13.8498 | 0 | 0 | -0.11322 | 0.58207 | 27.8 | 12.3009
[6 empty] | 2.27 | 2.16 | 5 | 0 | 7.2621 | 0.36311 | 87.5405 | -0.22621 | -13.5728 | 0 | 0 | -0.11096 | 0.58096 | 27.85 | 12.2687
[6 empty] | 2.28 | 2.17 | 5 | 0 | 7.2169 | 0.36084 | 85.2783 | -0.22169 | -13.3014 | 0 | 0 | -0.10874 | 0.57987 | 27.9 | 12.2368
[6 empty] | 2.29 | 2.18 | 5 | 0 | 7.1726 | 0.35863 | 83.0614 | -0.21726 | -13.0353 | 0 | 0 | -0.10656 | 0.57881 | 27.95 | 12.2052
[6 empty] | 2.3 | 2.19 | 5 | 0 | 7.1291 | 0.35646 | 80.8889 | -0.21291 | -12.7746 | 0 | 0 | -0.10443 | 0.57776 | 28 | 12.1739
[6 empty] | 2.31 | 2.2 | 5 | 0 | 7.0865 | 0.35433 | 78.7598 | -0.20865 | -12.5191 | 0 | 0 | -0.10234 | 0.57674 | 28.05 | 12.1429
[6 empty] | 2.32 | 2.21 | 5 | 0 | 7.0448 | 0.35224 | 76.6733 | -0.20448 | -12.2687 | 0 | 0 | -0.1003 | 0.57574 | 28.1 | 12.1121
[6 empty] | 2.33 | 2.22 | 5 | 0 | 7.0039 | 0.35019 | 74.6285 | -0.20039 | -12.0234 | 0 | 0 | -0.098291 | 0.57475 | 28.15 | 12.0815
[6 empty] | 2.34 | 2.23 | 5 | 0 | 6.9638 | 0.34819 | 72.6246 | -0.19638 | -11.7829 | 0 | 0 | -0.096325 | 0.57379 | 28.2 | 12.0513
[6 empty] | 2.35 | 2.24 | 5 | 0 | 6.9245 | 0.34623 | 70.6608 | -0.19245 | -11.5472 | 0 | 0 | -0.094399 | 0.57285 | 28.25 | 12.0213
[6 empty] | 2.36 | 2.25 | 5 | 0 | 6.8861 | 0.3443 | 68.7362 | -0.18861 | -11.3163 | 0 | 0 | -0.092511 | 0.57192 | 28.3 | 11.9915
[6 empty] | 2.37 | 2.26 | 5 | 0 | 6.8483 | 0.34242 | 66.8502 | -0.18483 | -11.09 | 0 | 0 | -0.090661 | 0.57101 | 28.35 | 11.962
[6 empty] | 2.38 | 2.27 | 5 | 0 | 6.8114 | 0.34057 | 65.0018 | -0.18114 | -10.8682 | 0 | 0 | -0.088847 | 0.57013 | 28.4 | 11.9328
[6 empty] | 2.39 | 2.28 | 5 | 0 | 6.7751 | 0.33876 | 63.1905 | -0.17751 | -10.6508 | 0 | 0 | -0.08707 | 0.56926 | 28.45 | 11.9038
[6 empty] | 2.4 | 2.29 | 5 | 0 | 6.7396 | 0.33698 | 61.4153 | -0.17396 | -10.4378 | 0 | 0 | -0.085329 | 0.5684 | 28.5 | 11.875
[6 empty] | 2.41 | 2.3 | 5 | 0 | 6.7048 | 0.33524 | 59.6757 | -0.17048 | -10.229 | 0 | 0 | -0.083622 | 0.56757 | 28.55 | 11.8465
[6 empty] | 2.42 | 2.31 | 5 | 0 | 6.6707 | 0.33354 | 57.9709 | -0.16707 | -10.0245 | 0 | 0 | -0.08195 | 0.56675 | 28.6 | 11.8182
[6 empty] | 2.43 | 2.32 | 5 | 0 | 6.6373 | 0.33187 | 56.3001 | -0.16373 | -9.824 | 0 | 0 | -0.080311 | 0.56594 | 28.65 | 11.7901
[6 empty] | 2.44 | 2.33 | 5 | 0 | 6.6046 | 0.33023 | 54.6628 | -0.16046 | -9.6275 | 0 | 0 | -0.078705 | 0.56516 | 28.7 | 11.7623
[6 empty] | 2.45 | 2.34 | 5 | 0 | 6.5725 | 0.32862 | 53.0582 | -0.15725 | -9.4349 | 0 | 0 | -0.077131 | 0.56439 | 28.75 | 11.7347
[6 empty] | 2.46 | 2.35 | 5 | 0 | 6.541 | 0.32705 | 51.4857 | -0.1541 | -9.2462 | 0 | 0 | -0.075588 | 0.56363 | 28.8 | 11.7073
[6 empty] | 2.47 | 2.36 | 5 | 0 | 6.5102 | 0.32551 | 49.9447 | -0.15102 | -9.0613 | 0 | 0 | -0.074076 | 0.56289 | 28.85 | 11.6802
[6 empty] | 2.48 | 2.37 | 5 | 0 | 6.48 | 0.324 | 48.4345 | -0.148 | -8.8801 | 0 | 0 | -0.072595 | 0.56216 | 28.9 | 11.6532
[6 empty] | 2.49 | 2.38 | 5 | 0 | 6.4504 | 0.32252 | 46.9544 | -0.14504 | -8.7025 | 0 | 0 | -0.071143 | 0.56145 | 28.95 | 11.6265
[6 empty] | 2.5 | 2.39 | 5 | 0 | 6.4214 | 0.32107 | 45.504 | -0.14214 | -8.5284 | 0 | 0 | -0.06972 | 0.56075 | 29 | 11.6
[6 empty] | 2.51 | 2.4 | 5 | 0 | 6.393 | 0.31965 | 44.0826 | -0.1393 | -8.3579 | 0 | 0 | -0.068326 | 0.56007 | 29.05 | 11.5737
[6 empty] | 2.52 | 2.41 | 5 | 0 | 6.3651 | 0.31826 | 42.6896 | -0.13651 | -8.1907 | 0 | 0 | -0.066959 | 0.5594 | 29.1 | 11.5476
[6 empty] | 2.53 | 2.42 | 5 | 0 | 6.3378 | 0.31689 | 41.3245 | -0.13378 | -8.0269 | 0 | 0 | -0.06562 | 0.55874 | 29.15 | 11.5217
[6 empty] | 2.54 | 2.43 | 5 | 0 | 6.3111 | 0.31555 | 39.9867 | -0.13111 | -7.8664 | 0 | 0 | -0.064308 | 0.5581 | 29.2 | 11.4961
[6 empty] | 2.55 | 2.44 | 5 | 0 | 6.2848 | 0.31424 | 38.6756 | -0.12848 | -7.709 | 0 | 0 | -0.063021 | 0.55747 | 29.25 | 11.4706
[6 empty] | 2.56 | 2.45 | 5 | 0 | 6.2591 | 0.31296 | 37.3908 | -0.12591 | -7.5549 | 0 | 0 | -0.061761 | 0.55685 | 29.3 | 11.4453
[6 empty] | 2.57 | 2.46 | 5 | 0 | 6.234 | 0.3117 | 36.1317 | -0.1234 | -7.4038 | 0 | 0 | -0.060526 | 0.55625 | 29.35 | 11.4202
[6 empty] | 2.58 | 2.47 | 5 | 0 | 6.2093 | 0.31046 | 34.8977 | -0.12093 | -7.2557 | 0 | 0 | -0.059315 | 0.55566 | 29.4 | 11.3953
[6 empty] | 2.59 | 2.48 | 5 | 0 | 6.1851 | 0.30925 | 33.6884 | -0.11851 | -7.1106 | 0 | 0 | -0.058129 | 0.55507 | 29.45 | 11.3707
[6 empty] | 2.6 | 2.49 | 5 | 0 | 6.1614 | 0.30807 | 32.5033 | -0.11614 | -6.9684 | 0 | 0 | -0.056966 | 0.5545 | 29.5 | 11.3462
[6 empty] | 2.61 | 2.5 | 5 | 0 | 6.1382 | 0.30691 | 31.3419 | -0.11382 | -6.829 | 0 | 0 | -0.055827 | 0.55395 | 29.55 | 11.3218
[6 empty] | 2.62 | 2.51 | 5 | 0 | 6.1154 | 0.30577 | 30.2038 | -0.11154 | -6.6924 | 0 | 0 | -0.05471 | 0.5534 | 29.6 | 11.2977
[6 empty] | 2.63 | 2.52 | 5 | 0 | 6.0931 | 0.30465 | 29.0884 | -0.10931 | -6.5586 | 0 | 0 | -0.053616 | 0.55286 | 29.65 | 11.2738
[6 empty] | 2.64 | 2.53 | 5 | 0 | 6.0712 | 0.30356 | 27.9953 | -0.10712 | -6.4274 | 0 | 0 | -0.052544 | 0.55234 | 29.7 | 11.25
[6 empty] | 2.65 | 2.54 | 5 | 0 | 6.0498 | 0.30249 | 26.924 | -0.10498 | -6.2988 | 0 | 0 | -0.051493 | 0.55182 | 29.75 | 11.2264
[6 empty] | 2.66 | 2.55 | 5 | 0 | 6.0288 | 0.30144 | 25.8742 | -0.10288 | -6.1729 | 0 | 0 | -0.050463 | 0.55132 | 29.8 | 11.203
[6 empty] | 2.67 | 2.56 | 5 | 0 | 6.0082 | 0.30041 | 24.8454 | -0.10082 | -6.0494 | 0 | 0 | -0.049454 | 0.55082 | 29.85 | 11.1798
[6 empty] | 2.68 | 2.57 | 5 | 0 | 5.9881 | 0.2994 | 23.8372 | -0.098807 | -5.9284 | 0 | 0 | -0.048465 | 0.55034 | 29.9 | 11.1567
[6 empty] | 2.69 | 2.58 | 5 | 0 | 5.9683 | 0.29842 | 22.8491 | -0.096831 | -5.8099 | 0 | 0 | -0.047496 | 0.54986 | 29.95 | 11.1338
[6 empty] | 2.7 | 2.59 | 5 | 0 | 5.9489 | 0.29745 | 21.8808 | -0.094894 | -5.6937 | 0 | 0 | -0.046546 | 0.5494 | 30 | 11.1111
[6 empty] | 2.71 | 2.6 | 5 | 0 | 5.93 | 0.2965 | 20.9319 | -0.092996 | -5.5798 | 0 | 0 | -0.045615 | 0.54894 | 30.05 | 11.0886
[6 empty] | 2.72 | 2.61 | 5 | 0 | 5.9114 | 0.29557 | 20.0019 | -0.091136 | -5.4682 | 0 | 0 | -0.044702 | 0.5485 | 30.1 | 11.0662
[6 empty] | 2.73 | 2.62 | 5 | 0 | 5.8931 | 0.29466 | 19.0905 | -0.089314 | -5.3588 | 0 | 0 | -0.043808 | 0.54806 | 30.15 | 11.044
[6 empty] | 2.74 | 2.63 | 5 | 0 | 5.8753 | 0.29376 | 18.1974 | -0.087527 | -5.2516 | 0 | 0 | -0.042932 | 0.54763 | 30.2 | 11.0219
[6 empty] | 2.75 | 2.64 | 5 | 0 | 5.8578 | 0.29289 | 17.3221 | -0.085777 | -5.1466 | 0 | 0 | -0.042074 | 0.54721 | 30.25 | 11
[6 empty] | 2.76 | 2.65 | 5 | 0 | 5.8406 | 0.29203 | 16.4644 | -0.084061 | -5.0437 | 0 | 0 | -0.041232 | 0.54679 | 30.3 | 10.9783
[6 empty] | 2.77 | 2.66 | 5 | 0 | 5.8238 | 0.29119 | 15.6237 | -0.08238 | -4.9428 | 0 | 0 | -0.040407 | 0.54639 | 30.35 | 10.9567
[6 empty] | 2.78 | 2.67 | 5 | 0 | 5.8073 | 0.29037 | 14.7999 | -0.080733 | -4.844 | 0 | 0 | -0.039599 | 0.54599 | 30.4 | 10.9353
[6 empty] | 2.79 | 2.68 | 5 | 0 | 5.7912 | 0.28956 | 13.9926 | -0.079118 | -4.7471 | 0 | 0 | -0.038807 | 0.54561 | 30.45 | 10.914
[6 empty] | 2.8 | 2.69 | 5 | 0 | 5.7754 | 0.28877 | 13.2014 | -0.077536 | -4.6521 | 0 | 0 | -0.038031 | 0.54523 | 30.5 | 10.8929
[6 empty] | 2.81 | 2.7 | 5 | 0 | 5.7598 | 0.28799 | 12.4261 | -0.075985 | -4.5591 | 0 | 0 | -0.037271 | 0.54485 | 30.55 | 10.8719
[6 empty] | 2.82 | 2.71 | 5 | 0 | 5.7447 | 0.28723 | 11.6662 | -0.074465 | -4.4679 | 0 | 0 | -0.036525 | 0.54449 | 30.6 | 10.8511
[6 empty] | 2.83 | 2.72 | 5 | 0 | 5.7298 | 0.28649 | 10.9216 | -0.072976 | -4.3785 | 0 | 0 | -0.035795 | 0.54413 | 30.65 | 10.8304
[6 empty] | 2.84 | 2.73 | 5 | 0 | 5.7152 | 0.28576 | 10.1918 | -0.071516 | -4.291 | 0 | 0 | -0.035079 | 0.54378 | 30.7 | 10.8099
[6 empty] | 2.85 | 2.74 | 5 | 0 | 5.7009 | 0.28504 | 9.4767 | -0.070086 | -4.2052 | 0 | 0 | -0.034377 | 0.54344 | 30.75 | 10.7895
[6 empty] | 2.86 | 2.75 | 5 | 0 | 5.6868 | 0.28434 | 8.7758 | -0.068684 | -4.1211 | 0 | 0 | -0.03369 | 0.5431 | 30.8 | 10.7692
[6 empty] | 2.87 | 2.76 | 5 | 0 | 5.6731 | 0.28366 | 8.089 | -0.067311 | -4.0386 | 0 | 0 | -0.033016 | 0.54277 | 30.85 | 10.7491
[6 empty] | 2.88 | 2.77 | 5 | 0 | 5.6596 | 0.28298 | 7.4159 | -0.065964 | -3.9579 | 0 | 0 | -0.032356 | 0.54245 | 30.9 | 10.7292
[6 empty] | 2.89 | 2.78 | 5 | 0 | 5.6465 | 0.28232 | 6.7562 | -0.064645 | -3.8787 | 0 | 0 | -0.031708 | 0.54213 | 30.95 | 10.7093
[6 empty] | 2.9 | 2.79 | 5 | 0 | 5.6335 | 0.28168 | 6.1098 | -0.063352 | -3.8011 | 0 | 0 | -0.031074 | 0.54182 | 31 | 10.6897
[6 empty] | 2.91 | 2.8 | 5 | 0 | 5.6209 | 0.28104 | 5.4762 | -0.062085 | -3.7251 | 0 | 0 | -0.030453 | 0.54151 | 31.05 | 10.6701
[6 empty] | 2.92 | 2.81 | 5 | 0 | 5.6084 | 0.28042 | 4.8554 | -0.060843 | -3.6506 | 0 | 0 | -0.029844 | 0.54121 | 31.1 | 10.6507
[6 empty] | 2.93 | 2.82 | 5 | 0 | 5.5963 | 0.27981 | 4.247 | -0.059627 | -3.5776 | 0 | 0 | -0.029247 | 0.54092 | 31.15 | 10.6314
[6 empty] | 2.94 | 2.83 | 5 | 0 | 5.5843 | 0.27922 | 3.6507 | -0.058434 | -3.506 | 0 | 0 | -0.028662 | 0.54064 | 31.2 | 10.6122
[6 empty] | 2.95 | 2.84 | 5 | 0 | 5.5727 | 0.27863 | 3.0663 | -0.057265 | -3.4359 | 0 | 0 | -0.028089 | 0.54035 | 31.25 | 10.5932
[6 empty] | 2.96 | 2.85 | 5 | 0 | 5.5612 | 0.27806 | 2.4937 | -0.05612 | -3.3672 | 0 | 0 | -0.027527 | 0.54008 | 31.3 | 10.5743
[6 empty] | 2.97 | 2.86 | 5 | 0 | 5.55 | 0.2775 | 1.9325 | -0.054998 | -3.2999 | 0 | 0 | -0.026976 | 0.53981 | 31.35 | 10.5556
[6 empty] | 2.98 | 2.87 | 5 | 0 | 5.539 | 0.27695 | 1.3825 | -0.053898 | -3.2339 | 0 | 0 | -0.026437 | 0.53955 | 31.4 | 10.5369
[6 empty] | 2.99 | 2.88 | 5 | 0 | 5.5282 | 0.27641 | 0.84354 | -0.05282 | -3.1692 | 0 | 0 | -0.025908 | 0.53929 | 31.45 | 10.5184
[6 empty] | 3 | 2.89 | 16 | 0 | 5.5176 | 0.27588 | 0.31534 | 1.0482 | 62.8942 | 0 | 0 | 1.6453 | 0.55574 | 31.61 | 10.5367
[6 empty] | 3.01 | 2.9 | 16 | 0 | 5.7273 | 0.28636 | 10.7977 | 1.0273 | 61.6363 | 0 | 0 | 1.6124 | 0.57186 | 31.77 | 10.5548
[6 empty] | 3.02 | 2.91 | 16 | 0 | 5.9327 | 0.29664 | 21.0704 | 1.0067 | 60.4036 | 0 | 0 | 1.5802 | 0.58766 | 31.93 | 10.5728
[6 empty] | 3.03 | 2.92 | 16 | 0 | 6.1341 | 0.3067 | 31.1377 | 0.98659 | 59.1955 | 0 | 0 | 1.5486 | 0.60315 | 32.09 | 10.5908
[6 empty] | 3.04 | 2.93 | 16 | 0 | 6.3314 | 0.31657 | 41.0036 | 0.96686 | 58.0116 | 0 | 0 | 1.5176 | 0.61833 | 32.25 | 10.6086
[6 empty] | 3.05 | 2.94 | 16 | 0 | 6.5248 | 0.32624 | 50.6722 | 0.94752 | 56.8514 | 0 | 0 | 1.4872 | 0.6332 | 32.41 | 10.6262
[6 empty] | 3.06 | 2.95 | 16 | 0 | 6.7143 | 0.33571 | 60.1474 | 0.92857 | 55.7143 | 0 | 0 | 1.4575 | 0.64777 | 32.57 | 10.6438
[6 empty] | 3.07 | 2.96 | 16 | 0 | 6.9 | 0.345 | 69.4332 | 0.91 | 54.6001 | 0 | 0 | 1.4283 | 0.66206 | 32.73 | 10.6612
[6 empty] | 3.08 | 2.97 | 16 | 0 | 7.082 | 0.3541 | 78.5332 | 0.8918 | 53.5081 | 0 | 0 | 1.3998 | 0.67605 | 32.89 | 10.6786
[6 empty] | 3.09 | 2.98 | 16 | 0 | 7.2604 | 0.36302 | 87.4512 | 0.87396 | 52.4379 | 0 | 0 | 1.3718 | 0.68977 | 33.05 | 10.6958
[6 empty] | 3.1 | 2.99 | 16 | 0 | 7.4351 | 0.37176 | 96.1908 | 0.85649 | 51.3891 | 0 | 0 | 1.3443 | 0.70322 | 33.21 | 10.7129
[6 empty] | 3.11 | 3 | 16 | 0 | 7.6064 | 0.38032 | 104.7557 | 0.83936 | 50.3614 | 0 | 0 | 1.3175 | 0.71639 | 33.37 | 10.7299
[6 empty] | 3.12 | 3.01 | 16 | 0 | 7.7743 | 0.38872 | 113.1493 | 0.82257 | 49.3541 | 0 | 0 | 1.2911 | 0.7293 | 33.53 | 10.7468
[6 empty] | 3.13 | 3.02 | 16 | 0 | 7.9388 | 0.39694 | 121.3749 | 0.80612 | 48.367 | 0 | 0 | 1.2653 | 0.74195 | 33.69 | 10.7636
[6 empty] | 3.14 | 3.03 | 16 | 0 | 8.1 | 0.405 | 129.4361 | 0.79 | 47.3997 | 0 | 0 | 1.24 | 0.75435 | 33.85 | 10.7803
[6 empty] | 3.15 | 3.04 | 16 | 0 | 8.258 | 0.4129 | 137.3361 | 0.7742 | 46.4517 | 0 | 0 | 1.2152 | 0.76651 | 34.01 | 10.7968
[6 empty] | 3.16 | 3.05 | 16 | 0 | 8.4129 | 0.42064 | 145.078 | 0.75871 | 45.5227 | 0 | 0 | 1.1909 | 0.77841 | 34.17 | 10.8133
[6 empty] | 3.17 | 3.06 | 16 | 0 | 8.5646 | 0.42823 | 152.6651 | 0.74354 | 44.6122 | 0 | 0 | 1.1671 | 0.79008 | 34.33 | 10.8297
[6 empty] | 3.18 | 3.07 | 16 | 0 | 8.7133 | 0.43567 | 160.1005 | 0.72867 | 43.72 | 0 | 0 | 1.1437 | 0.80152 | 34.49 | 10.8459
[6 empty] | 3.19 | 3.08 | 16 | 0 | 8.8591 | 0.44295 | 167.3872 | 0.71409 | 42.8456 | 0 | 0 | 1.1208 | 0.81273 | 34.65 | 10.8621
[6 empty] | 3.2 | 3.09 | 16 | 0 | 9.0019 | 0.45009 | 174.5281 | 0.69981 | 41.9887 | 0 | 0 | 1.0984 | 0.82371 | 34.81 | 10.8781
[6 empty] | 3.21 | 3.1 | 16 | 0 | 9.1419 | 0.45709 | 181.5262 | 0.68581 | 41.1489 | 0 | 0 | 1.0765 | 0.83448 | 34.97 | 10.8941
[6 empty] | 3.22 | 3.11 | 16 | 0 | 9.279 | 0.46395 | 188.3844 | 0.6721 | 40.3259 | 0 | 0 | 1.0549 | 0.84503 | 35.13 | 10.9099
[6 empty] | 3.23 | 3.12 | 16 | 0 | 9.4134 | 0.47067 | 195.1053 | 0.65866 | 39.5194 | 0 | 0 | 1.0338 | 0.85537 | 35.29 | 10.9257
[6 empty] | 3.24 | 3.13 | 16 | 0 | 9.5452 | 0.47726 | 201.6919 | 0.64548 | 38.729 | 0 | 0 | 1.0132 | 0.8655 | 35.45 | 10.9414
[6 empty] | 3.25 | 3.14 | 16 | 0 | 9.6743 | 0.48371 | 208.1467 | 0.63257 | 37.9544 | 0 | 0 | 0.99289 | 0.87543 | 35.61 | 10.9569
[6 empty] | 3.26 | 3.15 | 16 | 0 | 9.8008 | 0.49004 | 214.4725 | 0.61992 | 37.1953 | 0 | 0 | 0.97303 | 0.88516 | 35.77 | 10.9724
[6 empty] | 3.27 | 3.16 | 16 | 0 | 9.9248 | 0.49624 | 220.6717 | 0.60752 | 36.4514 | 0 | 0 | 0.95357 | 0.89469 | 35.93 | 10.9878
[6 empty] | 3.28 | 3.17 | 16 | 0 | 10.0463 | 0.50231 | 226.747 | 0.59537 | 35.7224 | 0 | 0 | 0.9345 | 0.90404 | 36.09 | 11.003
[6 empty] | 3.29 | 3.18 | 16 | 0 | 10.1653 | 0.50827 | 232.7007 | 0.58347 | 35.008 | 0 | 0 | 0.91581 | 0.9132 | 36.25 | 11.0182
[6 empty] | 3.3 | 3.19 | 16 | 0 | 10.282 | 0.5141 | 238.5353 | 0.5718 | 34.3078 | 0 | 0 | 0.89749 | 0.92217 | 36.41 | 11.0333
[6 empty] | 3.31 | 3.2 | 16 | 0 | 10.3964 | 0.51982 | 244.2533 | 0.56036 | 33.6216 | 0 | 0 | 0.87954 | 0.93097 | 36.57 | 11.0483
[6 empty] | 3.32 | 3.21 | 16 | 0 | 10.5085 | 0.52542 | 249.8569 | 0.54915 | 32.9492 | 0 | 0 | 0.86195 | 0.93959 | 36.73 | 11.0633
[6 empty] | 3.33 | 3.22 | 16 | 0 | 10.6183 | 0.53091 | 255.3485 | 0.53817 | 32.2902 | 0 | 0 | 0.84471 | 0.94803 | 36.89 | 11.0781
[6 empty] | 3.34 | 3.23 | 16 | 0 | 10.7259 | 0.5363 | 260.7302 | 0.52741 | 31.6444 | 0 | 0 | 0.82782 | 0.95631 | 37.05 | 11.0928
[6 empty] | 3.35 | 3.24 | 16 | 0 | 10.8314 | 0.54157 | 266.0042 | 0.51686 | 31.0115 | 0 | 0 | 0.81126 | 0.96442 | 37.21 | 11.1075
[6 empty] | 3.36 | 3.25 | 16 | 0 | 10.9348 | 0.54674 | 271.1728 | 0.50652 | 30.3913 | 0 | 0 | 0.79504 | 0.97237 | 37.37 | 11.122
[6 empty] | 3.37 | 3.26 | 16 | 0 | 11.0361 | 0.5518 | 276.238 | 0.49639 | 29.7835 | 0 | 0 | 0.77914 | 0.98017 | 37.53 | 11.1365
[6 empty] | 3.38 | 3.27 | 16 | 0 | 11.1354 | 0.55677 | 281.2019 | 0.48646 | 29.1878 | 0 | 0 | 0.76355 | 0.9878 | 37.69 | 11.1509
[6 empty] | 3.39 | 3.28 | 16 | 0 | 11.2327 | 0.56163 | 286.0666 | 0.47673 | 28.6041 | 0 | 0 | 0.74828 | 0.99528 | 37.85 | 11.1652
[6 empty] | 3.4 | 3.29 | 16 | 0 | 11.328 | 0.5664 | 290.8339 | 0.4672 | 28.032 | 0 | 0 | 0.73332 | 1.0026 | 38.01 | 11.1794
[6 empty] | 3.41 | 3.3 | 16 | 0 | 11.4214 | 0.57107 | 295.5059 | 0.45786 | 27.4713 | 0 | 0 | 0.71865 | 1.0098 | 38.17 | 11.1935
[6 empty] | 3.42 | 3.31 | 16 | 0 | 11.513 | 0.57565 | 300.0845 | 0.4487 | 26.9219 | 0 | 0 | 0.70428 | 1.0168 | 38.33 | 11.2076
[6 empty] | 3.43 | 3.32 | 16 | 0 | 11.6028 | 0.58014 | 304.5715 | 0.43972 | 26.3835 | 0 | 0 | 0.69019 | 1.0237 | 38.49 | 11.2216
[6 empty] | 3.44 | 3.33 | 16 | 0 | 11.6907 | 0.58454 | 308.9687 | 0.43093 | 25.8558 | 0 | 0 | 0.67639 | 1.0305 | 38.65 | 11.2355
[6 empty] | 3.45 | 3.34 | 16 | 0 | 11.7769 | 0.58884 | 313.278 | 0.42231 | 25.3387 | 0 | 0 | 0.66286 | 1.0371 | 38.81 | 11.2493
[6 empty] | 3.46 | 3.35 | 16 | 0 | 11.8613 | 0.59307 | 317.5011 | 0.41387 | 24.8319 | 0 | 0 | 0.6496 | 1.0436 | 38.97 | 11.263
[6 empty] | 3.47 | 3.36 | 16 | 0 | 11.9441 | 0.59721 | 321.6398 | 0.40559 | 24.3353 | 0 | 0 | 0.63661 | 1.05 | 39.13 | 11.2767
[6 empty] | 3.48 | 3.37 | 16 | 0 | 12.0252 | 0.60126 | 325.6956 | 0.39748 | 23.8486 | 0 | 0 | 0.62388 | 1.0562 | 39.29 | 11.2902
[6 empty] | 3.49 | 3.38 | 16 | 0 | 12.1047 | 0.60524 | 329.6704 | 0.38953 | 23.3716 | 0 | 0 | 0.6114 | 1.0624 | 39.45 | 11.3037
[6 empty] | 3.5 | 3.39 | 16 | 0 | 12.1826 | 0.60913 | 333.5657 | 0.38174 | 22.9042 | 0 | 0 | 0.59917 | 1.0683 | 39.61 | 11.3171
[6 empty] | 3.51 | 3.4 | 16 | 0 | 12.259 | 0.61295 | 337.383 | 0.3741 | 22.4461 | 0 | 0 | 0.58719 | 1.0742 | 39.77 | 11.3305
[6 empty] | 3.52 | 3.41 | 16 | 0 | 12.3338 | 0.61669 | 341.124 | 0.36662 | 21.9972 | 0 | 0 | 0.57545 | 1.08 | 39.93 | 11.3438
[6 empty] | 3.53 | 3.42 | 16 | 0 | 12.4071 | 0.62036 | 344.7902 | 0.35929 | 21.5572 | 0 | 0 | 0.56394 | 1.0856 | 40.09 | 11.3569
[6 empty] | 3.54 | 3.43 | 16 | 0 | 12.479 | 0.62395 | 348.3831 | 0.3521 | 21.1261 | 0 | 0 | 0.55266 | 1.0911 | 40.25 | 11.3701
[6 empty] | 3.55 | 3.44 | 16 | 0 | 12.5494 | 0.62747 | 351.9041 | 0.34506 | 20.7035 | 0 | 0 | 0.5416 | 1.0966 | 40.41 | 11.3831
[6 empty] | 3.56 | 3.45 | 16 | 0 | 12.6184 | 0.63092 | 355.3547 | 0.33816 | 20.2895 | 0 | 0 | 0.53077 | 1.1019 | 40.57 | 11.3961
[6 empty] | 3.57 | 3.46 | 16 | 0 | 12.6861 | 0.6343 | 358.7363 | 0.33139 | 19.8837 | 0 | 0 | 0.52016 | 1.1071 | 40.73 | 11.409
[6 empty] | 3.58 | 3.47 | 16 | 0 | 12.7523 | 0.63762 | 362.0502 | 0.32477 | 19.486 | 0 | 0 | 0.50975 | 1.1122 | 40.89 | 11.4218
[6 empty] | 3.59 | 3.48 | 16 | 0 | 12.8173 | 0.64086 | 365.2979 | 0.31827 | 19.0963 | 0 | 0 | 0.49956 | 1.1172 | 41.05 | 11.4345
[6 empty] | 3.6 | 3.49 | 16 | 0 | 12.8809 | 0.64405 | 368.4806 | 0.31191 | 18.7144 | 0 | 0 | 0.48957 | 1.1221 | 41.21 | 11.4472
[6 empty] | 3.61 | 3.5 | 16 | 0 | 12.9433 | 0.64717 | 371.5997 | 0.30567 | 18.3401 | 0 | 0 | 0.47978 | 1.1269 | 41.37 | 11.4598
[6 empty] | 3.62 | 3.51 | 16 | 0 | 13.0045 | 0.65022 | 374.6564 | 0.29955 | 17.9733 | 0 | 0 | 0.47018 | 1.1316 | 41.53 | 11.4724
[6 empty] | 3.63 | 3.52 | 16 | 0 | 13.0644 | 0.65322 | 377.6519 | 0.29356 | 17.6138 | 0 | 0 | 0.46078 | 1.1362 | 41.69 | 11.4848
[6 empty] | 3.64 | 3.53 | 16 | 0 | 13.1231 | 0.65615 | 380.5875 | 0.28769 | 17.2615 | 0 | 0 | 0.45156 | 1.1407 | 41.85 | 11.4973
[6 empty] | 3.65 | 3.54 | 16 | 0 | 13.1806 | 0.65903 | 383.4645 | 0.28194 | 16.9163 | 0 | 0 | 0.44253 | 1.1451 | 42.01 | 11.5096
[6 empty] | 3.66 | 3.55 | 16 | 0 | 13.237 | 0.66185 | 386.2838 | 0.2763 | 16.578 | 0 | 0 | 0.43368 | 1.1494 | 42.17 | 11.5219
[6 empty] | 3.67 | 3.56 | 16 | 0 | 13.2923 | 0.66461 | 389.0468 | 0.27077 | 16.2464 | 0 | 0 | 0.42501 | 1.1537 | 42.33 | 11.5341
[6 empty] | 3.68 | 3.57 | 16 | 0 | 13.3464 | 0.66732 | 391.7546 | 0.26536 | 15.9215 | 0 | 0 | 0.41651 | 1.1579 | 42.49 | 11.5462
[6 empty] | 3.69 | 3.58 | 16 | 0 | 13.3995 | 0.66997 | 394.4082 | 0.26005 | 15.6031 | 0 | 0 | 0.40818 | 1.1619 | 42.65 | 11.5583
[6 empty] | 3.7 | 3.59 | 16 | 0 | 13.4515 | 0.67257 | 397.0087 | 0.25485 | 15.291 | 0 | 0 | 0.40001 | 1.1659 | 42.81 | 11.5703
[6 empty] | 3.71 | 3.6 | 16 | 0 | 13.5025 | 0.67512 | 399.5572 | 0.24975 | 14.9852 | 0 | 0 | 0.39201 | 1.1699 | 42.97 | 11.5822
[6 empty] | 3.72 | 3.61 | 16 | 0 | 13.5524 | 0.67762 | 402.0547 | 0.24476 | 14.6855 | 0 | 0 | 0.38417 | 1.1737 | 43.13 | 11.5941
[6 empty] | 3.73 | 3.62 | 16 | 0 | 13.6014 | 0.68007 | 404.5023 | 0.23986 | 14.3918 | 0 | 0 | 0.37649 | 1.1775 | 43.29 | 11.6059
[6 empty] | 3.74 | 3.63 | 16 | 0 | 13.6493 | 0.68247 | 406.9009 | 0.23507 | 14.1039 | 0 | 0 | 0.36896 | 1.1812 | 43.45 | 11.6176
[6 empty] | 3.75 | 3.64 | 16 | 0 | 13.6964 | 0.68482 | 409.2516 | 0.23036 | 13.8219 | 0 | 0 | 0.36158 | 1.1848 | 43.61 | 11.6293
[6 empty] | 3.76 | 3.65 | 16 | 0 | 13.7424 | 0.68712 | 411.5552 | 0.22576 | 13.5454 | 0 | 0 | 0.35435 | 1.1883 | 43.77 | 11.641
[6 empty] | 3.77 | 3.66 | 16 | 0 | 13.7876 | 0.68938 | 413.8128 | 0.22124 | 13.2745 | 0 | 0 | 0.34726 | 1.1918 | 43.93 | 11.6525
[6 empty] | 3.78 | 3.67 | 16 | 0 | 13.8318 | 0.69159 | 416.0252 | 0.21682 | 13.009 | 0 | 0 | 0.34032 | 1.1952 | 44.09 | 11.664
[6 empty] | 3.79 | 3.68 | 16 | 0 | 13.8752 | 0.69376 | 418.1934 | 0.21248 | 12.7488 | 0 | 0 | 0.33351 | 1.1985 | 44.25 | 11.6755
[6 empty] | 3.8 | 3.69 | 16 | 0 | 13.9177 | 0.69588 | 420.3182 | 0.20823 | 12.4939 | 0 | 0 | 0.32684 | 1.2018 | 44.41 | 11.6868
[6 empty] | 3.81 | 3.7 | 16 | 0 | 13.9593 | 0.69797 | 422.4005 | 0.20407 | 12.244 | 0 | 0 | 0.3203 | 1.205 | 44.57 | 11.6982
[6 empty] | 3.82 | 3.71 | 16 | 0 | 14.0001 | 0.70001 | 424.4411 | 0.19999 | 11.9991 | 0 | 0 | 0.3139 | 1.2081 | 44.73 | 11.7094
[6 empty] | 3.83 | 3.72 | 16 | 0 | 14.0401 | 0.70201 | 426.441 | 0.19599 | 11.7591 | 0 | 0 | 0.30762 | 1.2112 | 44.89 | 11.7206
[6 empty] | 3.84 | 3.73 | 16 | 0 | 14.0793 | 0.70397 | 428.4009 | 0.19207 | 11.5239 | 0 | 0 | 0.30147 | 1.2142 | 45.05 | 11.7318
[6 empty] | 3.85 | 3.74 | 16 | 0 | 14.1178 | 0.70589 | 430.3215 | 0.18822 | 11.2935 | 0 | 0 | 0.29544 | 1.2172 | 45.21 | 11.7429
[6 empty] | 3.86 | 3.75 | 16 | 0 | 14.1554 | 0.70777 | 432.2038 | 0.18446 | 11.0676 | 0 | 0 | 0.28953 | 1.2201 | 45.37 | 11.7539
[6 empty] | 3.87 | 3.76 | 16 | 0 | 14.1923 | 0.70961 | 434.0483 | 0.18077 | 10.8462 | 0 | 0 | 0.28374 | 1.2229 | 45.53 | 11.7649
[6 empty] | 3.88 | 3.77 | 16 | 0 | 14.2284 | 0.71142 | 435.8561 | 0.17716 | 10.6293 | 0 | 0 | 0.27806 | 1.2257 | 45.69 | 11.7758
[6 empty] | 3.89 | 3.78 | 16 | 0 | 14.2639 | 0.71319 | 437.6276 | 0.17361 | 10.4167 | 0 | 0 | 0.2725 | 1.2284 | 45.85 | 11.7866
[6 empty] | 3.9 | 3.79 | 16 | 0 | 14.2986 | 0.71493 | 439.3637 | 0.17014 | 10.2084 | 0 | 0 | 0.26705 | 1.2311 | 46.01 | 11.7974
[6 empty] | 3.91 | 3.8 | 16 | 0 | 14.3326 | 0.71663 | 441.0651 | 0.16674 | 10.0042 | 0 | 0 | 0.26171 | 1.2337 | 46.17 | 11.8082
[6 empty] | 3.92 | 3.81 | 16 | 0 | 14.366 | 0.7183 | 442.7325 | 0.1634 | 9.8041 | 0 | 0 | 0.25648 | 1.2363 | 46.33 | 11.8189
[6 empty] | 3.93 | 3.82 | 16 | 0 | 14.3987 | 0.71993 | 444.3665 | 0.16013 | 9.6081 | 0 | 0 | 0.25135 | 1.2388 | 46.49 | 11.8295
[6 empty] | 3.94 | 3.83 | 16 | 0 | 14.4307 | 0.72153 | 445.9679 | 0.15693 | 9.4159 | 0 | 0 | 0.24632 | 1.2412 | 46.65 | 11.8401
[6 empty] | 3.95 | 3.84 | 16 | 0 | 14.4621 | 0.7231 | 447.5372 | 0.15379 | 9.2276 | 0 | 0 | 0.24139 | 1.2437 | 46.81 | 11.8506
[6 empty] | 3.96 | 3.85 | 16 | 0 | 14.4928 | 0.72464 | 449.0751 | 0.15072 | 9.043 | 0 | 0 | 0.23657 | 1.246 | 46.97 | 11.8611
[6 empty] | 3.97 | 3.86 | 16 | 0 | 14.523 | 0.72615 | 450.5823 | 0.1477 | 8.8622 | 0 | 0 | 0.23183 | 1.2483 | 47.13 | 11.8715
[6 empty] | 3.98 | 3.87 | 16 | 0 | 14.5525 | 0.72763 | 452.0593 | 0.14475 | 8.6849 | 0 | 0 | 0.2272 | 1.2506 | 47.29 | 11.8819
[6 empty] | 3.99 | 3.88 | 16 | 0 | 14.5815 | 0.72907 | 453.5068 | 0.14185 | 8.5112 | 0 | 0 | 0.22265 | 1.2528 | 47.45 | 11.8922
[6 empty] | 4 | 3.89 | 16 | 0 | 14.6098 | 0.73049 | 454.9253 | 0.13902 | 8.341 | 0 | 0 | 0.2182 | 1.255 | 47.61 | 11.9025
[6 empty] | 4.01 | 3.9 | 16 | 0 | 14.6376 | 0.73188 | 456.3155 | 0.13624 | 8.1742 | 0 | 0 | 0.21384 | 1.2572 | 47.77 | 11.9127
[6 empty] | 4.02 | 3.91 | 16 | 0 | 14.6649 | 0.73324 | 457.6779 | 0.13351 | 8.0107 | 0 | 0 | 0.20956 | 1.2593 | 47.93 | 11.9229
[6 empty] | 4.03 | 3.92 | 16 | 0 | 14.6916 | 0.73458 | 459.013 | 0.13084 | 7.8505 | 0 | 0 | 0.20537 | 1.2613 | 48.09 | 11.933
[6 empty] | 4.04 | 3.93 | 16 | 0 | 14.7178 | 0.73589 | 460.3214 | 0.12822 | 7.6935 | 0 | 0 | 0.20126 | 1.2633 | 48.25 | 11.9431
[6 empty] | 4.05 | 3.94 | 16 | 0 | 14.7434 | 0.73717 | 461.6036 | 0.12566 | 7.5396 | 0 | 0 | 0.19724 | 1.2653 | 48.41 | 11.9531
[6 empty] | 4.06 | 3.95 | 16 | 0 | 14.7685 | 0.73843 | 462.8602 | 0.12315 | 7.3888 | 0 | 0 | 0.19329 | 1.2672 | 48.57 | 11.9631
[6 empty] | 4.07 | 3.96 | 16 | 0 | 14.7932 | 0.73966 | 464.0917 | 0.12068 | 7.241 | 0 | 0 | 0.18943 | 1.2691 | 48.73 | 11.973
[6 empty] | 4.08 | 3.97 | 16 | 0 | 14.8173 | 0.74086 | 465.2986 | 0.11827 | 7.0962 | 0 | 0 | 0.18564 | 1.271 | 48.89 | 11.9828
[6 empty] | 4.09 | 3.98 | 16 | 0 | 14.841 | 0.74205 | 466.4813 | 0.1159 | 6.9543 | 0 | 0 | 0.18192 | 1.2728 | 49.05 | 11.9927
[6 empty] | 4.1 | 3.99 | 16 | 0 | 14.8641 | 0.74321 | 467.6403 | 0.11359 | 6.8152 | 0 | 0 | 0.17829 | 1.2746 | 49.21 | 12.0024
[6 empty] | 4.11 | 4 | 16 | 0 | 14.8868 | 0.74434 | 468.7762 | 0.11132 | 6.6789 | 0 | 0 | 0.17472 | 1.2763 | 49.37 | 12.0122
[6 empty] | 4.12 | 4.01 | 16 | 0 | 14.9091 | 0.74546 | 469.8893 | 0.10909 | 6.5453 | 0 | 0 | 0.17123 | 1.278 | 49.53 | 12.0218
[6 empty] | 4.13 | 4.02 | 16 | 0 | 14.9309 | 0.74655 | 470.9802 | 0.10691 | 6.4144 | 0 | 0 | 0.1678 | 1.2797 | 49.69 | 12.0315
[6 empty] | 4.14 | 4.03 | 16 | 0 | 14.9523 | 0.74762 | 472.0493 | 0.10477 | 6.2861 | 0 | 0 | 0.16445 | 1.2814 | 49.85 | 12.0411
[6 empty] | 4.15 | 4.04 | 16 | 0 | 14.9733 | 0.74866 | 473.097 | 0.10267 | 6.1604 | 0 | 0 | 0.16116 | 1.283 | 50.01 | 12.0506
[6 empty] | 4.16 | 4.05 | 16 | 0 | 14.9938 | 0.74969 | 474.1237 | 0.10062 | 6.0372 | 0 | 0 | 0.15793 | 1.2846 | 50.17 | 12.0601
[6 empty] | 4.17 | 4.06 | 16 | 0 | 15.0139 | 0.7507 | 475.1299 | 0.098608 | 5.9165 | 0 | 0 | 0.15477 | 1.2861 | 50.33 | 12.0695
[6 empty] | 4.18 | 4.07 | 16 | 0 | 15.0336 | 0.75168 | 476.116 | 0.096635 | 5.7981 | 0 | 0 | 0.15168 | 1.2876 | 50.49 | 12.0789
[6 empty] | 4.19 | 4.08 | 16 | 0 | 15.053 | 0.75265 | 477.0823 | 0.094703 | 5.6822 | 0 | 0 | 0.14865 | 1.2891 | 50.65 | 12.0883
[6 empty] | 4.2 | 4.09 | 16 | 0 | 15.0719 | 0.7536 | 478.0294 | 0.092809 | 5.5685 | 0 | 0 | 0.14567 | 1.2906 | 50.81 | 12.0976
[6 empty] | 4.21 | 4.1 | 16 | 0 | 15.0905 | 0.75452 | 478.9574 | 0.090952 | 5.4571 | 0 | 0 | 0.14276 | 1.292 | 50.97 | 12.1069
[6 empty] | 4.22 | 4.11 | 16 | 0 | 15.1087 | 0.75543 | 479.867 | 0.089133 | 5.348 | 0 | 0 | 0.1399 | 1.2934 | 51.13 | 12.1161
[6 empty] | 4.23 | 4.12 | 16 | 0 | 15.1265 | 0.75632 | 480.7583 | 0.087351 | 5.241 | 0 | 0 | 0.13711 | 1.2948 | 51.29 | 12.1253
[6 empty] | 4.24 | 4.13 | 16 | 0 | 15.144 | 0.7572 | 481.6318 | 0.085604 | 5.1362 | 0 | 0 | 0.13436 | 1.2961 | 51.45 | 12.1344
[6 empty] | 4.25 | 4.14 | 16 | 0 | 15.1611 | 0.75805 | 482.4878 | 0.083892 | 5.0335 | 0 | 0 | 0.13168 | 1.2974 | 51.61 | 12.1435
[6 empty] | 4.26 | 4.15 | 16 | 0 | 15.1779 | 0.75889 | 483.3268 | 0.082214 | 4.9328 | 0 | 0 | 0.12904 | 1.2987 | 51.77 | 12.1526
[6 empty] | 4.27 | 4.16 | 16 | 0 | 15.1943 | 0.75972 | 484.1489 | 0.08057 | 4.8342 | 0 | 0 | 0.12646 | 1.3 | 51.93 | 12.1616
[6 empty] | 4.28 | 4.17 | 16 | 0 | 15.2104 | 0.76052 | 484.9546 | 0.078958 | 4.7375 | 0 | 0 | 0.12393 | 1.3012 | 52.09 | 12.1706
[6 empty] | 4.29 | 4.18 | 16 | 0 | 15.2262 | 0.76131 | 485.7442 | 0.077379 | 4.6427 | 0 | 0 | 0.12145 | 1.3024 | 52.25 | 12.1795
[6 empty] | 4.3 | 4.19 | 16 | 0 | 15.2417 | 0.76208 | 486.518 | 0.075831 | 4.5499 | 0 | 0 | 0.11902 | 1.3036 | 52.41 | 12.1884
[6 empty] | 4.31 | 4.2 | 16 | 0 | 15.2569 | 0.76284 | 487.2763 | 0.074315 | 4.4589 | 0 | 0 | 0.11664 | 1.3048 | 52.57 | 12.1972
[6 empty] | 4.32 | 4.21 | 16 | 0 | 15.2717 | 0.76359 | 488.0194 | 0.072828 | 4.3697 | 0 | 0 | 0.11431 | 1.3059 | 52.73 | 12.206
[6 empty] | 4.33 | 4.22 | 16 | 0 | 15.2863 | 0.76431 | 488.7477 | 0.071372 | 4.2823 | 0 | 0 | 0.11203 | 1.307 | 52.89 | 12.2148
[6 empty] | 4.34 | 4.23 | 16 | 0 | 15.3006 | 0.76503 | 489.4614 | 0.069944 | 4.1967 | 0 | 0 | 0.10978 | 1.3081 | 53.05 | 12.2235
[6 empty] | 4.35 | 4.24 | 16 | 0 | 15.3145 | 0.76573 | 490.1609 | 0.068546 | 4.1127 | 0 | 0 | 0.10759 | 1.3092 | 53.21 | 12.2322
[6 empty] | 4.36 | 4.25 | 16 | 0 | 15.3283 | 0.76641 | 490.8463 | 0.067175 | 4.0305 | 0 | 0 | 0.10544 | 1.3103 | 53.37 | 12.2408
[6 empty] | 4.37 | 4.26 | 16 | 0 | 15.3417 | 0.76708 | 491.5181 | 0.065831 | 3.9499 | 0 | 0 | 0.10333 | 1.3113 | 53.53 | 12.2494
[6 empty] | 4.38 | 4.27 | 16 | 0 | 15.3549 | 0.76774 | 492.1764 | 0.064515 | 3.8709 | 0 | 0 | 0.10126 | 1.3123 | 53.69 | 12.258
[6 empty] | 4.39 | 4.28 | 16 | 0 | 15.3678 | 0.76839 | 492.8215 | 0.063224 | 3.7935 | 0 | 0 | 0.099237 | 1.3133 | 53.85 | 12.2665
[6 empty] | 4.4 | 4.29 | 16 | 0 | 15.3804 | 0.76902 | 493.4538 | 0.06196 | 3.7176 | 0 | 0 | 0.097252 | 1.3143 | 54.01 | 12.275
[6 empty] | 4.41 | 4.3 | 16 | 0 | 15.3928 | 0.76964 | 494.0734 | 0.060721 | 3.6432 | 0 | 0 | 0.095307 | 1.3152 | 54.17 | 12.2834
[6 empty] | 4.42 | 4.31 | 16 | 0 | 15.4049 | 0.77025 | 494.6806 | 0.059506 | 3.5704 | 0 | 0 | 0.093401 | 1.3162 | 54.33 | 12.2919
[6 empty] | 4.43 | 4.32 | 16 | 0 | 15.4168 | 0.77084 | 495.2756 | 0.058316 | 3.499 | 0 | 0 | 0.091533 | 1.3171 | 54.49 | 12.3002
[6 empty] | 4.44 | 4.33 | 16 | 0 | 15.4285 | 0.77143 | 495.8588 | 0.05715 | 3.429 | 0 | 0 | 0.089702 | 1.318 | 54.65 | 12.3086
[6 empty] | 4.45 | 4.34 | 16 | 0 | 15.4399 | 0.772 | 496.4303 | 0.056007 | 3.3604 | 0 | 0 | 0.087908 | 1.3189 | 54.81 | 12.3169
[6 empty] | 4.46 | 4.35 | 16 | 0 | 15.4511 | 0.77256 | 496.9904 | 0.054887 | 3.2932 | 0 | 0 | 0.08615 | 1.3197 | 54.97 | 12.3251
[6 empty] | 4.47 | 4.36 | 16 | 0 | 15.4621 | 0.77311 | 497.5392 | 0.053789 | 3.2273 | 0 | 0 | 0.084427 | 1.3206 | 55.13 | 12.3333
[6 empty] | 4.48 | 4.37 | 16 | 0 | 15.4729 | 0.77364 | 498.0771 | 0.052713 | 3.1628 | 0 | 0 | 0.082738 | 1.3214 | 55.29 | 12.3415
[6 empty] | 4.49 | 4.38 | 16 | 0 | 15.4834 | 0.77417 | 498.6043 | 0.051659 | 3.0995 | 0 | 0 | 0.081084 | 1.3222 | 55.45 | 12.3497
[6 empty] | 4.5 | 4.39 | 16 | 0 | 15.4937 | 0.77469 | 499.1208 | 0.050626 | 3.0375 | 0 | 0 | 0.079462 | 1.323 | 55.61 | 12.3578
[6 empty] | 4.51 | 4.4 | 5 | 0 | 15.5039 | 0.77519 | 499.6271 | -1.0504 | -63.0232 | 0 | 0 | -0.51521 | 1.3179 | 55.66 | 12.3415
[6 empty] | 4.52 | 4.41 | 5 | 0 | 15.2938 | 0.76469 | 489.1232 | -1.0294 | -61.7627 | 0 | 0 | -0.50491 | 1.3128 | 55.71 | 12.3252
[6 empty] | 4.53 | 4.42 | 5 | 0 | 15.0879 | 0.7544 | 478.8294 | -1.0088 | -60.5275 | 0 | 0 | -0.49481 | 1.3079 | 55.76 | 12.3091
[6 empty] | 4.54 | 4.43 | 5 | 0 | 14.8862 | 0.74431 | 468.7415 | -0.98862 | -59.3169 | 0 | 0 | -0.48492 | 1.303 | 55.81 | 12.293
[6 empty] | 4.55 | 4.44 | 5 | 0 | 14.6884 | 0.73442 | 458.8554 | -0.96884 | -58.1306 | 0 | 0 | -0.47522 | 1.2983 | 55.86 | 12.2769
[6 empty] | 4.56 | 4.45 | 5 | 0 | 14.4947 | 0.72473 | 449.1669 | -0.94947 | -56.968 | 0 | 0 | -0.46571 | 1.2936 | 55.91 | 12.261
[6 empty] | 4.57 | 4.46 | 5 | 0 | 14.3048 | 0.71524 | 439.6723 | -0.93048 | -55.8286 | 0 | 0 | -0.4564 | 1.289 | 55.96 | 12.2451
[6 empty] | 4.58 | 4.47 | 5 | 0 | 14.1187 | 0.70593 | 430.3675 | -0.91187 | -54.7121 | 0 | 0 | -0.44727 | 1.2846 | 56.01 | 12.2293
[6 empty] | 4.59 | 4.48 | 5 | 0 | 13.9363 | 0.69682 | 421.2488 | -0.89363 | -53.6178 | 0 | 0 | -0.43833 | 1.2802 | 56.06 | 12.2135
[6 empty] | 4.6 | 4.49 | 5 | 0 | 13.7576 | 0.68788 | 412.3125 | -0.87576 | -52.5455 | 0 | 0 | -0.42956 | 1.2759 | 56.11 | 12.1978
[6 empty] | 4.61 | 4.5 | 5 | 0 | 13.5824 | 0.67912 | 403.5549 | -0.85824 | -51.4946 | 0 | 0 | -0.42097 | 1.2717 | 56.16 | 12.1822
[6 empty] | 4.62 | 4.51 | 5 | 0 | 13.4108 | 0.67054 | 394.9725 | -0.84108 | -50.4647 | 0 | 0 | -0.41255 | 1.2675 | 56.21 | 12.1667
[6 empty] | 4.63 | 4.52 | 5 | 0 | 13.2426 | 0.66213 | 386.5617 | -0.82426 | -49.4554 | 0 | 0 | -0.4043 | 1.2635 | 56.26 | 12.1512
[6 empty] | 4.64 | 4.53 | 5 | 0 | 13.0777 | 0.65389 | 378.3192 | -0.80777 | -48.4663 | 0 | 0 | -0.39621 | 1.2595 | 56.31 | 12.1358
[6 empty] | 4.65 | 4.54 | 5 | 0 | 12.9162 | 0.64581 | 370.2415 | -0.79162 | -47.4969 | 0 | 0 | -0.38829 | 1.2557 | 56.36 | 12.1204
[6 empty] | 4.66 | 4.55 | 5 | 0 | 12.7578 | 0.63789 | 362.3253 | -0.77578 | -46.547 | 0 | 0 | -0.38052 | 1.2519 | 56.41 | 12.1052
[6 empty] | 4.67 | 4.56 | 5 | 0 | 12.6027 | 0.63013 | 354.5675 | -0.76027 | -45.6161 | 0 | 0 | -0.37291 | 1.2481 | 56.46 | 12.0899
[6 empty] | 4.68 | 4.57 | 5 | 0 | 12.4506 | 0.62253 | 346.9648 | -0.74506 | -44.7037 | 0 | 0 | -0.36545 | 1.2445 | 56.51 | 12.0748
[6 empty] | 4.69 | 4.58 | 5 | 0 | 12.3016 | 0.61508 | 339.5142 | -0.73016 | -43.8097 | 0 | 0 | -0.35814 | 1.2409 | 56.56 | 12.0597
[6 empty] | 4.7 | 4.59 | 5 | 0 | 12.1556 | 0.60778 | 332.2126 | -0.71556 | -42.9335 | 0 | 0 | -0.35098 | 1.2374 | 56.61 | 12.0447
[6 empty] | 4.71 | 4.6 | 5 | 0 | 12.0125 | 0.60062 | 325.057 | -0.70125 | -42.0748 | 0 | 0 | -0.34396 | 1.2339 | 56.66 | 12.0297
[6 empty] | 4.72 | 4.61 | 5 | 0 | 11.8722 | 0.59361 | 318.0445 | -0.68722 | -41.2333 | 0 | 0 | -0.33708 | 1.2306 | 56.71 | 12.0148
[6 empty] | 4.73 | 4.62 | 5 | 0 | 11.7348 | 0.58674 | 311.1723 | -0.67348 | -40.4086 | 0 | 0 | -0.33034 | 1.2273 | 56.76 | 12
[6 empty] | 4.74 | 4.63 | 5 | 0 | 11.6001 | 0.58 | 304.4375 | -0.66001 | -39.6005 | 0 | 0 | -0.32373 | 1.224 | 56.81 | 11.9852
[6 empty] | 4.75 | 4.64 | 5 | 0 | 11.4681 | 0.5734 | 297.8375 | -0.64681 | -38.8085 | 0 | 0 | -0.31726 | 1.2209 | 56.86 | 11.9705
[6 empty] | 4.76 | 4.65 | 5 | 0 | 11.3387 | 0.56694 | 291.3694 | -0.63387 | -38.0323 | 0 | 0 | -0.31091 | 1.2177 | 56.91 | 11.9559
[6 empty] | 4.77 | 4.66 | 5 | 0 | 11.2119 | 0.5606 | 285.0307 | -0.62119 | -37.2716 | 0 | 0 | -0.3047 | 1.2147 | 56.96 | 11.9413
[6 empty] | 4.78 | 4.67 | 5 | 0 | 11.0877 | 0.55439 | 278.8187 | -0.60877 | -36.5262 | 0 | 0 | -0.2986 | 1.2117 | 57.01 | 11.9268
[6 empty] | 4.79 | 4.68 | 5 | 0 | 10.9659 | 0.5483 | 272.731 | -0.59659 | -35.7957 | 0 | 0 | -0.29263 | 1.2088 | 57.06 | 11.9123
[6 empty] | 4.8 | 4.69 | 5 | 0 | 10.8466 | 0.54233 | 266.7651 | -0.58466 | -35.0798 | 0 | 0 | -0.28678 | 1.2059 | 57.11 | 11.8979
[6 empty] | 4.81 | 4.7 | 5 | 0 | 10.7297 | 0.53648 | 260.9185 | -0.57297 | -34.3782 | 0 | 0 | -0.28104 | 1.2031 | 57.16 | 11.8836
[6 empty] | 4.82 | 4.71 | 5 | 0 | 10.6151 | 0.53076 | 255.1888 | -0.56151 | -33.6906 | 0 | 0 | -0.27542 | 1.2004 | 57.21 | 11.8693
[6 empty] | 4.83 | 4.72 | 5 | 0 | 10.5028 | 0.52514 | 249.5737 | -0.55028 | -33.0168 | 0 | 0 | -0.26991 | 1.1977 | 57.26 | 11.8551
[6 empty] | 4.84 | 4.73 | 5 | 0 | 10.3927 | 0.51964 | 244.0709 | -0.53927 | -32.3565 | 0 | 0 | -0.26451 | 1.195 | 57.31 | 11.8409
[6 empty] | 4.85 | 4.74 | 5 | 0 | 10.2849 | 0.51424 | 238.6781 | -0.52849 | -31.7093 | 0 | 0 | -0.25922 | 1.1924 | 57.36 | 11.8268
[6 empty] | 4.86 | 4.75 | 5 | 0 | 10.1792 | 0.50896 | 233.3932 | -0.51792 | -31.0751 | 0 | 0 | -0.25404 | 1.1899 | 57.41 | 11.8128
[6 empty] | 4.87 | 4.76 | 5 | 0 | 10.0756 | 0.50378 | 228.214 | -0.50756 | -30.4536 | 0 | 0 | -0.24896 | 1.1874 | 57.46 | 11.7988
[6 empty] | 4.88 | 4.77 | 5 | 0 | 9.9741 | 0.4987 | 223.1384 | -0.49741 | -29.8446 | 0 | 0 | -0.24398 | 1.1849 | 57.51 | 11.7848
[6 empty] | 4.89 | 4.78 | 5 | 0 | 9.8746 | 0.49373 | 218.1643 | -0.48746 | -29.2477 | 0 | 0 | -0.2391 | 1.1826 | 57.56 | 11.771
[6 empty] | 4.9 | 4.79 | 5 | 0 | 9.7771 | 0.48886 | 213.2897 | -0.47771 | -28.6627 | 0 | 0 | -0.23432 | 1.1802 | 57.61 | 11.7571
[6 empty] | 4.91 | 4.8 | 5 | 0 | 9.6816 | 0.48408 | 208.5126 | -0.46816 | -28.0895 | 0 | 0 | -0.22963 | 1.1779 | 57.66 | 11.7434
[6 empty] | 4.92 | 4.81 | 5 | 0 | 9.5879 | 0.4794 | 203.831 | -0.45879 | -27.5277 | 0 | 0 | -0.22504 | 1.1757 | 57.71 | 11.7297
[6 empty] | 4.93 | 4.82 | 5 | 0 | 9.4962 | 0.47481 | 199.2431 | -0.44962 | -26.9771 | 0 | 0 | -0.22054 | 1.1735 | 57.76 | 11.716
[6 empty] | 4.94 | 4.83 | 5 | 0 | 9.4063 | 0.47031 | 194.7469 | -0.44063 | -26.4376 | 0 | 0 | -0.21613 | 1.1713 | 57.81 | 11.7024
[6 empty] | 4.95 | 4.84 | 5 | 0 | 9.3181 | 0.46591 | 190.3406 | -0.43181 | -25.9088 | 0 | 0 | -0.2118 | 1.1692 | 57.86 | 11.6889
[6 empty] | 4.96 | 4.85 | 5 | 0 | 9.2318 | 0.46159 | 186.0225 | -0.42318 | -25.3907 | 0 | 0 | -0.20757 | 1.1671 | 57.91 | 11.6754
[6 empty] | 4.97 | 4.86 | 5 | 0 | 9.1471 | 0.45736 | 181.7907 | -0.41471 | -24.8828 | 0 | 0 | -0.20342 | 1.1651 | 57.96 | 11.662
[6 empty] | 4.98 | 4.87 | 5 | 0 | 9.0642 | 0.45321 | 177.6436 | -0.40642 | -24.3852 | 0 | 0 | -0.19935 | 1.1631 | 58.01 | 11.6486
[6 empty] | 4.99 | 4.88 | 5 | 0 | 8.9829 | 0.44915 | 173.5794 | -0.39829 | -23.8975 | 0 | 0 | -0.19536 | 1.1611 | 58.06 | 11.6353
[6 empty] | 5 | 4.89 | 5 | 0 | 8.9033 | 0.44516 | 169.5965 | -0.39033 | -23.4195 | 0 | 0 | -0.19145 | 1.1592 | 58.11 | 11.622
[6 empty] | 5.01 | 4.9 | 5 | 0 | 8.8252 | 0.44126 | 165.6932 | -0.38252 | -22.9511 | 0 | 0 | -0.18763 | 1.1573 | 58.16 | 11.6088
[6 empty] | 5.02 | 4.91 | 5 | 0 | 8.7487 | 0.43743 | 161.868 | -0.37487 | -22.4921 | 0 | 0 | -0.18387 | 1.1555 | 58.21 | 11.5956
[6 empty] | 5.03 | 4.92 | 5 | 0 | 8.6737 | 0.43369 | 158.1193 | -0.36737 | -22.0423 | 0 | 0 | -0.1802 | 1.1537 | 58.26 | 11.5825
[6 empty] | 5.04 | 4.93 | 5 | 0 | 8.6002 | 0.43001 | 154.4456 | -0.36002 | -21.6014 | 0 | 0 | -0.17659 | 1.1519 | 58.31 | 11.5694
[6 empty] | 5.05 | 4.94 | 5 | 0 | 8.5282 | 0.42641 | 150.8454 | -0.35282 | -21.1694 | 0 | 0 | -0.17306 | 1.1502 | 58.36 | 11.5564
[6 empty] | 5.06 | 4.95 | 5 | 0 | 8.4577 | 0.42288 | 147.3171 | -0.34577 | -20.746 | 0 | 0 | -0.1696 | 1.1485 | 58.41 | 11.5435
[6 empty] | 5.07 | 4.96 | 5 | 0 | 8.3885 | 0.41943 | 143.8595 | -0.33885 | -20.3311 | 0 | 0 | -0.16621 | 1.1468 | 58.46 | 11.5306
[6 empty] | 5.08 | 4.97 | 5 | 0 | 8.3207 | 0.41604 | 140.471 | -0.33207 | -19.9245 | 0 | 0 | -0.16288 | 1.1452 | 58.51 | 11.5177
[6 empty] | 5.09 | 4.98 | 5 | 0 | 8.2543 | 0.41272 | 137.1502 | -0.32543 | -19.526 | 0 | 0 | -0.15962 | 1.1436 | 58.56 | 11.5049
[6 empty] | 5.1 | 4.99 | 5 | 0 | 8.1892 | 0.40946 | 133.8959 | -0.31892 | -19.1355 | 0 | 0 | -0.15643 | 1.1421 | 58.61 | 11.4922
[6 empty] | 5.11 | 5 | 5 | 0 | 8.1255 | 0.40627 | 130.7066 | -0.31255 | -18.7528 | 0 | 0 | -0.1533 | 1.1405 | 58.66 | 11.4795
[6 empty] | 5.12 | 5.01 | 5 | 0 | 8.0629 | 0.40315 | 127.5812 | -0.30629 | -18.3777 | 0 | 0 | -0.15024 | 1.139 | 58.71 | 11.4668
[6 empty] | 5.13 | 5.02 | 5 | 0 | 8.0017 | 0.40008 | 124.5182 | -0.30017 | -18.0101 | 0 | 0 | -0.14723 | 1.1375 | 58.76 | 11.4542
[6 empty] | 5.14 | 5.03 | 5 | 0 | 7.9417 | 0.39708 | 121.5165 | -0.29417 | -17.6499 | 0 | 0 | -0.14429 | 1.1361 | 58.81 | 11.4416
[6 empty] | 5.15 | 5.04 | 5 | 0 | 7.8828 | 0.39414 | 118.5749 | -0.28828 | -17.2969 | 0 | 0 | -0.1414 | 1.1347 | 58.86 | 11.4291
[6 empty] | 5.16 | 5.05 | 5 | 0 | 7.8252 | 0.39126 | 115.6921 | -0.28252 | -16.951 | 0 | 0 | -0.13857 | 1.1333 | 58.91 | 11.4167
[6 empty] | 5.17 | 5.06 | 5 | 0 | 7.7687 | 0.38843 | 112.8669 | -0.27687 | -16.612 | 0 | 0 | -0.1358 | 1.1319 | 58.96 | 11.4043
[6 empty] | 5.18 | 5.07 | 5 | 0 | 7.7133 | 0.38566 | 110.0982 | -0.27133 | -16.2797 | 0 | 0 | -0.13309 | 1.1306 | 59.01 | 11.3919
[6 empty] | 5.19 | 5.08 | 5 | 0 | 7.659 | 0.38295 | 107.3849 | -0.2659 | -15.9542 | 0 | 0 | -0.13043 | 1.1293 | 59.06 | 11.3796
[6 empty] | 5.2 | 5.09 | 5 | 0 | 7.6058 | 0.38029 | 104.7259 | -0.26058 | -15.6351 | 0 | 0 | -0.12782 | 1.128 | 59.11 | 11.3673
[6 empty] | 5.21 | 5.1 | 5 | 0 | 7.5537 | 0.37769 | 102.1201 | -0.25537 | -15.3224 | 0 | 0 | -0.12526 | 1.1268 | 59.16 | 11.3551
[6 empty] | 5.22 | 5.11 | 5 | 0 | 7.5027 | 0.37513 | 99.5663 | -0.25027 | -15.0159 | 0 | 0 | -0.12276 | 1.1255 | 59.21 | 11.3429
[6 empty] | 5.23 | 5.12 | 5 | 0 | 7.4526 | 0.37263 | 97.0637 | -0.24526 | -14.7156 | 0 | 0 | -0.1203 | 1.1243 | 59.26 | 11.3308
[6 empty] | 5.24 | 5.13 | 5 | 0 | 7.4035 | 0.37018 | 94.6111 | -0.24035 | -14.4213 | 0 | 0 | -0.11789 | 1.1232 | 59.31 | 11.3187
[6 empty] | 5.25 | 5.14 | 5 | 0 | 7.3555 | 0.36777 | 92.2075 | -0.23555 | -14.1329 | 0 | 0 | -0.11554 | 1.122 | 59.36 | 11.3067
[6 empty] | 5.26 | 5.15 | 5 | 0 | 7.3084 | 0.36542 | 89.8521 | -0.23084 | -13.8502 | 0 | 0 | -0.11323 | 1.1209 | 59.41 | 11.2947
[6 empty] | 5.27 | 5.16 | 5 | 0 | 7.2622 | 0.36311 | 87.5437 | -0.22622 | -13.5732 | 0 | 0 | -0.11096 | 1.1198 | 59.46 | 11.2827
[6 empty] | 5.28 | 5.17 | 5 | 0 | 7.217 | 0.36085 | 85.2815 | -0.2217 | -13.3017 | 0 | 0 | -0.10874 | 1.1187 | 59.51 | 11.2708
[6 empty] | 5.29 | 5.18 | 5 | 0 | 7.1726 | 0.35863 | 83.0645 | -0.21726 | -13.0357 | 0 | 0 | -0.10657 | 1.1176 | 59.56 | 11.259
[6 empty] | 5.3 | 5.19 | 5 | 0 | 7.1292 | 0.35646 | 80.8919 | -0.21292 | -12.775 | 0 | 0 | -0.10444 | 1.1166 | 59.61 | 11.2472
[6 empty] | 5.31 | 5.2 | 5 | 0 | 7.0866 | 0.35433 | 78.7627 | -0.20866 | -12.5195 | 0 | 0 | -0.10235 | 1.1155 | 59.66 | 11.2354
[6 empty] | 5.32 | 5.21 | 5 | 0 | 7.0448 | 0.35224 | 76.6762 | -0.20448 | -12.2691 | 0 | 0 | -0.1003 | 1.1145 | 59.71 | 11.2237
[6 empty] | 5.33 | 5.22 | 5 | 0 | 7.004 | 0.3502 | 74.6313 | -0.2004 | -12.0237 | 0 | 0 | -0.098294 | 1.1136 | 59.76 | 11.212
[6 empty] | 5.34 | 5.23 | 5 | 0 | 6.9639 | 0.34819 | 72.6274 | -0.19639 | -11.7832 | 0 | 0 | -0.096328 | 1.1126 | 59.81 | 11.2004
[6 empty] | 5.35 | 5.24 | 5 | 0 | 6.9246 | 0.34623 | 70.6635 | -0.19246 | -11.5476 | 0 | 0 | -0.094401 | 1.1117 | 59.86 | 11.1888
[6 empty] | 5.36 | 5.25 | 5 | 0 | 6.8861 | 0.34431 | 68.7389 | -0.18861 | -11.3166 | 0 | 0 | -0.092513 | 1.1107 | 59.91 | 11.1772
[6 empty] | 5.37 | 5.26 | 5 | 0 | 6.8484 | 0.34242 | 66.8528 | -0.18484 | -11.0903 | 0 | 0 | -0.090663 | 1.1098 | 59.96 | 11.1657
[6 empty] | 5.38 | 5.27 | 5 | 0 | 6.8114 | 0.34057 | 65.0044 | -0.18114 | -10.8685 | 0 | 0 | -0.08885 | 1.1089 | 60.01 | 11.1543
[6 empty] | 5.39 | 5.28 | 5 | 0 | 6.7752 | 0.33876 | 63.193 | -0.17752 | -10.6511 | 0 | 0 | -0.087073 | 1.1081 | 60.06 | 11.1429
[6 empty] | 5.4 | 5.29 | 5 | 0 | 6.7397 | 0.33698 | 61.4178 | -0.17397 | -10.4381 | 0 | 0 | -0.085331 | 1.1072 | 60.11 | 11.1315
[6 empty] | 5.41 | 5.3 | 5 | 0 | 6.7049 | 0.33524 | 59.6781 | -0.17049 | -10.2293 | 0 | 0 | -0.083625 | 1.1064 | 60.16 | 11.1201
[6 empty] | 5.42 | 5.31 | 5 | 0 | 6.6708 | 0.33354 | 57.9732 | -0.16708 | -10.0247 | 0 | 0 | -0.081952 | 1.1056 | 60.21 | 11.1089
[6 empty] | 5.43 | 5.32 | 5 | 0 | 6.6374 | 0.33187 | 56.3024 | -0.16374 | -9.8243 | 0 | 0 | -0.080313 | 1.1048 | 60.26 | 11.0976
[6 empty] | 5.44 | 5.33 | 5 | 0 | 6.6046 | 0.33023 | 54.6651 | -0.16046 | -9.6278 | 0 | 0 | -0.078707 | 1.104 | 60.31 | 11.0864
[6 empty] | 5.45 | 5.34 | 5 | 0 | 6.5725 | 0.32863 | 53.0604 | -0.15725 | -9.4352 | 0 | 0 | -0.077133 | 1.1032 | 60.36 | 11.0752
[6 empty] | 5.46 | 5.35 | 5 | 0 | 6.5411 | 0.32705 | 51.4879 | -0.15411 | -9.2465 | 0 | 0 | -0.07559 | 1.1024 | 60.41 | 11.0641
[6 empty] | 5.47 | 5.36 | 5 | 0 | 6.5103 | 0.32551 | 49.9468 | -0.15103 | -9.0616 | 0 | 0 | -0.074078 | 1.1017 | 60.46 | 11.053
[6 empty] | 5.48 | 5.37 | 5 | 0 | 6.4801 | 0.324 | 48.4366 | -0.14801 | -8.8803 | 0 | 0 | -0.072597 | 1.101 | 60.51 | 11.042
[6 empty] | 5.49 | 5.38 | 5 | 0 | 6.4505 | 0.32252 | 46.9565 | -0.14505 | -8.7027 | 0 | 0 | -0.071145 | 1.1003 | 60.56 | 11.031
[6 empty] | 5.5 | 5.39 | 5 | 0 | 6.4214 | 0.32107 | 45.506 | -0.14214 | -8.5287 | 0 | 0 | -0.069722 | 1.0996 | 60.61 | 11.02
[6 empty] | 5.51 | 5.4 | 5 | 0 | 6.393 | 0.31965 | 44.0846 | -0.1393 | -8.3581 | 0 | 0 | -0.068328 | 1.0989 | 60.66 | 11.0091
[6 empty] | 5.52 | 5.41 | 5 | 0 | 6.3652 | 0.31826 | 42.6916 | -0.13652 | -8.1909 | 0 | 0 | -0.066961 | 1.0982 | 60.71 | 10.9982
[6 empty] | 5.53 | 5.42 | 5 | 0 | 6.3379 | 0.31689 | 41.3264 | -0.13379 | -8.0271 | 0 | 0 | -0.065622 | 1.0976 | 60.76 | 10.9873
[6 empty] | 5.54 | 5.43 | 5 | 0 | 6.3111 | 0.31555 | 39.9886 | -0.13111 | -7.8666 | 0 | 0 | -0.064309 | 1.0969 | 60.81 | 10.9765
[6 empty] | 5.55 | 5.44 | 5 | 0 | 6.2849 | 0.31424 | 38.6775 | -0.12849 | -7.7093 | 0 | 0 | -0.063023 | 1.0963 | 60.86 | 10.9658
[6 empty] | 5.56 | 5.45 | 5 | 0 | 6.2592 | 0.31296 | 37.3926 | -0.12592 | -7.5551 | 0 | 0 | -0.061763 | 1.0957 | 60.91 | 10.955
[6 empty] | 5.57 | 5.46 | 5 | 0 | 6.234 | 0.3117 | 36.1334 | -0.1234 | -7.404 | 0 | 0 | -0.060527 | 1.0951 | 60.96 | 10.9443
[6 empty] | 5.58 | 5.47 | 5 | 0 | 6.2093 | 0.31047 | 34.8994 | -0.12093 | -7.2559 | 0 | 0 | -0.059317 | 1.0945 | 61.01 | 10.9337
[6 empty] | 5.59 | 5.48 | 5 | 0 | 6.1851 | 0.30926 | 33.6901 | -0.11851 | -7.1108 | 0 | 0 | -0.058131 | 1.0939 | 61.06 | 10.9231
[6 empty] | 5.6 | 5.49 | 5 | 0 | 6.1614 | 0.30807 | 32.505 | -0.11614 | -6.9686 | 0 | 0 | -0.056968 | 1.0933 | 61.11 | 10.9125
[6 empty] | 5.61 | 5.5 | 5 | 0 | 6.1382 | 0.30691 | 31.3435 | -0.11382 | -6.8292 | 0 | 0 | -0.055829 | 1.0928 | 61.16 | 10.902
[6 empty] | 5.62 | 5.51 | 5 | 0 | 6.1154 | 0.30577 | 30.2054 | -0.11154 | -6.6926 | 0 | 0 | -0.054712 | 1.0922 | 61.21 | 10.8915
[6 empty] | 5.63 | 5.52 | 5 | 0 | 6.0931 | 0.30466 | 29.0899 | -0.10931 | -6.5587 | 0 | 0 | -0.053618 | 1.0917 | 61.26 | 10.881
[6 empty] | 5.64 | 5.53 | 5 | 0 | 6.0713 | 0.30356 | 27.9968 | -0.10713 | -6.4276 | 0 | 0 | -0.052545 | 1.0911 | 61.31 | 10.8706
[6 empty] | 5.65 | 5.54 | 5 | 0 | 6.0498 | 0.30249 | 26.9255 | -0.10498 | -6.299 | 0 | 0 | -0.051495 | 1.0906 | 61.36 | 10.8602
[6 empty] | 5.66 | 5.55 | 5 | 0 | 6.0288 | 0.30144 | 25.8757 | -0.10288 | -6.173 | 0 | 0 | -0.050465 | 1.0901 | 61.41 | 10.8498
[6 empty] | 5.67 | 5.56 | 5 | 0 | 6.0083 | 0.30041 | 24.8469 | -0.10083 | -6.0496 | 0 | 0 | -0.049455 | 1.0896 | 61.46 | 10.8395
[6 empty] | 5.68 | 5.57 | 5 | 0 | 5.9881 | 0.2994 | 23.8386 | -0.09881 | -5.9286 | 0 | 0 | -0.048466 | 1.0891 | 61.51 | 10.8292
[6 empty] | 5.69 | 5.58 | 5 | 0 | 5.9683 | 0.29842 | 22.8505 | -0.096834 | -5.81 | 0 | 0 | -0.047497 | 1.0887 | 61.56 | 10.819
[6 empty] | 5.7 | 5.59 | 5 | 0 | 5.949 | 0.29745 | 21.8822 | -0.094897 | -5.6938 | 0 | 0 | -0.046547 | 1.0882 | 61.61 | 10.8088
[6 empty] | 5.71 | 5.6 | 5 | 0 | 5.93 | 0.2965 | 20.9332 | -0.092999 | -5.5799 | 0 | 0 | -0.045616 | 1.0878 | 61.66 | 10.7986
[6 empty] | 5.72 | 5.61 | 5 | 0 | 5.9114 | 0.29557 | 20.0032 | -0.091139 | -5.4683 | 0 | 0 | -0.044704 | 1.0873 | 61.71 | 10.7885
[6 empty] | 5.73 | 5.62 | 5 | 0 | 5.8932 | 0.29466 | 19.0918 | -0.089316 | -5.359 | 0 | 0 | -0.04381 | 1.0869 | 61.76 | 10.7784
[6 empty] | 5.74 | 5.63 | 5 | 0 | 5.8753 | 0.29376 | 18.1986 | -0.08753 | -5.2518 | 0 | 0 | -0.042933 | 1.0864 | 61.81 | 10.7683
[6 empty] | 5.75 | 5.64 | 5 | 0 | 5.8578 | 0.29289 | 17.3233 | -0.085779 | -5.1468 | 0 | 0 | -0.042075 | 1.086 | 61.86 | 10.7583
[6 empty] | 5.76 | 5.65 | 5 | 0 | 5.8406 | 0.29203 | 16.4656 | -0.084064 | -5.0438 | 0 | 0 | -0.041233 | 1.0856 | 61.91 | 10.7483
[6 empty] | 5.77 | 5.66 | 5 | 0 | 5.8238 | 0.29119 | 15.6249 | -0.082382 | -4.9429 | 0 | 0 | -0.040409 | 1.0852 | 61.96 | 10.7383
[6 empty] | 5.78 | 5.67 | 5 | 0 | 5.8073 | 0.29037 | 14.8011 | -0.080735 | -4.8441 | 0 | 0 | -0.0396 | 1.0848 | 62.01 | 10.7284
[6 empty] | 5.79 | 5.68 | 5 | 0 | 5.7912 | 0.28956 | 13.9937 | -0.07912 | -4.7472 | 0 | 0 | -0.038808 | 1.0844 | 62.06 | 10.7185
[6 empty] | 5.8 | 5.69 | 5 | 0 | 5.7754 | 0.28877 | 13.2025 | -0.077538 | -4.6523 | 0 | 0 | -0.038032 | 1.084 | 62.11 | 10.7086
[6 empty] | 5.81 | 5.7 | 5 | 0 | 5.7599 | 0.28799 | 12.4272 | -0.075987 | -4.5592 | 0 | 0 | -0.037272 | 1.0837 | 62.16 | 10.6988
[6 empty] | 5.82 | 5.71 | 5 | 0 | 5.7447 | 0.28723 | 11.6673 | -0.074467 | -4.468 | 0 | 0 | -0.036526 | 1.0833 | 62.21 | 10.689
[6 empty] | 5.83 | 5.72 | 5 | 0 | 5.7298 | 0.28649 | 10.9226 | -0.072978 | -4.3787 | 0 | 0 | -0.035796 | 1.0829 | 62.26 | 10.6792
[6 empty] | 5.84 | 5.73 | 5 | 0 | 5.7152 | 0.28576 | 10.1928 | -0.071518 | -4.2911 | 0 | 0 | -0.03508 | 1.0826 | 62.31 | 10.6695
[6 empty] | 5.85 | 5.74 | 5 | 0 | 5.7009 | 0.28504 | 9.4777 | -0.070088 | -4.2053 | 0 | 0 | -0.034378 | 1.0822 | 62.36 | 10.6598
[6 empty] | 5.86 | 5.75 | 5 | 0 | 5.6869 | 0.28434 | 8.7768 | -0.068686 | -4.1212 | 0 | 0 | -0.033691 | 1.0819 | 62.41 | 10.6502
[6 empty] | 5.87 | 5.76 | 5 | 0 | 5.6731 | 0.28366 | 8.0899 | -0.067312 | -4.0387 | 0 | 0 | -0.033017 | 1.0816 | 62.46 | 10.6405
[6 empty] | 5.88 | 5.77 | 5 | 0 | 5.6597 | 0.28298 | 7.4168 | -0.065966 | -3.958 | 0 | 0 | -0.032356 | 1.0813 | 62.51 | 10.631
[6 empty] | 5.89 | 5.78 | 5 | 0 | 5.6465 | 0.28232 | 6.7571 | -0.064647 | -3.8788 | 0 | 0 | -0.031709 | 1.0809 | 62.56 | 10.6214
[6 empty] | 5.9 | 5.79 | 5 | 0 | 5.6335 | 0.28168 | 6.1107 | -0.063354 | -3.8012 | 0 | 0 | -0.031075 | 1.0806 | 62.61 | 10.6119 |  |  | y: 0.5
[6 empty] | 5.91 | 5.8 | 5 | 0 | 5.6209 | 0.28104 | 5.4771 | -0.062087 | -3.7252 | 0 | 0 | -0.030454 | 1.0803 | 62.66 | 10.6024 |  |  | y: 0.5
[6 empty] | 5.92 | 5.81 | 5 | 0 | 5.6085 | 0.28042 | 4.8563 | -0.060845 | -3.6507 | 0 | 0 | -0.029845 | 1.08 | 62.71 | 10.5929 |  |  | y: 0.5
[6 empty] | 5.93 | 5.82 | 5 | 0 | 5.5963 | 0.27981 | 4.2478 | -0.059628 | -3.5777 | 0 | 0 | -0.029248 | 1.0797 | 62.76 | 10.5835 |  |  | y: 0.5
[6 empty] | 5.94 | 5.83 | 5 | 0 | 5.5844 | 0.27922 | 3.6515 | -0.058436 | -3.5061 | 0 | 0 | -0.028663 | 1.0794 | 62.81 | 10.5741 |  |  | y: 0.5
[6 empty] | 5.95 | 5.84 | 5 | 0 | 5.5727 | 0.27863 | 3.0672 | -0.057267 | -3.436 | 0 | 0 | -0.028089 | 1.0792 | 62.86 | 10.5647 |  |  | y: 0.5
[6 empty] | 5.96 | 5.85 | 5 | 0 | 5.5612 | 0.27806 | 2.4945 | -0.056122 | -3.3673 | 0 | 0 | -0.027528 | 1.0789 | 62.91 | 10.5554 |  |  | y: 0.5
[6 empty] | 5.97 | 5.86 | 5 | 0 | 5.55 | 0.2775 | 1.9333 | -0.054999 | -3.3 | 0 | 0 | -0.026977 | 1.0786 | 62.96 | 10.5461 |  |  | y: 0.5
[6 empty] | 5.98 | 5.87 | 5 | 0 | 5.539 | 0.27695 | 1.3833 | -0.053899 | -3.234 | 0 | 0 | -0.026438 | 1.0784 | 63.01 | 10.5368
[6 empty] | 5.99 | 5.88 | 5 | 0 | 5.5282 | 0.27641 | 0.84429 | -0.052821 | -3.1693 | 0 | 0 | -0.025909 | 1.0781 | 63.06 | 10.5275
[6 empty] | 6 | 5.89 | 5 | 0 | 5.5176 | 0.27588 | 0.31608 | -0.051765 | -3.1059 | 0 | 0 | -0.025391 | 1.0778 | 63.11 | 10.5183
Menu
Passive Pressure -Volume Curve
Blue = Static Green = Dynamic

## PC (IP)
Pressure Control
Lung Model
|  | Elastance (cm H2O/L) | 20
|  | Compliance (L/cm H2O) | 0.05 |  | 50
|  | Resistance (cm H2O/L/s) | 10 |  | 10
|  | Time Constant (s) | 0.5
|  | Tidal Volume (mL) | 432.7303
|  | Potential display error (%) | 0.30486 |  | Ref.
Effort Model
|  | Delay Before Effort (s) | 0 |  | 0
|  | Duration of Effort (s) | 0.3 |  | 6
|  | Freq (hz) | 0.83333
|  | Period (s) | 0.6
|  | max Pmus (cm H2O) | 0 | 49 | 0
Ventilator Settings
|  | Inspiratory Pressure Target (cm H2O) | 10
|  | Peak Inspiratory Pressure (cm H2O) | 20 |  | 10
|  | Frequency (breaths/min) | 17 |  | 17
|  | Inflation Pressure Time (s) | 1 |  | 10
|  | PEEP (cm H2O) | 10 |  | 10
|  | autoPEEP (cm H2O) | 0.15627
|  | Pause Time (s) | 0 |  | 0
|  | Peak Airway Pressure (cm H2O) | 20
|  | Total Cycle Time (s) | 3.5294
|  | Inspirtory Time (s) | 1
|  | Expiratory Time (s) | 2.5294
|  | I:E | 1:2.5
|  | Duty Cycle (%) | 28.3333
|  | Plateau Pressure (cm H2O) | 18.6808
|  | Tidal (Driving) Pressure (cm H2O) | 8.6546
|  | Peak Inspiratory Flow (L/min) | 59.6847
|  | Peak Expiratory Flow (L/min) | -52.2429
|  | Mean Inspiratory Pressure | 20
|  | Mean Airway Pressure | 12.849
|  | End Inspiratory Volume (mL) | 431.4111
|  | Inspiratory Flow Time (s) | 1
Passive |  | Ventilator Elastic Work (J) | 0.57355
|  | Ventilator Resistive Work (J) | 0.23793
|  | Ventilator Total Work (J) | 0.81147
Active |  | Ventilator Work (J) | 0.84902
|  | Patient Work (J) | 0
|  | Total Work (J) | 0.84902
|  | Work Shift Index (%) | 0
|  | Ventilator Inspiratory Power (W) | 0.24055
|  | Ventilator Inspiratory Power (J/min) | 14.4333
pause start time
1
[11 empty] | pause stop time | actual VT | [7 empty] | Mean Arwy Pressure
[11 empty] | 1 | 0.43273 |  | Peak Insp Flow |  |  |  | Elastance |  | 12.849 |  |  | …………………….Pressure Volume Curve Data……………….
[9 empty] | Vmus equiv to constant mean Pmus |  | Initial | Passive VT |  | 59.6847 |  |  |  | 0.02 |  | Mean Insp Pressure | [6 empty] | Seg |  | Seg |  | Seg |  | Seg
[9 empty] | 0 |  | 0.50263 | 413.5946 |  | Peak Exp Flow | [5 empty] | 20 | Pinsp | 10 | Pstat |  | Pdyn |  | L1 |  | L2 |  | U1 |  | U2
[11 empty] | includes est autoPEEP |  |  | -52.2429 |  |  |  |  | Integral | Mean | R | 10 | X | Y | X | Y | x | y | x | y | x | y | x | y
[8 empty] | Vent | Muscle | Lung | Lung | Tidal | Tidal | Tidal | Muscle | Muscle | Ventilator |  | Total | Total | C | 0.05 | 10 | 0 | 20 | 0 | 0 | 0 | 10 | 0 | 0 | 413.5946 | 18.2719 | 413.5946
[7 empty] | Relative | Pressure | Pressure | Pressure | Volume | Volume | Flow | Flow | Power | Work | Power | Work | Pressure | Pressure | RC | 0.5 | 18.2719 | 413.5946 | 20 | 413.5946 | 10 | 0 | 20 | 0 | 18.2719 | 413.5946 | 20 | 413.5946
[6 empty] | Time | Time | (cm H2O) | (cm H2O) | (cm H2O) | (L) | (mL) | (L/s) | (L/min) | (W) | (J) | (W) | (J) | (cm H2O) | (cm H2O) | Inspiratory Pressure Time | 1
[6 empty] | 0 | 0 | 20 | 0 | 10.0526 | 0.50263 | 0 | 0.99474 | 59.6847 | 0 | 0 | 1.9517 | 0.019517 | 0 | 0 | Total Cycle Time | 3.5294
[6 empty] | 0.01 | 0.01 | 20 | 0 | 10.2515 | 0.51258 | 9.9474 | 0.97485 | 58.491 | 0 | 0 | 1.9127 | 0.038643 | 0.2 | 20 | D | 0.28333
[6 empty] | 0.02 | 0.02 | 20 | 0 | 10.4465 | 0.52232 | 19.6959 | 0.95535 | 57.3212 | 0 | 0 | 1.8744 | 0.057387 | 0.4 | 20
[6 empty] | 0.03 | 0.03 | 20 | 0 | 10.6375 | 0.53188 | 29.2495 | 0.93625 | 56.1748 | 0 | 0 | 1.8369 | 0.075757 | 0.6 | 20 | f | 0.28333
[6 empty] | 0.04 | 0.04 | 20 | 0 | 10.8248 | 0.54124 | 38.6119 | 0.91752 | 55.0513 | 0 | 0 | 1.8002 | 0.093758 | 0.8 | 20 | fRC | 0.14167
[6 empty] | 0.05 | 0.05 | 20 | 0 | 11.0083 | 0.55041 | 47.7871 | 0.89917 | 53.9502 | 0 | 0 | 1.7642 | 0.1114 | 1 | 20 | -D/fRC | -2
[6 empty] | 0.06 | 0.06 | 20 | 0 | 11.1881 | 0.55941 | 56.7789 | 0.88119 | 52.8712 | 0 | 0 | 1.7289 | 0.12869 | 1.2 | 20 | -1-D/fRC | -3
[6 empty] | 0.07 | 0.07 | 20 | 0 | 11.3644 | 0.56822 | 65.5907 | 0.86356 | 51.8138 | 0 | 0 | 1.6943 | 0.14563 | 1.4 | 20 | exp(-D/fRC) | 0.13534
[6 empty] | 0.08 | 0.08 | 20 | 0 | 11.5371 | 0.57685 | 74.2264 | 0.84629 | 50.7775 | 0 | 0 | 1.6604 | 0.16224 | 1.6 | 20 | exp(-1-D/fRC) | 0.049787
[6 empty] | 0.09 | 0.09 | 20 | 0 | 11.7063 | 0.58532 | 82.6893 | 0.82937 | 49.762 | 0 | 0 | 1.6272 | 0.17851 | 1.8 | 20
[6 empty] | 0.1 | 0.1 | 20 | 0 | 11.8722 | 0.59361 | 90.9829 | 0.81278 | 48.7667 | 0 | 0 | 1.5947 | 0.19446 | 2 | 20 | 1-exp(-D/fRC) | 0.86466
[6 empty] | 0.11 | 0.11 | 20 | 0 | 12.0348 | 0.60174 | 99.1107 | 0.79652 | 47.7914 | 0 | 0 | 1.5628 | 0.21008 | 2.2 | 20 | 1-exp(-1-D/fRC) | 0.95021
[6 empty] | 0.12 | 0.12 | 20 | 0 | 12.1941 | 0.6097 | 107.076 | 0.78059 | 46.8356 | 0 | 0 | 1.5315 | 0.2254 | 2.4 | 20 | exp(-1-D/fRC) | 0.049787
[6 empty] | 0.13 | 0.13 | 20 | 0 | 12.3502 | 0.61751 | 114.8819 | 0.76498 | 45.8989 | 0 | 0 | 1.5009 | 0.24041 | 2.6 | 20
[6 empty] | 0.14 | 0.14 | 20 | 0 | 12.5032 | 0.62516 | 122.5317 | 0.74968 | 44.9809 | 0 | 0 | 1.4709 | 0.25512 | 2.8 | 20 | 1-exp(-D/fRC)*exp(-1-D/fRC) | 0.82162
[6 empty] | 0.15 | 0.15 | 20 | 0 | 12.6531 | 0.63266 | 130.0285 | 0.73469 | 44.0813 | 0 | 0 | 1.4415 | 0.26953 | 3 | 20 | (1-exp(-D/fRC))*(exp(-1-D/fRC)) | 0.99326
[6 empty] | 0.16 | 0.16 | 20 | 0 | 12.8001 | 0.64 | 137.3754 | 0.71999 | 43.1996 | 0 | 0 | 1.4126 | 0.28366 | 3.2 | 20
[6 empty] | 0.17 | 0.17 | 20 | 0 | 12.9441 | 0.6472 | 144.5753 | 0.70559 | 42.3357 | 0 | 0 | 1.3844 | 0.2975 | 3.4 | 20 | VT | 0.41359
[6 empty] | 0.18 | 0.18 | 20 | 0 | 13.0852 | 0.65426 | 151.6313 | 0.69148 | 41.4889 | 0 | 0 | 1.3567 | 0.31107 | 3.6 | 20
[6 empty] | 0.19 | 0.19 | 20 | 0 | 13.2235 | 0.66117 | 158.5461 | 0.67765 | 40.6592 | 0 | 0 | 1.3296 | 0.32436 | 3.8 | 20
[6 empty] | 0.2 | 0.2 | 20 | 0 | 13.359 | 0.66795 | 165.3226 | 0.6641 | 39.846 | 0 | 0 | 1.303 | 0.33739 | 4 | 20
[6 empty] | 0.21 | 0.21 | 20 | 0 | 13.4918 | 0.67459 | 171.9636 | 0.65082 | 39.0491 | 0 | 0 | 1.2769 | 0.35016 | 4.2 | 20
[6 empty] | 0.22 | 0.22 | 20 | 0 | 13.622 | 0.6811 | 178.4718 | 0.6378 | 38.2681 | 0 | 0 | 1.2514 | 0.36268 | 4.4 | 20
[6 empty] | 0.23 | 0.23 | 20 | 0 | 13.7495 | 0.68748 | 184.8498 | 0.62505 | 37.5027 | 0 | 0 | 1.2263 | 0.37494 | 4.6 | 20
[6 empty] | 0.24 | 0.24 | 20 | 0 | 13.8746 | 0.69373 | 191.1003 | 0.61254 | 36.7527 | 0 | 0 | 1.2018 | 0.38696 | 4.8 | 20
[6 empty] | 0.25 | 0.25 | 20 | 0 | 13.9971 | 0.69985 | 197.2257 | 0.60029 | 36.0176 | 0 | 0 | 1.1778 | 0.39873 | 5 | 20
[6 empty] | 0.26 | 0.26 | 20 | 0 | 14.1171 | 0.70586 | 203.2286 | 0.58829 | 35.2973 | 0 | 0 | 1.1542 | 0.41028 | 5.2 | 20
[6 empty] | 0.27 | 0.27 | 20 | 0 | 14.2348 | 0.71174 | 209.1115 | 0.57652 | 34.5913 | 0 | 0 | 1.1311 | 0.42159 | 5.4 | 20
[6 empty] | 0.28 | 0.28 | 20 | 0 | 14.3501 | 0.7175 | 214.8767 | 0.56499 | 33.8995 | 0 | 0 | 1.1085 | 0.43267 | 5.6 | 20
[6 empty] | 0.29 | 0.29 | 20 | 0 | 14.4631 | 0.72315 | 220.5266 | 0.55369 | 33.2215 | 0 | 0 | 1.0863 | 0.44354 | 5.8 | 20
[6 empty] | 0.3 | 0.3 | 20 | 0 | 14.5738 | 0.72869 | 226.0636 | 0.54262 | 32.5571 | 0 | 0 | 1.0646 | 0.45418 | 6 | 20
[6 empty] | 0.31 | 0.31 | 20 | 0 | 14.6823 | 0.73412 | 231.4897 | 0.53177 | 31.9059 | 0 | 0 | 1.0433 | 0.46462 | 6.2 | 20
[6 empty] | 0.32 | 0.32 | 20 | 0 | 14.7887 | 0.73943 | 236.8074 | 0.52113 | 31.2678 | 0 | 0 | 1.0225 | 0.47484 | 6.4 | 20
[6 empty] | 0.33 | 0.33 | 20 | 0 | 14.8929 | 0.74465 | 242.0187 | 0.51071 | 30.6424 | 0 | 0 | 1.002 | 0.48486 | 6.6 | 20
[6 empty] | 0.34 | 0.34 | 20 | 0 | 14.9951 | 0.74975 | 247.1258 | 0.50049 | 30.0296 | 0 | 0 | 0.98197 | 0.49468 | 6.8 | 20
[6 empty] | 0.35 | 0.35 | 20 | 0 | 15.0952 | 0.75476 | 252.1307 | 0.49048 | 29.429 | 0 | 0 | 0.96233 | 0.5043 | 7 | 20
[6 empty] | 0.36 | 0.36 | 20 | 0 | 15.1933 | 0.75966 | 257.0355 | 0.48067 | 28.8404 | 0 | 0 | 0.94308 | 0.51373 | 7.2 | 20
[6 empty] | 0.37 | 0.37 | 20 | 0 | 15.2894 | 0.76447 | 261.8423 | 0.47106 | 28.2636 | 0 | 0 | 0.92422 | 0.52298 | 7.4 | 20
[6 empty] | 0.38 | 0.38 | 20 | 0 | 15.3836 | 0.76918 | 266.5529 | 0.46164 | 27.6983 | 0 | 0 | 0.90574 | 0.53203 | 7.6 | 20
[6 empty] | 0.39 | 0.39 | 20 | 0 | 15.4759 | 0.7738 | 271.1693 | 0.45241 | 27.1444 | 0 | 0 | 0.88762 | 0.54091 | 7.8 | 20
[6 empty] | 0.4 | 0.4 | 20 | 0 | 15.5664 | 0.77832 | 275.6933 | 0.44336 | 26.6015 | 0 | 0 | 0.86987 | 0.54961 | 8 | 20
[6 empty] | 0.41 | 0.41 | 20 | 0 | 15.6551 | 0.78275 | 280.1269 | 0.43449 | 26.0695 | 0 | 0 | 0.85247 | 0.55813 | 8.2 | 20
[6 empty] | 0.42 | 0.42 | 20 | 0 | 15.742 | 0.7871 | 284.4718 | 0.4258 | 25.5481 | 0 | 0 | 0.83542 | 0.56649 | 8.4 | 20
[6 empty] | 0.43 | 0.43 | 20 | 0 | 15.8271 | 0.79136 | 288.7298 | 0.41729 | 25.0371 | 0 | 0 | 0.81871 | 0.57468 | 8.6 | 20
[6 empty] | 0.44 | 0.44 | 20 | 0 | 15.9106 | 0.79553 | 292.9027 | 0.40894 | 24.5364 | 0 | 0 | 0.80234 | 0.5827 | 8.8 | 20
[6 empty] | 0.45 | 0.45 | 20 | 0 | 15.9924 | 0.79962 | 296.9921 | 0.40076 | 24.0456 | 0 | 0 | 0.78629 | 0.59056 | 9 | 20
[6 empty] | 0.46 | 0.46 | 20 | 0 | 16.0725 | 0.80363 | 300.9997 | 0.39275 | 23.5647 | 0 | 0 | 0.77057 | 0.59827 | 9.2 | 20
[6 empty] | 0.47 | 0.47 | 20 | 0 | 16.1511 | 0.80755 | 304.9271 | 0.38489 | 23.0934 | 0 | 0 | 0.75516 | 0.60582 | 9.4 | 20
[6 empty] | 0.48 | 0.48 | 20 | 0 | 16.2281 | 0.8114 | 308.776 | 0.37719 | 22.6316 | 0 | 0 | 0.74005 | 0.61322 | 9.6 | 20
[6 empty] | 0.49 | 0.49 | 20 | 0 | 16.3035 | 0.81518 | 312.548 | 0.36965 | 22.1789 | 0 | 0 | 0.72525 | 0.62047 | 9.8 | 20
[6 empty] | 0.5 | 0.5 | 20 | 0 | 16.3774 | 0.81887 | 316.2445 | 0.36226 | 21.7354 | 0 | 0 | 0.71075 | 0.62758 | 10 | 20
[6 empty] | 0.51 | 0.51 | 20 | 0 | 16.4499 | 0.82249 | 319.867 | 0.35501 | 21.3006 | 0 | 0 | 0.69653 | 0.63454 | 10.2 | 20
[6 empty] | 0.52 | 0.52 | 20 | 0 | 16.5209 | 0.82604 | 323.4171 | 0.34791 | 20.8746 | 0 | 0 | 0.6826 | 0.64137 | 10.4 | 20
[6 empty] | 0.53 | 0.53 | 20 | 0 | 16.5905 | 0.82952 | 326.8962 | 0.34095 | 20.4571 | 0 | 0 | 0.66895 | 0.64806 | 10.6 | 20
[6 empty] | 0.54 | 0.54 | 20 | 0 | 16.6587 | 0.83293 | 330.3058 | 0.33413 | 20.048 | 0 | 0 | 0.65557 | 0.65462 | 10.8 | 20
[6 empty] | 0.55 | 0.55 | 20 | 0 | 16.7255 | 0.83627 | 333.6471 | 0.32745 | 19.647 | 0 | 0 | 0.64246 | 0.66104 | 11 | 20
[6 empty] | 0.56 | 0.56 | 20 | 0 | 16.791 | 0.83955 | 336.9216 | 0.3209 | 19.2541 | 0 | 0 | 0.62961 | 0.66734 | 11.2 | 20
[6 empty] | 0.57 | 0.57 | 20 | 0 | 16.8552 | 0.84276 | 340.1306 | 0.31448 | 18.869 | 0 | 0 | 0.61702 | 0.67351 | 11.4 | 20
[6 empty] | 0.58 | 0.58 | 20 | 0 | 16.9181 | 0.8459 | 343.2755 | 0.30819 | 18.4916 | 0 | 0 | 0.60468 | 0.67955 | 11.6 | 20
[6 empty] | 0.59 | 0.59 | 20 | 0 | 16.9797 | 0.84898 | 346.3574 | 0.30203 | 18.1218 | 0 | 0 | 0.59258 | 0.68548 | 11.8 | 20
[6 empty] | 0.6 | 0.6 | 20 | 0 | 17.0401 | 0.85201 | 349.3777 | 0.29599 | 17.7594 | 0 | 0 | 0.58073 | 0.69129 | 12 | 20
[6 empty] | 0.61 | 0.61 | 20 | 0 | 17.0993 | 0.85497 | 352.3376 | 0.29007 | 17.4042 | 0 | 0 | 0.56912 | 0.69698 | 12.2 | 20
[6 empty] | 0.62 | 0.62 | 20 | 0 | 17.1573 | 0.85787 | 355.2383 | 0.28427 | 17.0561 | 0 | 0 | 0.55773 | 0.70255 | 12.4 | 20
[6 empty] | 0.63 | 0.63 | 20 | 0 | 17.2142 | 0.86071 | 358.081 | 0.27858 | 16.715 | 0 | 0 | 0.54658 | 0.70802 | 12.6 | 20
[6 empty] | 0.64 | 0.64 | 20 | 0 | 17.2699 | 0.86349 | 360.8668 | 0.27301 | 16.3807 | 0 | 0 | 0.53565 | 0.71338 | 12.8 | 20
[6 empty] | 0.65 | 0.65 | 20 | 0 | 17.3245 | 0.86622 | 363.5969 | 0.26755 | 16.0531 | 0 | 0 | 0.52494 | 0.71863 | 13 | 20
[6 empty] | 0.66 | 0.66 | 20 | 0 | 17.378 | 0.8689 | 366.2724 | 0.2622 | 15.732 | 0 | 0 | 0.51444 | 0.72377 | 13.2 | 20
[6 empty] | 0.67 | 0.67 | 20 | 0 | 17.4304 | 0.87152 | 368.8944 | 0.25696 | 15.4174 | 0 | 0 | 0.50415 | 0.72881 | 13.4 | 20
[6 empty] | 0.68 | 0.68 | 20 | 0 | 17.4818 | 0.87409 | 371.464 | 0.25182 | 15.109 | 0 | 0 | 0.49406 | 0.73375 | 13.6 | 20
[6 empty] | 0.69 | 0.69 | 20 | 0 | 17.5322 | 0.87661 | 373.9821 | 0.24678 | 14.8068 | 0 | 0 | 0.48418 | 0.73859 | 13.8 | 20
[6 empty] | 0.7 | 0.7 | 20 | 0 | 17.5816 | 0.87908 | 376.45 | 0.24184 | 14.5107 | 0 | 0 | 0.4745 | 0.74334 | 14 | 20
[6 empty] | 0.71 | 0.71 | 20 | 0 | 17.6299 | 0.8815 | 378.8684 | 0.23701 | 14.2205 | 0 | 0 | 0.46501 | 0.74799 | 14.2 | 20
[6 empty] | 0.72 | 0.72 | 20 | 0 | 17.6773 | 0.88387 | 381.2385 | 0.23227 | 13.9361 | 0 | 0 | 0.45571 | 0.75255 | 14.4 | 20
[6 empty] | 0.73 | 0.73 | 20 | 0 | 17.7238 | 0.88619 | 383.5612 | 0.22762 | 13.6574 | 0 | 0 | 0.4466 | 0.75701 | 14.6 | 20
[6 empty] | 0.74 | 0.74 | 20 | 0 | 17.7693 | 0.88846 | 385.8374 | 0.22307 | 13.3842 | 0 | 0 | 0.43766 | 0.76139 | 14.8 | 20
[6 empty] | 0.75 | 0.75 | 20 | 0 | 17.8139 | 0.8907 | 388.0681 | 0.21861 | 13.1165 | 0 | 0 | 0.42891 | 0.76568 | 15 | 20
[6 empty] | 0.76 | 0.76 | 20 | 0 | 17.8576 | 0.89288 | 390.2542 | 0.21424 | 12.8542 | 0 | 0 | 0.42033 | 0.76988 | 15.2 | 20
[6 empty] | 0.77 | 0.77 | 20 | 0 | 17.9005 | 0.89502 | 392.3965 | 0.20995 | 12.5971 | 0 | 0 | 0.41193 | 0.774 | 15.4 | 20
[6 empty] | 0.78 | 0.78 | 20 | 0 | 17.9425 | 0.89712 | 394.4961 | 0.20575 | 12.3452 | 0 | 0 | 0.40369 | 0.77804 | 15.6 | 20
[6 empty] | 0.79 | 0.79 | 20 | 0 | 17.9836 | 0.89918 | 396.5536 | 0.20164 | 12.0983 | 0 | 0 | 0.39561 | 0.78199 | 15.8 | 20
[6 empty] | 0.8 | 0.8 | 20 | 0 | 18.024 | 0.9012 | 398.57 | 0.1976 | 11.8563 | 0 | 0 | 0.3877 | 0.78587 | 16 | 20
[6 empty] | 0.81 | 0.81 | 20 | 0 | 18.0635 | 0.90317 | 400.546 | 0.19365 | 11.6192 | 0 | 0 | 0.37995 | 0.78967 | 16.2 | 20
[6 empty] | 0.82 | 0.82 | 20 | 0 | 18.1022 | 0.90511 | 402.4825 | 0.18978 | 11.3868 | 0 | 0 | 0.37235 | 0.79339 | 16.4 | 20
[6 empty] | 0.83 | 0.83 | 20 | 0 | 18.1402 | 0.90701 | 404.3803 | 0.18598 | 11.1591 | 0 | 0 | 0.3649 | 0.79704 | 16.6 | 20
[6 empty] | 0.84 | 0.84 | 20 | 0 | 18.1774 | 0.90887 | 406.2402 | 0.18226 | 10.9359 | 0 | 0 | 0.3576 | 0.80062 | 16.8 | 20
[6 empty] | 0.85 | 0.85 | 20 | 0 | 18.2138 | 0.91069 | 408.0628 | 0.17862 | 10.7172 | 0 | 0 | 0.35045 | 0.80412 | 17 | 20
[6 empty] | 0.86 | 0.86 | 20 | 0 | 18.2495 | 0.91248 | 409.849 | 0.17505 | 10.5028 | 0 | 0 | 0.34344 | 0.80756 | 17.2 | 20
[6 empty] | 0.87 | 0.87 | 20 | 0 | 18.2845 | 0.91423 | 411.5995 | 0.17155 | 10.2928 | 0 | 0 | 0.33657 | 0.81092 | 17.4 | 20
[6 empty] | 0.88 | 0.88 | 20 | 0 | 18.3189 | 0.91594 | 413.3149 | 0.16811 | 10.0869 | 0 | 0 | 0.32984 | 0.81422 | 17.6 | 20
[6 empty] | 0.89 | 0.89 | 20 | 0 | 18.3525 | 0.91762 | 414.9961 | 0.16475 | 9.8852 | 0 | 0 | 0.32324 | 0.81745 | 17.8 | 20
[6 empty] | 0.9 | 0.9 | 20 | 0 | 18.3854 | 0.91927 | 416.6436 | 0.16146 | 9.6875 | 0 | 0 | 0.31678 | 0.82062 | 18 | 20
[6 empty] | 0.91 | 0.91 | 20 | 0 | 18.4177 | 0.92089 | 418.2582 | 0.15823 | 9.4937 | 0 | 0 | 0.31044 | 0.82373 | 18.2 | 20
[6 empty] | 0.92 | 0.92 | 20 | 0 | 18.4494 | 0.92247 | 419.8405 | 0.15506 | 9.3038 | 0 | 0 | 0.30424 | 0.82677 | 18.4 | 20
[6 empty] | 0.93 | 0.93 | 20 | 0 | 18.4804 | 0.92402 | 421.3911 | 0.15196 | 9.1178 | 0 | 0 | 0.29815 | 0.82975 | 18.6 | 20
[6 empty] | 0.94 | 0.94 | 20 | 0 | 18.5108 | 0.92554 | 422.9107 | 0.14892 | 8.9354 | 0 | 0 | 0.29219 | 0.83267 | 18.8 | 20
[6 empty] | 0.95 | 0.95 | 20 | 0 | 18.5406 | 0.92703 | 424.4 | 0.14594 | 8.7567 | 0 | 0 | 0.28634 | 0.83554 | 19 | 20
[6 empty] | 0.96 | 0.96 | 20 | 0 | 18.5697 | 0.92849 | 425.8594 | 0.14303 | 8.5816 | 0 | 0 | 0.28062 | 0.83834 | 19.2 | 20
[6 empty] | 0.97 | 0.97 | 20 | 0 | 18.5983 | 0.92992 | 427.2897 | 0.14017 | 8.4099 | 0 | 0 | 0.275 | 0.84109 | 19.4 | 20
[6 empty] | 0.98 | 0.98 | 20 | 0 | 18.6264 | 0.93132 | 428.6913 | 0.13736 | 8.2417 | 0 | 0 | 0.2695 | 0.84379 | 19.6 | 20
[6 empty] | 0.99 | 0.99 | 20 | 0 | 18.6539 | 0.93269 | 430.065 | 0.13461 | 8.0769 | 0 | 0 | 0.26411 | 0.84643 | 19.8 | 20
[6 empty] | 1 | 1 | 20 | 0 | 18.6808 | 0.93404 | 431.4111 | 0.13192 | 7.9154 | 0 | 0 | 0.25883 | 0.84902 | 20 | 20
[6 empty] | 1.01 | 1.01 | 10 | 0 | 18.7072 | 0.93536 | 432.7303 | -0.87072 | -52.2429 | 0 | 0 | -0.85417 | 0.84048 | 20.1 | 19.901
[6 empty] | 1.02 | 1.02 | 10 | 0 | 18.533 | 0.92665 | 424.0232 | -0.8533 | -51.1981 | 0 | 0 | -0.83709 | 0.8321 | 20.2 | 19.8039
[6 empty] | 1.03 | 1.03 | 10 | 0 | 18.3624 | 0.91812 | 415.4902 | -0.83624 | -50.1741 | 0 | 0 | -0.82035 | 0.8239 | 20.3 | 19.7087
[6 empty] | 1.04 | 1.04 | 10 | 0 | 18.1951 | 0.90976 | 407.1278 | -0.81951 | -49.1706 | 0 | 0 | -0.80394 | 0.81586 | 20.4 | 19.6154
[6 empty] | 1.05 | 1.05 | 10 | 0 | 18.0312 | 0.90156 | 398.9327 | -0.80312 | -48.1872 | 0 | 0 | -0.78786 | 0.80798 | 20.5 | 19.5238
[6 empty] | 1.06 | 1.06 | 10 | 0 | 17.8706 | 0.89353 | 390.9015 | -0.78706 | -47.2235 | 0 | 0 | -0.7721 | 0.80026 | 20.6 | 19.434
[6 empty] | 1.07 | 1.07 | 10 | 0 | 17.7132 | 0.88566 | 383.0309 | -0.77132 | -46.279 | 0 | 0 | -0.75666 | 0.7927 | 20.7 | 19.3458
[6 empty] | 1.08 | 1.08 | 10 | 0 | 17.5589 | 0.87795 | 375.3177 | -0.75589 | -45.3534 | 0 | 0 | -0.74153 | 0.78528 | 20.8 | 19.2593
[6 empty] | 1.09 | 1.09 | 10 | 0 | 17.4077 | 0.87039 | 367.7588 | -0.74077 | -44.4464 | 0 | 0 | -0.7267 | 0.77801 | 20.9 | 19.1743
[6 empty] | 1.1 | 1.1 | 10 | 0 | 17.2596 | 0.86298 | 360.3511 | -0.72596 | -43.5574 | 0 | 0 | -0.71216 | 0.77089 | 21 | 19.0909
[6 empty] | 1.11 | 1.11 | 10 | 0 | 17.1144 | 0.85572 | 353.0915 | -0.71144 | -42.6863 | 0 | 0 | -0.69792 | 0.76391 | 21.1 | 19.009
[6 empty] | 1.12 | 1.12 | 10 | 0 | 16.9721 | 0.8486 | 345.9772 | -0.69721 | -41.8326 | 0 | 0 | -0.68396 | 0.75707 | 21.2 | 18.9286
[6 empty] | 1.13 | 1.13 | 10 | 0 | 16.8327 | 0.84163 | 339.0051 | -0.68327 | -40.9959 | 0 | 0 | -0.67028 | 0.75037 | 21.3 | 18.8496
[6 empty] | 1.14 | 1.14 | 10 | 0 | 16.696 | 0.8348 | 332.1724 | -0.6696 | -40.176 | 0 | 0 | -0.65688 | 0.7438 | 21.4 | 18.7719
[6 empty] | 1.15 | 1.15 | 10 | 0 | 16.5621 | 0.8281 | 325.4764 | -0.65621 | -39.3725 | 0 | 0 | -0.64374 | 0.73736 | 21.5 | 18.6957
[6 empty] | 1.16 | 1.16 | 10 | 0 | 16.4308 | 0.82154 | 318.9143 | -0.64308 | -38.585 | 0 | 0 | -0.63087 | 0.73105 | 21.6 | 18.6207
[6 empty] | 1.17 | 1.17 | 10 | 0 | 16.3022 | 0.81511 | 312.4835 | -0.63022 | -37.8133 | 0 | 0 | -0.61825 | 0.72487 | 21.7 | 18.547
[6 empty] | 1.18 | 1.18 | 10 | 0 | 16.1762 | 0.80881 | 306.1813 | -0.61762 | -37.0571 | 0 | 0 | -0.60588 | 0.71881 | 21.8 | 18.4746
[6 empty] | 1.19 | 1.19 | 10 | 0 | 16.0527 | 0.80263 | 300.0051 | -0.60527 | -36.3159 | 0 | 0 | -0.59377 | 0.71288 | 21.9 | 18.4034
[6 empty] | 1.2 | 1.2 | 10 | 0 | 15.9316 | 0.79658 | 293.9524 | -0.59316 | -35.5896 | 0 | 0 | -0.58189 | 0.70706 | 22 | 18.3333
[6 empty] | 1.21 | 1.21 | 10 | 0 | 15.813 | 0.79065 | 288.0208 | -0.5813 | -34.8778 | 0 | 0 | -0.57025 | 0.70135 | 22.1 | 18.2645
[6 empty] | 1.22 | 1.22 | 10 | 0 | 15.6967 | 0.78484 | 282.2079 | -0.56967 | -34.1803 | 0 | 0 | -0.55885 | 0.69577 | 22.2 | 18.1967
[6 empty] | 1.23 | 1.23 | 10 | 0 | 15.5828 | 0.77914 | 276.5112 | -0.55828 | -33.4966 | 0 | 0 | -0.54767 | 0.69029 | 22.3 | 18.1301
[6 empty] | 1.24 | 1.24 | 10 | 0 | 15.4711 | 0.77356 | 270.9284 | -0.54711 | -32.8267 | 0 | 0 | -0.53672 | 0.68492 | 22.4 | 18.0645
[6 empty] | 1.25 | 1.25 | 10 | 0 | 15.3617 | 0.76808 | 265.4573 | -0.53617 | -32.1702 | 0 | 0 | -0.52598 | 0.67966 | 22.5 | 18
[6 empty] | 1.26 | 1.26 | 10 | 0 | 15.2545 | 0.76272 | 260.0956 | -0.52545 | -31.5268 | 0 | 0 | -0.51546 | 0.67451 | 22.6 | 17.9365
[6 empty] | 1.27 | 1.27 | 10 | 0 | 15.1494 | 0.75747 | 254.8411 | -0.51494 | -30.8962 | 0 | 0 | -0.50515 | 0.66946 | 22.7 | 17.874
[6 empty] | 1.28 | 1.28 | 10 | 0 | 15.0464 | 0.75232 | 249.6917 | -0.50464 | -30.2783 | 0 | 0 | -0.49505 | 0.66451 | 22.8 | 17.8125
[6 empty] | 1.29 | 1.29 | 10 | 0 | 14.9455 | 0.74727 | 244.6453 | -0.49455 | -29.6728 | 0 | 0 | -0.48515 | 0.65965 | 22.9 | 17.7519
[6 empty] | 1.3 | 1.3 | 10 | 0 | 14.8465 | 0.74233 | 239.6999 | -0.48465 | -29.0793 | 0 | 0 | -0.47545 | 0.6549 | 23 | 17.6923
[6 empty] | 1.31 | 1.31 | 10 | 0 | 14.7496 | 0.73748 | 234.8533 | -0.47496 | -28.4977 | 0 | 0 | -0.46594 | 0.65024 | 23.1 | 17.6336
[6 empty] | 1.32 | 1.32 | 10 | 0 | 14.6546 | 0.73273 | 230.1037 | -0.46546 | -27.9278 | 0 | 0 | -0.45662 | 0.64567 | 23.2 | 17.5758
[6 empty] | 1.33 | 1.33 | 10 | 0 | 14.5615 | 0.72808 | 225.4491 | -0.45615 | -27.3692 | 0 | 0 | -0.44749 | 0.6412 | 23.3 | 17.5188
[6 empty] | 1.34 | 1.34 | 10 | 0 | 14.4703 | 0.72352 | 220.8876 | -0.44703 | -26.8218 | 0 | 0 | -0.43854 | 0.63681 | 23.4 | 17.4627
[6 empty] | 1.35 | 1.35 | 10 | 0 | 14.3809 | 0.71904 | 216.4173 | -0.43809 | -26.2854 | 0 | 0 | -0.42977 | 0.63252 | 23.5 | 17.4074
[6 empty] | 1.36 | 1.36 | 10 | 0 | 14.2933 | 0.71466 | 212.0364 | -0.42933 | -25.7597 | 0 | 0 | -0.42117 | 0.6283 | 23.6 | 17.3529
[6 empty] | 1.37 | 1.37 | 10 | 0 | 14.2074 | 0.71037 | 207.7431 | -0.42074 | -25.2445 | 0 | 0 | -0.41275 | 0.62418 | 23.7 | 17.2993
[6 empty] | 1.38 | 1.38 | 10 | 0 | 14.1233 | 0.70616 | 203.5357 | -0.41233 | -24.7396 | 0 | 0 | -0.40449 | 0.62013 | 23.8 | 17.2464
[6 empty] | 1.39 | 1.39 | 10 | 0 | 14.0408 | 0.70204 | 199.4124 | -0.40408 | -24.2448 | 0 | 0 | -0.3964 | 0.61617 | 23.9 | 17.1942
[6 empty] | 1.4 | 1.4 | 10 | 0 | 13.96 | 0.698 | 195.3716 | -0.396 | -23.7599 | 0 | 0 | -0.38847 | 0.61228 | 24 | 17.1429
[6 empty] | 1.41 | 1.41 | 10 | 0 | 13.8808 | 0.69404 | 191.4116 | -0.38808 | -23.2847 | 0 | 0 | -0.3807 | 0.60848 | 24.1 | 17.0922
[6 empty] | 1.42 | 1.42 | 10 | 0 | 13.8032 | 0.69016 | 187.5308 | -0.38032 | -22.819 | 0 | 0 | -0.37309 | 0.60475 | 24.2 | 17.0423
[6 empty] | 1.43 | 1.43 | 10 | 0 | 13.7271 | 0.68636 | 183.7277 | -0.37271 | -22.3626 | 0 | 0 | -0.36563 | 0.60109 | 24.3 | 16.993
[6 empty] | 1.44 | 1.44 | 10 | 0 | 13.6526 | 0.68263 | 180.0006 | -0.36526 | -21.9154 | 0 | 0 | -0.35832 | 0.59751 | 24.4 | 16.9444
[6 empty] | 1.45 | 1.45 | 10 | 0 | 13.5795 | 0.67898 | 176.348 | -0.35795 | -21.4771 | 0 | 0 | -0.35115 | 0.59399 | 24.5 | 16.8966
[6 empty] | 1.46 | 1.46 | 10 | 0 | 13.5079 | 0.6754 | 172.7685 | -0.35079 | -21.0475 | 0 | 0 | -0.34413 | 0.59055 | 24.6 | 16.8493
[6 empty] | 1.47 | 1.47 | 10 | 0 | 13.4378 | 0.67189 | 169.2606 | -0.34378 | -20.6266 | 0 | 0 | -0.33724 | 0.58718 | 24.7 | 16.8027
[6 empty] | 1.48 | 1.48 | 10 | 0 | 13.369 | 0.66845 | 165.8228 | -0.3369 | -20.214 | 0 | 0 | -0.3305 | 0.58388 | 24.8 | 16.7568
[6 empty] | 1.49 | 1.49 | 10 | 0 | 13.3016 | 0.66508 | 162.4538 | -0.33016 | -19.8098 | 0 | 0 | -0.32389 | 0.58064 | 24.9 | 16.7114
[6 empty] | 1.5 | 1.5 | 10 | 0 | 13.2356 | 0.66178 | 159.1522 | -0.32356 | -19.4136 | 0 | 0 | -0.31741 | 0.57746 | 25 | 16.6667
[6 empty] | 1.51 | 1.51 | 10 | 0 | 13.1709 | 0.65854 | 155.9166 | -0.31709 | -19.0253 | 0 | 0 | -0.31106 | 0.57435 | 25.1 | 16.6225
[6 empty] | 1.52 | 1.52 | 10 | 0 | 13.1075 | 0.65537 | 152.7457 | -0.31075 | -18.6448 | 0 | 0 | -0.30484 | 0.5713 | 25.2 | 16.5789
[6 empty] | 1.53 | 1.53 | 10 | 0 | 13.0453 | 0.65227 | 149.6382 | -0.30453 | -18.2719 | 0 | 0 | -0.29875 | 0.56832 | 25.3 | 16.5359
[6 empty] | 1.54 | 1.54 | 10 | 0 | 12.9844 | 0.64922 | 146.5929 | -0.29844 | -17.9065 | 0 | 0 | -0.29277 | 0.56539 | 25.4 | 16.4935
[6 empty] | 1.55 | 1.55 | 10 | 0 | 12.9247 | 0.64624 | 143.6085 | -0.29247 | -17.5483 | 0 | 0 | -0.28692 | 0.56252 | 25.5 | 16.4516
[6 empty] | 1.56 | 1.56 | 10 | 0 | 12.8662 | 0.64331 | 140.6838 | -0.28662 | -17.1974 | 0 | 0 | -0.28118 | 0.55971 | 25.6 | 16.4103
[6 empty] | 1.57 | 1.57 | 10 | 0 | 12.8089 | 0.64045 | 137.8176 | -0.28089 | -16.8534 | 0 | 0 | -0.27555 | 0.55695 | 25.7 | 16.3694
[6 empty] | 1.58 | 1.58 | 10 | 0 | 12.7527 | 0.63764 | 135.0086 | -0.27527 | -16.5163 | 0 | 0 | -0.27004 | 0.55425 | 25.8 | 16.3291
[6 empty] | 1.59 | 1.59 | 10 | 0 | 12.6977 | 0.63488 | 132.2559 | -0.26977 | -16.186 | 0 | 0 | -0.26464 | 0.55161 | 25.9 | 16.2893
[6 empty] | 1.6 | 1.6 | 10 | 0 | 12.6437 | 0.63219 | 129.5583 | -0.26437 | -15.8623 | 0 | 0 | -0.25935 | 0.54901 | 26 | 16.25
[6 empty] | 1.61 | 1.61 | 10 | 0 | 12.5908 | 0.62954 | 126.9145 | -0.25908 | -15.5451 | 0 | 0 | -0.25416 | 0.54647 | 26.1 | 16.2112
[6 empty] | 1.62 | 1.62 | 10 | 0 | 12.539 | 0.62695 | 124.3237 | -0.2539 | -15.2342 | 0 | 0 | -0.24908 | 0.54398 | 26.2 | 16.1728
[6 empty] | 1.63 | 1.63 | 10 | 0 | 12.4882 | 0.62441 | 121.7847 | -0.24882 | -14.9295 | 0 | 0 | -0.2441 | 0.54154 | 26.3 | 16.135
[6 empty] | 1.64 | 1.64 | 10 | 0 | 12.4385 | 0.62192 | 119.2964 | -0.24385 | -14.6309 | 0 | 0 | -0.23921 | 0.53915 | 26.4 | 16.0976
[6 empty] | 1.65 | 1.65 | 10 | 0 | 12.3897 | 0.61949 | 116.8579 | -0.23897 | -14.3383 | 0 | 0 | -0.23443 | 0.5368 | 26.5 | 16.0606
[6 empty] | 1.66 | 1.66 | 10 | 0 | 12.3419 | 0.6171 | 114.4682 | -0.23419 | -14.0515 | 0 | 0 | -0.22974 | 0.5345 | 26.6 | 16.0241
[6 empty] | 1.67 | 1.67 | 10 | 0 | 12.2951 | 0.61475 | 112.1263 | -0.22951 | -13.7705 | 0 | 0 | -0.22515 | 0.53225 | 26.7 | 15.988
[6 empty] | 1.68 | 1.68 | 10 | 0 | 12.2492 | 0.61246 | 109.8312 | -0.22492 | -13.4951 | 0 | 0 | -0.22064 | 0.53005 | 26.8 | 15.9524
[6 empty] | 1.69 | 1.69 | 10 | 0 | 12.2042 | 0.61021 | 107.5821 | -0.22042 | -13.2252 | 0 | 0 | -0.21623 | 0.52788 | 26.9 | 15.9172
[6 empty] | 1.7 | 1.7 | 10 | 0 | 12.1601 | 0.60801 | 105.3779 | -0.21601 | -12.9607 | 0 | 0 | -0.21191 | 0.52577 | 27 | 15.8824
[6 empty] | 1.71 | 1.71 | 10 | 0 | 12.1169 | 0.60585 | 103.2178 | -0.21169 | -12.7014 | 0 | 0 | -0.20767 | 0.52369 | 27.1 | 15.848
[6 empty] | 1.72 | 1.72 | 10 | 0 | 12.0746 | 0.60373 | 101.1009 | -0.20746 | -12.4474 | 0 | 0 | -0.20352 | 0.52165 | 27.2 | 15.814
[6 empty] | 1.73 | 1.73 | 10 | 0 | 12.0331 | 0.60165 | 99.0263 | -0.20331 | -12.1985 | 0 | 0 | -0.19944 | 0.51966 | 27.3 | 15.7803
[6 empty] | 1.74 | 1.74 | 10 | 0 | 11.9924 | 0.59962 | 96.9932 | -0.19924 | -11.9545 | 0 | 0 | -0.19546 | 0.5177 | 27.4 | 15.7471
[6 empty] | 1.75 | 1.75 | 10 | 0 | 11.9526 | 0.59763 | 95.0008 | -0.19526 | -11.7154 | 0 | 0 | -0.19155 | 0.51579 | 27.5 | 15.7143
[6 empty] | 1.76 | 1.76 | 10 | 0 | 11.9135 | 0.59568 | 93.0482 | -0.19135 | -11.4811 | 0 | 0 | -0.18772 | 0.51391 | 27.6 | 15.6818
[6 empty] | 1.77 | 1.77 | 10 | 0 | 11.8752 | 0.59376 | 91.1347 | -0.18752 | -11.2515 | 0 | 0 | -0.18396 | 0.51207 | 27.7 | 15.6497
[6 empty] | 1.78 | 1.78 | 10 | 0 | 11.8377 | 0.59189 | 89.2595 | -0.18377 | -11.0264 | 0 | 0 | -0.18028 | 0.51027 | 27.8 | 15.618
[6 empty] | 1.79 | 1.79 | 10 | 0 | 11.801 | 0.59005 | 87.4217 | -0.1801 | -10.8059 | 0 | 0 | -0.17668 | 0.5085 | 27.9 | 15.5866
[6 empty] | 1.8 | 1.8 | 10 | 0 | 11.765 | 0.58825 | 85.6207 | -0.1765 | -10.5898 | 0 | 0 | -0.17314 | 0.50677 | 28 | 15.5556
[6 empty] | 1.81 | 1.81 | 10 | 0 | 11.7297 | 0.58648 | 83.8558 | -0.17297 | -10.378 | 0 | 0 | -0.16968 | 0.50507 | 28.1 | 15.5249
[6 empty] | 1.82 | 1.82 | 10 | 0 | 11.6951 | 0.58475 | 82.1261 | -0.16951 | -10.1704 | 0 | 0 | -0.16629 | 0.50341 | 28.2 | 15.4945
[6 empty] | 1.83 | 1.83 | 10 | 0 | 11.6612 | 0.58306 | 80.431 | -0.16612 | -9.967 | 0 | 0 | -0.16296 | 0.50178 | 28.3 | 15.4645
[6 empty] | 1.84 | 1.84 | 10 | 0 | 11.6279 | 0.5814 | 78.7699 | -0.16279 | -9.7677 | 0 | 0 | -0.1597 | 0.50018 | 28.4 | 15.4348
[6 empty] | 1.85 | 1.85 | 10 | 0 | 11.5954 | 0.57977 | 77.1419 | -0.15954 | -9.5723 | 0 | 0 | -0.15651 | 0.49862 | 28.5 | 15.4054
[6 empty] | 1.86 | 1.86 | 10 | 0 | 11.5635 | 0.57817 | 75.5465 | -0.15635 | -9.3809 | 0 | 0 | -0.15338 | 0.49709 | 28.6 | 15.3763
[6 empty] | 1.87 | 1.87 | 10 | 0 | 11.5322 | 0.57661 | 73.983 | -0.15322 | -9.1933 | 0 | 0 | -0.15031 | 0.49558 | 28.7 | 15.3476
[6 empty] | 1.88 | 1.88 | 10 | 0 | 11.5016 | 0.57508 | 72.4508 | -0.15016 | -9.0094 | 0 | 0 | -0.1473 | 0.49411 | 28.8 | 15.3191
[6 empty] | 1.89 | 1.89 | 10 | 0 | 11.4715 | 0.57358 | 70.9493 | -0.14715 | -8.8292 | 0 | 0 | -0.14436 | 0.49267 | 28.9 | 15.291
[6 empty] | 1.9 | 1.9 | 10 | 0 | 11.4421 | 0.57211 | 69.4777 | -0.14421 | -8.6526 | 0 | 0 | -0.14147 | 0.49125 | 29 | 15.2632
[6 empty] | 1.91 | 1.91 | 10 | 0 | 11.4133 | 0.57066 | 68.0356 | -0.14133 | -8.4796 | 0 | 0 | -0.13864 | 0.48986 | 29.1 | 15.2356
[6 empty] | 1.92 | 1.92 | 10 | 0 | 11.385 | 0.56925 | 66.6224 | -0.1385 | -8.31 | 0 | 0 | -0.13587 | 0.48851 | 29.2 | 15.2083
[6 empty] | 1.93 | 1.93 | 10 | 0 | 11.3573 | 0.56786 | 65.2374 | -0.13573 | -8.1438 | 0 | 0 | -0.13315 | 0.48717 | 29.3 | 15.1813
[6 empty] | 1.94 | 1.94 | 10 | 0 | 11.3302 | 0.56651 | 63.8801 | -0.13302 | -7.9809 | 0 | 0 | -0.13049 | 0.48587 | 29.4 | 15.1546
[6 empty] | 1.95 | 1.95 | 10 | 0 | 11.3035 | 0.56518 | 62.5499 | -0.13035 | -7.8213 | 0 | 0 | -0.12788 | 0.48459 | 29.5 | 15.1282
[6 empty] | 1.96 | 1.96 | 10 | 0 | 11.2775 | 0.56387 | 61.2464 | -0.12775 | -7.6649 | 0 | 0 | -0.12532 | 0.48334 | 29.6 | 15.102
[6 empty] | 1.97 | 1.97 | 10 | 0 | 11.2519 | 0.5626 | 59.9689 | -0.12519 | -7.5116 | 0 | 0 | -0.12281 | 0.48211 | 29.7 | 15.0761
[6 empty] | 1.98 | 1.98 | 10 | 0 | 11.2269 | 0.56134 | 58.7169 | -0.12269 | -7.3613 | 0 | 0 | -0.12036 | 0.48091 | 29.8 | 15.0505
[6 empty] | 1.99 | 1.99 | 10 | 0 | 11.2024 | 0.56012 | 57.4901 | -0.12024 | -7.2141 | 0 | 0 | -0.11795 | 0.47973 | 29.9 | 15.0251
[6 empty] | 2 | 2 | 10 | 0 | 11.1783 | 0.55892 | 56.2877 | -0.11783 | -7.0698 | 0 | 0 | -0.11559 | 0.47857 | 30 | 15
[6 empty] | 2.01 | 2.01 | 10 | 0 | 11.1547 | 0.55774 | 55.1094 | -0.11547 | -6.9284 | 0 | 0 | -0.11328 | 0.47744 | 30.1 | 14.9751
[6 empty] | 2.02 | 2.02 | 10 | 0 | 11.1316 | 0.55658 | 53.9547 | -0.11316 | -6.7899 | 0 | 0 | -0.11101 | 0.47633 | 30.2 | 14.9505
[6 empty] | 2.03 | 2.03 | 10 | 0 | 11.109 | 0.55545 | 52.823 | -0.1109 | -6.6541 | 0 | 0 | -0.10879 | 0.47524 | 30.3 | 14.9261
[6 empty] | 2.04 | 2.04 | 10 | 0 | 11.0868 | 0.55434 | 51.714 | -0.10868 | -6.521 | 0 | 0 | -0.10662 | 0.47417 | 30.4 | 14.902
[6 empty] | 2.05 | 2.05 | 10 | 0 | 11.0651 | 0.55325 | 50.6272 | -0.10651 | -6.3906 | 0 | 0 | -0.10449 | 0.47313 | 30.5 | 14.878
[6 empty] | 2.06 | 2.06 | 10 | 0 | 11.0438 | 0.55219 | 49.5621 | -0.10438 | -6.2628 | 0 | 0 | -0.1024 | 0.4721 | 30.6 | 14.8544
[6 empty] | 2.07 | 2.07 | 10 | 0 | 11.0229 | 0.55115 | 48.5183 | -0.10229 | -6.1375 | 0 | 0 | -0.10035 | 0.4711 | 30.7 | 14.8309
[6 empty] | 2.08 | 2.08 | 10 | 0 | 11.0025 | 0.55012 | 47.4954 | -0.10025 | -6.0148 | 0 | 0 | -0.098341 | 0.47012 | 30.8 | 14.8077
[6 empty] | 2.09 | 2.09 | 10 | 0 | 10.9824 | 0.54912 | 46.4929 | -0.098241 | -5.8945 | 0 | 0 | -0.096374 | 0.46915 | 30.9 | 14.7847
[6 empty] | 2.1 | 2.1 | 10 | 0 | 10.9628 | 0.54814 | 45.5105 | -0.096276 | -5.7766 | 0 | 0 | -0.094447 | 0.46821 | 31 | 14.7619
[6 empty] | 2.11 | 2.11 | 10 | 0 | 10.9435 | 0.54718 | 44.5477 | -0.094351 | -5.661 | 0 | 0 | -0.092558 | 0.46728 | 31.1 | 14.7393
[6 empty] | 2.12 | 2.12 | 10 | 0 | 10.9246 | 0.54623 | 43.6042 | -0.092464 | -5.5478 | 0 | 0 | -0.090707 | 0.46638 | 31.2 | 14.717
[6 empty] | 2.13 | 2.13 | 10 | 0 | 10.9061 | 0.54531 | 42.6796 | -0.090614 | -5.4369 | 0 | 0 | -0.088893 | 0.46549 | 31.3 | 14.6948
[6 empty] | 2.14 | 2.14 | 10 | 0 | 10.888 | 0.5444 | 41.7734 | -0.088802 | -5.3281 | 0 | 0 | -0.087115 | 0.46462 | 31.4 | 14.6729
[6 empty] | 2.15 | 2.15 | 10 | 0 | 10.8703 | 0.54351 | 40.8854 | -0.087026 | -5.2216 | 0 | 0 | -0.085373 | 0.46376 | 31.5 | 14.6512
[6 empty] | 2.16 | 2.16 | 10 | 0 | 10.8529 | 0.54264 | 40.0152 | -0.085285 | -5.1171 | 0 | 0 | -0.083665 | 0.46293 | 31.6 | 14.6296
[6 empty] | 2.17 | 2.17 | 10 | 0 | 10.8358 | 0.54179 | 39.1623 | -0.08358 | -5.0148 | 0 | 0 | -0.081992 | 0.46211 | 31.7 | 14.6083
[6 empty] | 2.18 | 2.18 | 10 | 0 | 10.8191 | 0.54095 | 38.3265 | -0.081908 | -4.9145 | 0 | 0 | -0.080352 | 0.4613 | 31.8 | 14.5872
[6 empty] | 2.19 | 2.19 | 10 | 0 | 10.8027 | 0.54014 | 37.5074 | -0.08027 | -4.8162 | 0 | 0 | -0.078745 | 0.46052 | 31.9 | 14.5662
[6 empty] | 2.2 | 2.2 | 10 | 0 | 10.7866 | 0.53933 | 36.7047 | -0.078665 | -4.7199 | 0 | 0 | -0.07717 | 0.45974 | 32 | 14.5455
[6 empty] | 2.21 | 2.21 | 10 | 0 | 10.7709 | 0.53855 | 35.9181 | -0.077091 | -4.6255 | 0 | 0 | -0.075627 | 0.45899 | 32.1 | 14.5249
[6 empty] | 2.22 | 2.22 | 10 | 0 | 10.7555 | 0.53777 | 35.1472 | -0.07555 | -4.533 | 0 | 0 | -0.074114 | 0.45825 | 32.2 | 14.5045
[6 empty] | 2.23 | 2.23 | 10 | 0 | 10.7404 | 0.53702 | 34.3917 | -0.074039 | -4.4423 | 0 | 0 | -0.072632 | 0.45752 | 32.3 | 14.4843
[6 empty] | 2.24 | 2.24 | 10 | 0 | 10.7256 | 0.53628 | 33.6513 | -0.072558 | -4.3535 | 0 | 0 | -0.071179 | 0.45681 | 32.4 | 14.4643
[6 empty] | 2.25 | 2.25 | 10 | 0 | 10.7111 | 0.53555 | 32.9257 | -0.071107 | -4.2664 | 0 | 0 | -0.069756 | 0.45611 | 32.5 | 14.4444
[6 empty] | 2.26 | 2.26 | 10 | 0 | 10.6968 | 0.53484 | 32.2146 | -0.069684 | -4.1811 | 0 | 0 | -0.06836 | 0.45543 | 32.6 | 14.4248
[6 empty] | 2.27 | 2.27 | 10 | 0 | 10.6829 | 0.53415 | 31.5178 | -0.068291 | -4.0974 | 0 | 0 | -0.066993 | 0.45476 | 32.7 | 14.4053
[6 empty] | 2.28 | 2.28 | 10 | 0 | 10.6692 | 0.53346 | 30.8349 | -0.066925 | -4.0155 | 0 | 0 | -0.065653 | 0.4541 | 32.8 | 14.386
[6 empty] | 2.29 | 2.29 | 10 | 0 | 10.6559 | 0.53279 | 30.1656 | -0.065586 | -3.9352 | 0 | 0 | -0.06434 | 0.45346 | 32.9 | 14.3668
[6 empty] | 2.3 | 2.3 | 10 | 0 | 10.6427 | 0.53214 | 29.5098 | -0.064275 | -3.8565 | 0 | 0 | -0.063054 | 0.45283 | 33 | 14.3478
[6 empty] | 2.31 | 2.31 | 10 | 0 | 10.6299 | 0.53149 | 28.867 | -0.062989 | -3.7794 | 0 | 0 | -0.061792 | 0.45221 | 33.1 | 14.329
[6 empty] | 2.32 | 2.32 | 10 | 0 | 10.6173 | 0.53086 | 28.2371 | -0.061729 | -3.7038 | 0 | 0 | -0.060557 | 0.4516 | 33.2 | 14.3103
[6 empty] | 2.33 | 2.33 | 10 | 0 | 10.6049 | 0.53025 | 27.6198 | -0.060495 | -3.6297 | 0 | 0 | -0.059345 | 0.45101 | 33.3 | 14.2918
[6 empty] | 2.34 | 2.34 | 10 | 0 | 10.5928 | 0.52964 | 27.0149 | -0.059285 | -3.5571 | 0 | 0 | -0.058159 | 0.45043 | 33.4 | 14.2735
[6 empty] | 2.35 | 2.35 | 10 | 0 | 10.581 | 0.52905 | 26.4221 | -0.058099 | -3.486 | 0 | 0 | -0.056995 | 0.44986 | 33.5 | 14.2553
[6 empty] | 2.36 | 2.36 | 10 | 0 | 10.5694 | 0.52847 | 25.8411 | -0.056937 | -3.4162 | 0 | 0 | -0.055855 | 0.4493 | 33.6 | 14.2373
[6 empty] | 2.37 | 2.37 | 10 | 0 | 10.558 | 0.5279 | 25.2717 | -0.055799 | -3.3479 | 0 | 0 | -0.054738 | 0.44875 | 33.7 | 14.2194
[6 empty] | 2.38 | 2.38 | 10 | 0 | 10.5468 | 0.52734 | 24.7137 | -0.054683 | -3.281 | 0 | 0 | -0.053644 | 0.44822 | 33.8 | 14.2017
[6 empty] | 2.39 | 2.39 | 10 | 0 | 10.5359 | 0.52679 | 24.1669 | -0.053589 | -3.2153 | 0 | 0 | -0.052571 | 0.44769 | 33.9 | 14.1841
[6 empty] | 2.4 | 2.4 | 10 | 0 | 10.5252 | 0.52626 | 23.631 | -0.052517 | -3.151 | 0 | 0 | -0.051519 | 0.44718 | 34 | 14.1667
[6 empty] | 2.41 | 2.41 | 10 | 0 | 10.5147 | 0.52573 | 23.1058 | -0.051467 | -3.088 | 0 | 0 | -0.050489 | 0.44667 | 34.1 | 14.1494
[6 empty] | 2.42 | 2.42 | 10 | 0 | 10.5044 | 0.52522 | 22.5911 | -0.050437 | -3.0262 | 0 | 0 | -0.049479 | 0.44618 | 34.2 | 14.1322
[6 empty] | 2.43 | 2.43 | 10 | 0 | 10.4943 | 0.52471 | 22.0868 | -0.049429 | -2.9657 | 0 | 0 | -0.04849 | 0.44569 | 34.3 | 14.1152
[6 empty] | 2.44 | 2.44 | 10 | 0 | 10.4844 | 0.52422 | 21.5925 | -0.04844 | -2.9064 | 0 | 0 | -0.04752 | 0.44522 | 34.4 | 14.0984
[6 empty] | 2.45 | 2.45 | 10 | 0 | 10.4747 | 0.52374 | 21.1081 | -0.047471 | -2.8483 | 0 | 0 | -0.046569 | 0.44475 | 34.5 | 14.0816
[6 empty] | 2.46 | 2.46 | 10 | 0 | 10.4652 | 0.52326 | 20.6334 | -0.046522 | -2.7913 | 0 | 0 | -0.045638 | 0.44429 | 34.6 | 14.065
[6 empty] | 2.47 | 2.47 | 10 | 0 | 10.4559 | 0.5228 | 20.1682 | -0.045591 | -2.7355 | 0 | 0 | -0.044725 | 0.44385 | 34.7 | 14.0486
[6 empty] | 2.48 | 2.48 | 10 | 0 | 10.4468 | 0.52234 | 19.7122 | -0.04468 | -2.6808 | 0 | 0 | -0.043831 | 0.44341 | 34.8 | 14.0323
[6 empty] | 2.49 | 2.49 | 10 | 0 | 10.4379 | 0.52189 | 19.2654 | -0.043786 | -2.6272 | 0 | 0 | -0.042954 | 0.44298 | 34.9 | 14.0161
[6 empty] | 2.5 | 2.5 | 10 | 0 | 10.4291 | 0.52146 | 18.8276 | -0.04291 | -2.5746 | 0 | 0 | -0.042095 | 0.44256 | 35 | 14
[6 empty] | 2.51 | 2.51 | 10 | 0 | 10.4205 | 0.52103 | 18.3985 | -0.042052 | -2.5231 | 0 | 0 | -0.041253 | 0.44214 | 35.1 | 13.9841
[6 empty] | 2.52 | 2.52 | 10 | 0 | 10.4121 | 0.52061 | 17.978 | -0.041211 | -2.4727 | 0 | 0 | -0.040428 | 0.44174 | 35.2 | 13.9683
[6 empty] | 2.53 | 2.53 | 10 | 0 | 10.4039 | 0.52019 | 17.5658 | -0.040387 | -2.4232 | 0 | 0 | -0.039619 | 0.44134 | 35.3 | 13.9526
[6 empty] | 2.54 | 2.54 | 10 | 0 | 10.3958 | 0.51979 | 17.162 | -0.039579 | -2.3747 | 0 | 0 | -0.038827 | 0.44096 | 35.4 | 13.937
[6 empty] | 2.55 | 2.55 | 10 | 0 | 10.3879 | 0.51939 | 16.7662 | -0.038788 | -2.3273 | 0 | 0 | -0.038051 | 0.44058 | 35.5 | 13.9216
[6 empty] | 2.56 | 2.56 | 10 | 0 | 10.3801 | 0.51901 | 16.3783 | -0.038012 | -2.2807 | 0 | 0 | -0.03729 | 0.4402 | 35.6 | 13.9063
[6 empty] | 2.57 | 2.57 | 10 | 0 | 10.3725 | 0.51863 | 15.9982 | -0.037252 | -2.2351 | 0 | 0 | -0.036544 | 0.43984 | 35.7 | 13.8911
[6 empty] | 2.58 | 2.58 | 10 | 0 | 10.3651 | 0.51825 | 15.6257 | -0.036507 | -2.1904 | 0 | 0 | -0.035813 | 0.43948 | 35.8 | 13.876
[6 empty] | 2.59 | 2.59 | 10 | 0 | 10.3578 | 0.51789 | 15.2606 | -0.035776 | -2.1466 | 0 | 0 | -0.035097 | 0.43913 | 35.9 | 13.861
[6 empty] | 2.6 | 2.6 | 10 | 0 | 10.3506 | 0.51753 | 14.9028 | -0.035061 | -2.1037 | 0 | 0 | -0.034395 | 0.43878 | 36 | 13.8462
[6 empty] | 2.61 | 2.61 | 10 | 0 | 10.3436 | 0.51718 | 14.5522 | -0.03436 | -2.0616 | 0 | 0 | -0.033707 | 0.43845 | 36.1 | 13.8314
[6 empty] | 2.62 | 2.62 | 10 | 0 | 10.3367 | 0.51684 | 14.2086 | -0.033672 | -2.0203 | 0 | 0 | -0.033033 | 0.43812 | 36.2 | 13.8168
[6 empty] | 2.63 | 2.63 | 10 | 0 | 10.33 | 0.5165 | 13.8719 | -0.032999 | -1.9799 | 0 | 0 | -0.032372 | 0.43779 | 36.3 | 13.8023
[6 empty] | 2.64 | 2.64 | 10 | 0 | 10.3234 | 0.51617 | 13.5419 | -0.032339 | -1.9403 | 0 | 0 | -0.031725 | 0.43748 | 36.4 | 13.7879
[6 empty] | 2.65 | 2.65 | 10 | 0 | 10.3169 | 0.51585 | 13.2185 | -0.031692 | -1.9015 | 0 | 0 | -0.03109 | 0.43716 | 36.5 | 13.7736
[6 empty] | 2.66 | 2.66 | 10 | 0 | 10.3106 | 0.51553 | 12.9016 | -0.031058 | -1.8635 | 0 | 0 | -0.030468 | 0.43686 | 36.6 | 13.7594
[6 empty] | 2.67 | 2.67 | 10 | 0 | 10.3044 | 0.51522 | 12.591 | -0.030437 | -1.8262 | 0 | 0 | -0.029859 | 0.43656 | 36.7 | 13.7453
[6 empty] | 2.68 | 2.68 | 10 | 0 | 10.2983 | 0.51491 | 12.2867 | -0.029828 | -1.7897 | 0 | 0 | -0.029262 | 0.43627 | 36.8 | 13.7313
[6 empty] | 2.69 | 2.69 | 10 | 0 | 10.2923 | 0.51462 | 11.9884 | -0.029232 | -1.7539 | 0 | 0 | -0.028677 | 0.43598 | 36.9 | 13.7175
[6 empty] | 2.7 | 2.7 | 10 | 0 | 10.2865 | 0.51432 | 11.6961 | -0.028647 | -1.7188 | 0 | 0 | -0.028103 | 0.4357 | 37 | 13.7037
[6 empty] | 2.71 | 2.71 | 10 | 0 | 10.2807 | 0.51404 | 11.4096 | -0.028074 | -1.6845 | 0 | 0 | -0.027541 | 0.43543 | 37.1 | 13.69
[6 empty] | 2.72 | 2.72 | 10 | 0 | 10.2751 | 0.51376 | 11.1288 | -0.027513 | -1.6508 | 0 | 0 | -0.02699 | 0.43516 | 37.2 | 13.6765
[6 empty] | 2.73 | 2.73 | 10 | 0 | 10.2696 | 0.51348 | 10.8537 | -0.026963 | -1.6178 | 0 | 0 | -0.02645 | 0.43489 | 37.3 | 13.663
[6 empty] | 2.74 | 2.74 | 10 | 0 | 10.2642 | 0.51321 | 10.5841 | -0.026423 | -1.5854 | 0 | 0 | -0.025921 | 0.43463 | 37.4 | 13.6496
[6 empty] | 2.75 | 2.75 | 10 | 0 | 10.2589 | 0.51295 | 10.3199 | -0.025895 | -1.5537 | 0 | 0 | -0.025403 | 0.43438 | 37.5 | 13.6364
[6 empty] | 2.76 | 2.76 | 10 | 0 | 10.2538 | 0.51269 | 10.0609 | -0.025377 | -1.5226 | 0 | 0 | -0.024895 | 0.43413 | 37.6 | 13.6232
[6 empty] | 2.77 | 2.77 | 10 | 0 | 10.2487 | 0.51243 | 9.8071 | -0.024869 | -1.4922 | 0 | 0 | -0.024397 | 0.43389 | 37.7 | 13.6101
[6 empty] | 2.78 | 2.78 | 10 | 0 | 10.2437 | 0.51219 | 9.5584 | -0.024372 | -1.4623 | 0 | 0 | -0.023909 | 0.43365 | 37.8 | 13.5971
[6 empty] | 2.79 | 2.79 | 10 | 0 | 10.2388 | 0.51194 | 9.3147 | -0.023885 | -1.4331 | 0 | 0 | -0.023431 | 0.43341 | 37.9 | 13.5842
[6 empty] | 2.8 | 2.8 | 10 | 0 | 10.2341 | 0.5117 | 9.0759 | -0.023407 | -1.4044 | 0 | 0 | -0.022962 | 0.43318 | 38 | 13.5714
[6 empty] | 2.81 | 2.81 | 10 | 0 | 10.2294 | 0.51147 | 8.8418 | -0.022939 | -1.3763 | 0 | 0 | -0.022503 | 0.43296 | 38.1 | 13.5587
[6 empty] | 2.82 | 2.82 | 10 | 0 | 10.2248 | 0.51124 | 8.6124 | -0.02248 | -1.3488 | 0 | 0 | -0.022053 | 0.43274 | 38.2 | 13.5461
[6 empty] | 2.83 | 2.83 | 10 | 0 | 10.2203 | 0.51102 | 8.3876 | -0.02203 | -1.3218 | 0 | 0 | -0.021612 | 0.43252 | 38.3 | 13.5336
[6 empty] | 2.84 | 2.84 | 10 | 0 | 10.2159 | 0.51079 | 8.1673 | -0.02159 | -1.2954 | 0 | 0 | -0.02118 | 0.43231 | 38.4 | 13.5211
[6 empty] | 2.85 | 2.85 | 10 | 0 | 10.2116 | 0.51058 | 7.9514 | -0.021158 | -1.2695 | 0 | 0 | -0.020756 | 0.4321 | 38.5 | 13.5088
[6 empty] | 2.86 | 2.86 | 10 | 0 | 10.2073 | 0.51037 | 7.7398 | -0.020735 | -1.2441 | 0 | 0 | -0.020341 | 0.4319 | 38.6 | 13.4965
[6 empty] | 2.87 | 2.87 | 10 | 0 | 10.2032 | 0.51016 | 7.5325 | -0.02032 | -1.2192 | 0 | 0 | -0.019934 | 0.4317 | 38.7 | 13.4843
[6 empty] | 2.88 | 2.88 | 10 | 0 | 10.1991 | 0.50996 | 7.3293 | -0.019914 | -1.1948 | 0 | 0 | -0.019535 | 0.4315 | 38.8 | 13.4722
[6 empty] | 2.89 | 2.89 | 10 | 0 | 10.1952 | 0.50976 | 7.1301 | -0.019515 | -1.1709 | 0 | 0 | -0.019145 | 0.43131 | 38.9 | 13.4602
[6 empty] | 2.9 | 2.9 | 10 | 0 | 10.1913 | 0.50956 | 6.935 | -0.019125 | -1.1475 | 0 | 0 | -0.018762 | 0.43112 | 39 | 13.4483
[6 empty] | 2.91 | 2.91 | 10 | 0 | 10.1874 | 0.50937 | 6.7437 | -0.018743 | -1.1246 | 0 | 0 | -0.018387 | 0.43094 | 39.1 | 13.4364
[6 empty] | 2.92 | 2.92 | 10 | 0 | 10.1837 | 0.50918 | 6.5563 | -0.018368 | -1.1021 | 0 | 0 | -0.018019 | 0.43076 | 39.2 | 13.4247
[6 empty] | 2.93 | 2.93 | 10 | 0 | 10.18 | 0.509 | 6.3726 | -0.018 | -1.08 | 0 | 0 | -0.017658 | 0.43058 | 39.3 | 13.413
[6 empty] | 2.94 | 2.94 | 10 | 0 | 10.1764 | 0.50882 | 6.1926 | -0.01764 | -1.0584 | 0 | 0 | -0.017305 | 0.43041 | 39.4 | 13.4014
[6 empty] | 2.95 | 2.95 | 10 | 0 | 10.1729 | 0.50864 | 6.0162 | -0.017288 | -1.0373 | 0 | 0 | -0.016959 | 0.43024 | 39.5 | 13.3898
[6 empty] | 2.96 | 2.96 | 10 | 0 | 10.1694 | 0.50847 | 5.8434 | -0.016942 | -1.0165 | 0 | 0 | -0.01662 | 0.43007 | 39.6 | 13.3784
[6 empty] | 2.97 | 2.97 | 10 | 0 | 10.166 | 0.5083 | 5.6739 | -0.016603 | -0.99618 | 0 | 0 | -0.016288 | 0.42991 | 39.7 | 13.367
[6 empty] | 2.98 | 2.98 | 10 | 0 | 10.1627 | 0.50814 | 5.5079 | -0.016271 | -0.97626 | 0 | 0 | -0.015962 | 0.42975 | 39.8 | 13.3557
[6 empty] | 2.99 | 2.99 | 10 | 0 | 10.1595 | 0.50797 | 5.3452 | -0.015946 | -0.95673 | 0 | 0 | -0.015643 | 0.4296 | 39.9 | 13.3445
[6 empty] | 3 | 3 | 10 | 0 | 10.1563 | 0.50781 | 5.1857 | -0.015627 | -0.9376 | 0 | 0 | -0.01533 | 0.42944 | 40 | 13.3333
[6 empty] | 3.01 | 3.01 | 10 | 0 | 10.1531 | 0.50766 | 5.0295 | -0.015314 | -0.91885 | 0 | 0 | -0.015023 | 0.42929 | 40.1 | 13.3223
[6 empty] | 3.02 | 3.02 | 10 | 0 | 10.1501 | 0.5075 | 4.8763 | -0.015008 | -0.90047 | 0 | 0 | -0.014723 | 0.42914 | 40.2 | 13.3113
[6 empty] | 3.03 | 3.03 | 10 | 0 | 10.1471 | 0.50735 | 4.7263 | -0.014708 | -0.88246 | 0 | 0 | -0.014428 | 0.429 | 40.3 | 13.3003
[6 empty] | 3.04 | 3.04 | 10 | 0 | 10.1441 | 0.50721 | 4.5792 | -0.014414 | -0.86481 | 0 | 0 | -0.01414 | 0.42886 | 40.4 | 13.2895
[6 empty] | 3.05 | 3.05 | 10 | 0 | 10.1413 | 0.50706 | 4.435 | -0.014125 | -0.84751 | 0 | 0 | -0.013857 | 0.42872 | 40.5 | 13.2787
[6 empty] | 3.06 | 3.06 | 10 | 0 | 10.1384 | 0.50692 | 4.2938 | -0.013843 | -0.83056 | 0 | 0 | -0.01358 | 0.42858 | 40.6 | 13.268
[6 empty] | 3.07 | 3.07 | 10 | 0 | 10.1357 | 0.50678 | 4.1554 | -0.013566 | -0.81395 | 0 | 0 | -0.013308 | 0.42845 | 40.7 | 13.2573
[6 empty] | 3.08 | 3.08 | 10 | 0 | 10.1329 | 0.50665 | 4.0197 | -0.013295 | -0.79767 | 0 | 0 | -0.013042 | 0.42832 | 40.8 | 13.2468
[6 empty] | 3.09 | 3.09 | 10 | 0 | 10.1303 | 0.50651 | 3.8868 | -0.013029 | -0.78172 | 0 | 0 | -0.012781 | 0.42819 | 40.9 | 13.2362
[6 empty] | 3.1 | 3.1 | 10 | 0 | 10.1277 | 0.50638 | 3.7565 | -0.012768 | -0.76609 | 0 | 0 | -0.012526 | 0.42807 | 41 | 13.2258
[6 empty] | 3.11 | 3.11 | 10 | 0 | 10.1251 | 0.50626 | 3.6288 | -0.012513 | -0.75076 | 0 | 0 | -0.012275 | 0.42795 | 41.1 | 13.2154
[6 empty] | 3.12 | 3.12 | 10 | 0 | 10.1226 | 0.50613 | 3.5037 | -0.012262 | -0.73575 | 0 | 0 | -0.012029 | 0.42783 | 41.2 | 13.2051
[6 empty] | 3.13 | 3.13 | 10 | 0 | 10.1202 | 0.50601 | 3.381 | -0.012017 | -0.72103 | 0 | 0 | -0.011789 | 0.42771 | 41.3 | 13.1949
[6 empty] | 3.14 | 3.14 | 10 | 0 | 10.1178 | 0.50589 | 3.2609 | -0.011777 | -0.70661 | 0 | 0 | -0.011553 | 0.42759 | 41.4 | 13.1847
[6 empty] | 3.15 | 3.15 | 10 | 0 | 10.1154 | 0.50577 | 3.1431 | -0.011541 | -0.69248 | 0 | 0 | -0.011322 | 0.42748 | 41.5 | 13.1746
[6 empty] | 3.16 | 3.16 | 10 | 0 | 10.1131 | 0.50566 | 3.0277 | -0.011311 | -0.67863 | 0 | 0 | -0.011096 | 0.42737 | 41.6 | 13.1646
[6 empty] | 3.17 | 3.17 | 10 | 0 | 10.1108 | 0.50554 | 2.9146 | -0.011084 | -0.66506 | 0 | 0 | -0.010874 | 0.42726 | 41.7 | 13.1546
[6 empty] | 3.18 | 3.18 | 10 | 0 | 10.1086 | 0.50543 | 2.8037 | -0.010863 | -0.65176 | 0 | 0 | -0.010656 | 0.42715 | 41.8 | 13.1447
[6 empty] | 3.19 | 3.19 | 10 | 0 | 10.1065 | 0.50532 | 2.6951 | -0.010645 | -0.63872 | 0 | 0 | -0.010443 | 0.42705 | 41.9 | 13.1348
[6 empty] | 3.2 | 3.2 | 10 | 0 | 10.1043 | 0.50522 | 2.5887 | -0.010432 | -0.62595 | 0 | 0 | -0.010234 | 0.42695 | 42 | 13.125
[6 empty] | 3.21 | 3.21 | 10 | 0 | 10.1022 | 0.50511 | 2.4843 | -0.010224 | -0.61343 | 0 | 0 | -0.01003 | 0.42685 | 42.1 | 13.1153
[6 empty] | 3.22 | 3.22 | 10 | 0 | 10.1002 | 0.50501 | 2.3821 | -0.010019 | -0.60116 | 0 | 0 | -0.009829 | 0.42675 | 42.2 | 13.1056
[6 empty] | 3.23 | 3.23 | 10 | 0 | 10.0982 | 0.50491 | 2.2819 | -0.009819 | -0.58914 | 0 | 0 | -0.0096324 | 0.42665 | 42.3 | 13.096
[6 empty] | 3.24 | 3.24 | 10 | 0 | 10.0962 | 0.50481 | 2.1837 | -0.0096226 | -0.57735 | 0 | 0 | -0.0094397 | 0.42656 | 42.4 | 13.0864
[6 empty] | 3.25 | 3.25 | 10 | 0 | 10.0943 | 0.50472 | 2.0875 | -0.0094301 | -0.56581 | 0 | 0 | -0.009251 | 0.42646 | 42.5 | 13.0769
[6 empty] | 3.26 | 3.26 | 10 | 0 | 10.0924 | 0.50462 | 1.9932 | -0.0092415 | -0.55449 | 0 | 0 | -0.0090659 | 0.42637 | 42.6 | 13.0675
[6 empty] | 3.27 | 3.27 | 10 | 0 | 10.0906 | 0.50453 | 1.9008 | -0.0090567 | -0.5434 | 0 | 0 | -0.0088846 | 0.42628 | 42.7 | 13.0581
[6 empty] | 3.28 | 3.28 | 10 | 0 | 10.0888 | 0.50444 | 1.8102 | -0.0088756 | -0.53253 | 0 | 0 | -0.0087069 | 0.4262 | 42.8 | 13.0488
[6 empty] | 3.29 | 3.29 | 10 | 0 | 10.087 | 0.50435 | 1.7214 | -0.008698 | -0.52188 | 0 | 0 | -0.0085328 | 0.42611 | 42.9 | 13.0395
[6 empty] | 3.3 | 3.3 | 10 | 0 | 10.0852 | 0.50426 | 1.6345 | -0.0085241 | -0.51145 | 0 | 0 | -0.0083621 | 0.42603 | 43 | 13.0303
[6 empty] | 3.31 | 3.31 | 10 | 0 | 10.0835 | 0.50418 | 1.5492 | -0.0083536 | -0.50122 | 0 | 0 | -0.0081949 | 0.42595 | 43.1 | 13.0211
[6 empty] | 3.32 | 3.32 | 10 | 0 | 10.0819 | 0.50409 | 1.4657 | -0.0081865 | -0.49119 | 0 | 0 | -0.008031 | 0.42587 | 43.2 | 13.012
[6 empty] | 3.33 | 3.33 | 10 | 0 | 10.0802 | 0.50401 | 1.3838 | -0.0080228 | -0.48137 | 0 | 0 | -0.0078704 | 0.42579 | 43.3 | 13.003
[6 empty] | 3.34 | 3.34 | 10 | 0 | 10.0786 | 0.50393 | 1.3036 | -0.0078623 | -0.47174 | 0 | 0 | -0.007713 | 0.42571 | 43.4 | 12.994
[6 empty] | 3.35 | 3.35 | 10 | 0 | 10.0771 | 0.50385 | 1.225 | -0.0077051 | -0.46231 | 0 | 0 | -0.0075587 | 0.42563 | 43.5 | 12.9851
[6 empty] | 3.36 | 3.36 | 10 | 0 | 10.0755 | 0.50378 | 1.1479 | -0.007551 | -0.45306 | 0 | 0 | -0.0074075 | 0.42556 | 43.6 | 12.9762
[6 empty] | 3.37 | 3.37 | 10 | 0 | 10.074 | 0.5037 | 1.0724 | -0.0074 | -0.444 | 0 | 0 | -0.0072594 | 0.42549 | 43.7 | 12.9674
[6 empty] | 3.38 | 3.38 | 10 | 0 | 10.0725 | 0.50363 | 0.99841 | -0.007252 | -0.43512 | 0 | 0 | -0.0071142 | 0.42542 | 43.8 | 12.9586
[6 empty] | 3.39 | 3.39 | 10 | 0 | 10.0711 | 0.50355 | 0.92589 | -0.0071069 | -0.42642 | 0 | 0 | -0.0069719 | 0.42535 | 43.9 | 12.9499
[6 empty] | 3.4 | 3.4 | 10 | 0 | 10.0696 | 0.50348 | 0.85482 | -0.0069648 | -0.41789 | 0 | 0 | -0.0068325 | 0.42528 | 44 | 12.9412
[6 empty] | 3.41 | 3.41 | 10 | 0 | 10.0683 | 0.50341 | 0.78517 | -0.0068255 | -0.40953 | 0 | 0 | -0.0066958 | 0.42521 | 44.1 | 12.9326
[6 empty] | 3.42 | 3.42 | 10 | 0 | 10.0669 | 0.50334 | 0.71692 | -0.006689 | -0.40134 | 0 | 0 | -0.0065619 | 0.42515 | 44.2 | 12.924
[6 empty] | 3.43 | 3.43 | 10 | 0 | 10.0656 | 0.50328 | 0.65003 | -0.0065552 | -0.39331 | 0 | 0 | -0.0064307 | 0.42508 | 44.3 | 12.9155
[6 empty] | 3.44 | 3.44 | 10 | 0 | 10.0642 | 0.50321 | 0.58448 | -0.0064241 | -0.38545 | 0 | 0 | -0.006302 | 0.42502 | 44.4 | 12.907
[6 empty] | 3.45 | 3.45 | 10 | 0 | 10.063 | 0.50315 | 0.52023 | -0.0062956 | -0.37774 | 0 | 0 | -0.006176 | 0.42496 | 44.5 | 12.8986
[6 empty] | 3.46 | 3.46 | 10 | 0 | 10.0617 | 0.50308 | 0.45728 | -0.0061697 | -0.37018 | 0 | 0 | -0.0060525 | 0.4249 | 44.6 | 12.8902
[6 empty] | 3.47 | 3.47 | 10 | 0 | 10.0605 | 0.50302 | 0.39558 | -0.0060463 | -0.36278 | 0 | 0 | -0.0059314 | 0.42484 | 44.7 | 12.8818
[6 empty] | 3.48 | 3.48 | 10 | 0 | 10.0593 | 0.50296 | 0.33512 | -0.0059254 | -0.35552 | 0 | 0 | -0.0058128 | 0.42478 | 44.8 | 12.8736
[6 empty] | 3.49 | 3.49 | 10 | 0 | 10.0581 | 0.5029 | 0.27586 | -0.0058069 | -0.34841 | 0 | 0 | -0.0056966 | 0.42472 | 44.9 | 12.8653
[6 empty] | 3.5 | 3.5 | 10 | 0 | 10.0569 | 0.50285 | 0.2178 | -0.0056907 | -0.34144 | 0 | 0 | -0.0055826 | 0.42467 | 45 | 12.8571
[6 empty] | 3.51 | 3.51 | 10 | 0 | 10.0558 | 0.50279 | 0.16089 | -0.0055769 | -0.33462 | 0 | 0 | -0.005471 | 0.42461 | 45.1 | 12.849
[6 empty] | 3.52 | 3.52 | 10 | 0 | 10.0547 | 0.50273 | 0.10512 | -0.0054654 | -0.32792 | 0 | 0 | -0.0053616 | 0.42456 | 45.2 | 12.8409
[6 empty] | 3.53 | 3.53 | 20 | 0 | 10.0536 | 0.50268 | 0.050464 | 0.99464 | 59.6786 | 0 | 0 | 1.9515 | 0.44407 | 45.4 | 12.8612
[6 empty] | 3.54 | 3.54 | 20 | 0 | 10.2525 | 0.51262 | 9.9969 | 0.97475 | 58.4851 | 0 | 0 | 1.9125 | 0.4632 | 45.6 | 12.8814
[6 empty] | 3.55 | 3.55 | 20 | 0 | 10.4474 | 0.52237 | 19.7444 | 0.95526 | 57.3154 | 0 | 0 | 1.8742 | 0.48194 | 45.8 | 12.9014
[6 empty] | 3.56 | 3.56 | 20 | 0 | 10.6385 | 0.53192 | 29.297 | 0.93615 | 56.1691 | 0 | 0 | 1.8367 | 0.50031 | 46 | 12.9213
[6 empty] | 3.57 | 3.57 | 20 | 0 | 10.8257 | 0.54129 | 38.6585 | 0.91743 | 55.0457 | 0 | 0 | 1.8 | 0.51831 | 46.2 | 12.9412
[6 empty] | 3.58 | 3.58 | 20 | 0 | 11.0092 | 0.55046 | 47.8328 | 0.89908 | 53.9448 | 0 | 0 | 1.764 | 0.53595 | 46.4 | 12.9609
[6 empty] | 3.59 | 3.59 | 20 | 0 | 11.189 | 0.55945 | 56.8236 | 0.8811 | 52.8659 | 0 | 0 | 1.7287 | 0.55323 | 46.6 | 12.9805
[6 empty] | 3.6 | 3.6 | 20 | 0 | 11.3652 | 0.56826 | 65.6345 | 0.86348 | 51.8085 | 0 | 0 | 1.6941 | 0.57018 | 46.8 | 13
[6 empty] | 3.61 | 3.61 | 20 | 0 | 11.5379 | 0.5769 | 74.2693 | 0.84621 | 50.7724 | 0 | 0 | 1.6603 | 0.58678 | 47 | 13.0194
[6 empty] | 3.62 | 3.62 | 20 | 0 | 11.7072 | 0.58536 | 82.7314 | 0.82928 | 49.7569 | 0 | 0 | 1.6271 | 0.60305 | 47.2 | 13.0387
[6 empty] | 3.63 | 3.63 | 20 | 0 | 11.873 | 0.59365 | 91.0242 | 0.8127 | 48.7618 | 0 | 0 | 1.5945 | 0.61899 | 47.4 | 13.0579
[6 empty] | 3.64 | 3.64 | 20 | 0 | 12.0356 | 0.60178 | 99.1511 | 0.79644 | 47.7866 | 0 | 0 | 1.5626 | 0.63462 | 47.6 | 13.0769
[6 empty] | 3.65 | 3.65 | 20 | 0 | 12.1949 | 0.60974 | 107.1156 | 0.78051 | 46.8308 | 0 | 0 | 1.5314 | 0.64993 | 47.8 | 13.0959
[6 empty] | 3.66 | 3.66 | 20 | 0 | 12.351 | 0.61755 | 114.9207 | 0.7649 | 45.8942 | 0 | 0 | 1.5007 | 0.66494 | 48 | 13.1148
[6 empty] | 3.67 | 3.67 | 20 | 0 | 12.5039 | 0.6252 | 122.5697 | 0.74961 | 44.9763 | 0 | 0 | 1.4707 | 0.67965 | 48.2 | 13.1335
[6 empty] | 3.68 | 3.68 | 20 | 0 | 12.6539 | 0.63269 | 130.0658 | 0.73461 | 44.0768 | 0 | 0 | 1.4413 | 0.69406 | 48.4 | 13.1522
[6 empty] | 3.69 | 3.69 | 20 | 0 | 12.8008 | 0.64004 | 137.4119 | 0.71992 | 43.1953 | 0 | 0 | 1.4125 | 0.70819 | 48.6 | 13.1707
[6 empty] | 3.7 | 3.7 | 20 | 0 | 12.9448 | 0.64724 | 144.6111 | 0.70552 | 42.3314 | 0 | 0 | 1.3842 | 0.72203 | 48.8 | 13.1892
[6 empty] | 3.71 | 3.71 | 20 | 0 | 13.0859 | 0.65429 | 151.6664 | 0.69141 | 41.4847 | 0 | 0 | 1.3566 | 0.73559 | 49 | 13.2075
[6 empty] | 3.72 | 3.72 | 20 | 0 | 13.2242 | 0.66121 | 158.5805 | 0.67758 | 40.655 | 0 | 0 | 1.3294 | 0.74889 | 49.2 | 13.2258
[6 empty] | 3.73 | 3.73 | 20 | 0 | 13.3597 | 0.66798 | 165.3563 | 0.66403 | 39.8419 | 0 | 0 | 1.3028 | 0.76192 | 49.4 | 13.244
[6 empty] | 3.74 | 3.74 | 20 | 0 | 13.4925 | 0.67462 | 171.9966 | 0.65075 | 39.0451 | 0 | 0 | 1.2768 | 0.77468 | 49.6 | 13.262
[6 empty] | 3.75 | 3.75 | 20 | 0 | 13.6226 | 0.68113 | 178.5042 | 0.63774 | 38.2642 | 0 | 0 | 1.2512 | 0.7872 | 49.8 | 13.28
[6 empty] | 3.76 | 3.76 | 20 | 0 | 13.7502 | 0.68751 | 184.8815 | 0.62498 | 37.4989 | 0 | 0 | 1.2262 | 0.79946 | 50 | 13.2979
[6 empty] | 3.77 | 3.77 | 20 | 0 | 13.8752 | 0.69376 | 191.1313 | 0.61248 | 36.7489 | 0 | 0 | 1.2017 | 0.81148 | 50.2 | 13.3156
[6 empty] | 3.78 | 3.78 | 20 | 0 | 13.9977 | 0.69988 | 197.2562 | 0.60023 | 36.014 | 0 | 0 | 1.1777 | 0.82325 | 50.4 | 13.3333
[6 empty] | 3.79 | 3.79 | 20 | 0 | 14.1177 | 0.70589 | 203.2585 | 0.58823 | 35.2937 | 0 | 0 | 1.1541 | 0.83479 | 50.6 | 13.3509
[6 empty] | 3.8 | 3.8 | 20 | 0 | 14.2354 | 0.71177 | 209.1408 | 0.57646 | 34.5878 | 0 | 0 | 1.131 | 0.8461 | 50.8 | 13.3684
[6 empty] | 3.81 | 3.81 | 20 | 0 | 14.3507 | 0.71753 | 214.9054 | 0.56493 | 33.896 | 0 | 0 | 1.1084 | 0.85719 | 51 | 13.3858
[6 empty] | 3.82 | 3.82 | 20 | 0 | 14.4636 | 0.72318 | 220.5547 | 0.55364 | 33.2181 | 0 | 0 | 1.0862 | 0.86805 | 51.2 | 13.4031
[6 empty] | 3.83 | 3.83 | 20 | 0 | 14.5744 | 0.72872 | 226.0911 | 0.54256 | 32.5538 | 0 | 0 | 1.0645 | 0.87869 | 51.4 | 13.4204
[6 empty] | 3.84 | 3.84 | 20 | 0 | 14.6829 | 0.73414 | 231.5167 | 0.53171 | 31.9027 | 0 | 0 | 1.0432 | 0.88913 | 51.6 | 13.4375
[6 empty] | 3.85 | 3.85 | 20 | 0 | 14.7892 | 0.73946 | 236.8338 | 0.52108 | 31.2646 | 0 | 0 | 1.0224 | 0.89935 | 51.8 | 13.4545
[6 empty] | 3.86 | 3.86 | 20 | 0 | 14.8934 | 0.74467 | 242.0446 | 0.51066 | 30.6393 | 0 | 0 | 1.0019 | 0.90937 | 52 | 13.4715
[6 empty] | 3.87 | 3.87 | 20 | 0 | 14.9956 | 0.74978 | 247.1512 | 0.50044 | 30.0266 | 0 | 0 | 0.98187 | 0.91919 | 52.2 | 13.4884
[6 empty] | 3.88 | 3.88 | 20 | 0 | 15.0957 | 0.75478 | 252.1556 | 0.49043 | 29.426 | 0 | 0 | 0.96223 | 0.92881 | 52.4 | 13.5052
[6 empty] | 3.89 | 3.89 | 20 | 0 | 15.1938 | 0.75969 | 257.0599 | 0.48062 | 28.8375 | 0 | 0 | 0.94299 | 0.93824 | 52.6 | 13.5219
[6 empty] | 3.9 | 3.9 | 20 | 0 | 15.2899 | 0.76449 | 261.8662 | 0.47101 | 28.2607 | 0 | 0 | 0.92413 | 0.94748 | 52.8 | 13.5385
[6 empty] | 3.91 | 3.91 | 20 | 0 | 15.3841 | 0.7692 | 266.5763 | 0.46159 | 27.6955 | 0 | 0 | 0.90564 | 0.95654 | 53 | 13.555
[6 empty] | 3.92 | 3.92 | 20 | 0 | 15.4764 | 0.77382 | 271.1922 | 0.45236 | 27.1416 | 0 | 0 | 0.88753 | 0.96541 | 53.2 | 13.5714
[6 empty] | 3.93 | 3.93 | 20 | 0 | 15.5669 | 0.77834 | 275.7158 | 0.44331 | 26.5988 | 0 | 0 | 0.86978 | 0.97411 | 53.4 | 13.5878
[6 empty] | 3.94 | 3.94 | 20 | 0 | 15.6555 | 0.78278 | 280.149 | 0.43445 | 26.0668 | 0 | 0 | 0.85238 | 0.98264 | 53.6 | 13.6041
[6 empty] | 3.95 | 3.95 | 20 | 0 | 15.7424 | 0.78712 | 284.4934 | 0.42576 | 25.5455 | 0 | 0 | 0.83534 | 0.99099 | 53.8 | 13.6203
[6 empty] | 3.96 | 3.96 | 20 | 0 | 15.8276 | 0.79138 | 288.751 | 0.41724 | 25.0346 | 0 | 0 | 0.81863 | 0.99917 | 54 | 13.6364
[6 empty] | 3.97 | 3.97 | 20 | 0 | 15.911 | 0.79555 | 292.9234 | 0.4089 | 24.5339 | 0 | 0 | 0.80226 | 1.0072 | 54.2 | 13.6524
[6 empty] | 3.98 | 3.98 | 20 | 0 | 15.9928 | 0.79964 | 297.0124 | 0.40072 | 24.0432 | 0 | 0 | 0.78621 | 1.0151 | 54.4 | 13.6683
[6 empty] | 3.99 | 3.99 | 20 | 0 | 16.0729 | 0.80365 | 301.0196 | 0.39271 | 23.5623 | 0 | 0 | 0.77049 | 1.0228 | 54.6 | 13.6842
[6 empty] | 4 | 4 | 20 | 0 | 16.1515 | 0.80757 | 304.9467 | 0.38485 | 23.0911 | 0 | 0 | 0.75508 | 1.0303 | 54.8 | 13.7
[6 empty] | 4.01 | 4.01 | 20 | 0 | 16.2285 | 0.81142 | 308.7952 | 0.37715 | 22.6293 | 0 | 0 | 0.73998 | 1.0377 | 55 | 13.7157
[6 empty] | 4.02 | 4.02 | 20 | 0 | 16.3039 | 0.81519 | 312.5667 | 0.36961 | 22.1767 | 0 | 0 | 0.72518 | 1.045 | 55.2 | 13.7313
[6 empty] | 4.03 | 4.03 | 20 | 0 | 16.3778 | 0.81889 | 316.2628 | 0.36222 | 21.7331 | 0 | 0 | 0.71067 | 1.0521 | 55.4 | 13.7469
[6 empty] | 4.04 | 4.04 | 20 | 0 | 16.4503 | 0.82251 | 319.885 | 0.35497 | 21.2985 | 0 | 0 | 0.69646 | 1.059 | 55.6 | 13.7624
[6 empty] | 4.05 | 4.05 | 20 | 0 | 16.5212 | 0.82606 | 323.4348 | 0.34788 | 20.8725 | 0 | 0 | 0.68253 | 1.0659 | 55.8 | 13.7778
[6 empty] | 4.06 | 4.06 | 20 | 0 | 16.5908 | 0.82954 | 326.9135 | 0.34092 | 20.4551 | 0 | 0 | 0.66888 | 1.0726 | 56 | 13.7931
[6 empty] | 4.07 | 4.07 | 20 | 0 | 16.659 | 0.83295 | 330.3227 | 0.3341 | 20.046 | 0 | 0 | 0.6555 | 1.0791 | 56.2 | 13.8084
[6 empty] | 4.08 | 4.08 | 20 | 0 | 16.7258 | 0.83629 | 333.6637 | 0.32742 | 19.645 | 0 | 0 | 0.64239 | 1.0855 | 56.4 | 13.8235
[6 empty] | 4.09 | 4.09 | 20 | 0 | 16.7913 | 0.83957 | 336.9379 | 0.32087 | 19.2521 | 0 | 0 | 0.62955 | 1.0918 | 56.6 | 13.8386
[6 empty] | 4.1 | 4.1 | 20 | 0 | 16.8555 | 0.84277 | 340.1466 | 0.31445 | 18.8671 | 0 | 0 | 0.61695 | 1.098 | 56.8 | 13.8537
[6 empty] | 4.11 | 4.11 | 20 | 0 | 16.9184 | 0.84592 | 343.2911 | 0.30816 | 18.4898 | 0 | 0 | 0.60462 | 1.104 | 57 | 13.8686
[6 empty] | 4.12 | 4.12 | 20 | 0 | 16.98 | 0.849 | 346.3727 | 0.302 | 18.12 | 0 | 0 | 0.59252 | 1.11 | 57.2 | 13.8835
[6 empty] | 4.13 | 4.13 | 20 | 0 | 17.0404 | 0.85202 | 349.3927 | 0.29596 | 17.7576 | 0 | 0 | 0.58067 | 1.1158 | 57.4 | 13.8983
[6 empty] | 4.14 | 4.14 | 20 | 0 | 17.0996 | 0.85498 | 352.3523 | 0.29004 | 17.4024 | 0 | 0 | 0.56906 | 1.1215 | 57.6 | 13.913
[6 empty] | 4.15 | 4.15 | 20 | 0 | 17.1576 | 0.85788 | 355.2527 | 0.28424 | 17.0544 | 0 | 0 | 0.55768 | 1.127 | 57.8 | 13.9277
[6 empty] | 4.16 | 4.16 | 20 | 0 | 17.2145 | 0.86072 | 358.0951 | 0.27855 | 16.7133 | 0 | 0 | 0.54652 | 1.1325 | 58 | 13.9423
[6 empty] | 4.17 | 4.17 | 20 | 0 | 17.2702 | 0.86351 | 360.8806 | 0.27298 | 16.379 | 0 | 0 | 0.53559 | 1.1379 | 58.2 | 13.9568
[6 empty] | 4.18 | 4.18 | 20 | 0 | 17.3248 | 0.86624 | 363.6105 | 0.26752 | 16.0514 | 0 | 0 | 0.52488 | 1.1431 | 58.4 | 13.9713
[6 empty] | 4.19 | 4.19 | 20 | 0 | 17.3783 | 0.86891 | 366.2857 | 0.26217 | 15.7304 | 0 | 0 | 0.51438 | 1.1483 | 58.6 | 13.9857
[6 empty] | 4.2 | 4.2 | 20 | 0 | 17.4307 | 0.87154 | 368.9075 | 0.25693 | 15.4158 | 0 | 0 | 0.5041 | 1.1533 | 58.8 | 14
[6 empty] | 4.21 | 4.21 | 20 | 0 | 17.4821 | 0.8741 | 371.4768 | 0.25179 | 15.1075 | 0 | 0 | 0.49401 | 1.1582 | 59 | 14.0143
[6 empty] | 4.22 | 4.22 | 20 | 0 | 17.5324 | 0.87662 | 373.9947 | 0.24676 | 14.8053 | 0 | 0 | 0.48413 | 1.1631 | 59.2 | 14.0284
[6 empty] | 4.23 | 4.23 | 20 | 0 | 17.5818 | 0.87909 | 376.4622 | 0.24182 | 14.5092 | 0 | 0 | 0.47445 | 1.1678 | 59.4 | 14.0426
[6 empty] | 4.24 | 4.24 | 20 | 0 | 17.6302 | 0.88151 | 378.8804 | 0.23698 | 14.219 | 0 | 0 | 0.46496 | 1.1725 | 59.6 | 14.0566
[6 empty] | 4.25 | 4.25 | 20 | 0 | 17.6776 | 0.88388 | 381.2503 | 0.23224 | 13.9347 | 0 | 0 | 0.45566 | 1.177 | 59.8 | 14.0706
[6 empty] | 4.26 | 4.26 | 20 | 0 | 17.724 | 0.8862 | 383.5727 | 0.2276 | 13.656 | 0 | 0 | 0.44655 | 1.1815 | 60 | 14.0845
[6 empty] | 4.27 | 4.27 | 20 | 0 | 17.7695 | 0.88848 | 385.8487 | 0.22305 | 13.3828 | 0 | 0 | 0.43762 | 1.1859 | 60.2 | 14.0984
[6 empty] | 4.28 | 4.28 | 20 | 0 | 17.8141 | 0.89071 | 388.0792 | 0.21859 | 13.1152 | 0 | 0 | 0.42887 | 1.1902 | 60.4 | 14.1121
[6 empty] | 4.29 | 4.29 | 20 | 0 | 17.8579 | 0.89289 | 390.265 | 0.21421 | 12.8529 | 0 | 0 | 0.42029 | 1.1944 | 60.6 | 14.1259
[6 empty] | 4.3 | 4.3 | 20 | 0 | 17.9007 | 0.89503 | 392.4072 | 0.20993 | 12.5958 | 0 | 0 | 0.41188 | 1.1985 | 60.8 | 14.1395
[6 empty] | 4.31 | 4.31 | 20 | 0 | 17.9427 | 0.89713 | 394.5065 | 0.20573 | 12.3439 | 0 | 0 | 0.40365 | 1.2025 | 61 | 14.1531
[6 empty] | 4.32 | 4.32 | 20 | 0 | 17.9838 | 0.89919 | 396.5638 | 0.20162 | 12.097 | 0 | 0 | 0.39557 | 1.2065 | 61.2 | 14.1667
[6 empty] | 4.33 | 4.33 | 20 | 0 | 18.0242 | 0.90121 | 398.58 | 0.19758 | 11.8551 | 0 | 0 | 0.38766 | 1.2103 | 61.4 | 14.1801
[6 empty] | 4.34 | 4.34 | 20 | 0 | 18.0637 | 0.90318 | 400.5558 | 0.19363 | 11.618 | 0 | 0 | 0.37991 | 1.2141 | 61.6 | 14.1935
[6 empty] | 4.35 | 4.35 | 20 | 0 | 18.1024 | 0.90512 | 402.4922 | 0.18976 | 11.3856 | 0 | 0 | 0.37231 | 1.2179 | 61.8 | 14.2069
[6 empty] | 4.36 | 4.36 | 20 | 0 | 18.1403 | 0.90702 | 404.3898 | 0.18597 | 11.1579 | 0 | 0 | 0.36486 | 1.2215 | 62 | 14.2202
[6 empty] | 4.37 | 4.37 | 20 | 0 | 18.1775 | 0.90888 | 406.2494 | 0.18225 | 10.9348 | 0 | 0 | 0.35757 | 1.2251 | 62.2 | 14.2334
[6 empty] | 4.38 | 4.38 | 20 | 0 | 18.214 | 0.9107 | 408.0719 | 0.1786 | 10.7161 | 0 | 0 | 0.35042 | 1.2286 | 62.4 | 14.2466
[6 empty] | 4.39 | 4.39 | 20 | 0 | 18.2497 | 0.91249 | 409.8579 | 0.17503 | 10.5017 | 0 | 0 | 0.34341 | 1.232 | 62.6 | 14.2597
[6 empty] | 4.4 | 4.4 | 20 | 0 | 18.2847 | 0.91424 | 411.6082 | 0.17153 | 10.2917 | 0 | 0 | 0.33654 | 1.2354 | 62.8 | 14.2727
[6 empty] | 4.41 | 4.41 | 20 | 0 | 18.319 | 0.91595 | 413.3235 | 0.1681 | 10.0859 | 0 | 0 | 0.32981 | 1.2387 | 63 | 14.2857
[6 empty] | 4.42 | 4.42 | 20 | 0 | 18.3526 | 0.91763 | 415.0045 | 0.16474 | 9.8842 | 0 | 0 | 0.32321 | 1.2419 | 63.2 | 14.2986
[6 empty] | 4.43 | 4.43 | 20 | 0 | 18.3856 | 0.91928 | 416.6518 | 0.16144 | 9.6865 | 0 | 0 | 0.31675 | 1.2451 | 63.4 | 14.3115
[6 empty] | 4.44 | 4.44 | 20 | 0 | 18.4179 | 0.92089 | 418.2662 | 0.15821 | 9.4927 | 0 | 0 | 0.31041 | 1.2482 | 63.6 | 14.3243
[6 empty] | 4.45 | 4.45 | 20 | 0 | 18.4495 | 0.92248 | 419.8483 | 0.15505 | 9.3029 | 0 | 0 | 0.3042 | 1.2512 | 63.8 | 14.3371
[6 empty] | 4.46 | 4.46 | 20 | 0 | 18.4805 | 0.92403 | 421.3988 | 0.15195 | 9.1168 | 0 | 0 | 0.29812 | 1.2542 | 64 | 14.3498
[6 empty] | 4.47 | 4.47 | 20 | 0 | 18.5109 | 0.92555 | 422.9183 | 0.14891 | 8.9345 | 0 | 0 | 0.29216 | 1.2571 | 64.2 | 14.3624
[6 empty] | 4.48 | 4.48 | 20 | 0 | 18.5407 | 0.92703 | 424.4074 | 0.14593 | 8.7558 | 0 | 0 | 0.28631 | 1.26 | 64.4 | 14.375
[6 empty] | 4.49 | 4.49 | 20 | 0 | 18.5699 | 0.92849 | 425.8667 | 0.14301 | 8.5807 | 0 | 0 | 0.28059 | 1.2628 | 64.6 | 14.3875
[6 empty] | 4.5 | 4.5 | 20 | 0 | 18.5985 | 0.92992 | 427.2968 | 0.14015 | 8.4091 | 0 | 0 | 0.27498 | 1.2656 | 64.8 | 14.4
[6 empty] | 4.51 | 4.51 | 20 | 0 | 18.6265 | 0.93133 | 428.6983 | 0.13735 | 8.2409 | 0 | 0 | 0.26948 | 1.2683 | 65 | 14.4124
[6 empty] | 4.52 | 4.52 | 20 | 0 | 18.654 | 0.9327 | 430.0718 | 0.1346 | 8.0761 | 0 | 0 | 0.26409 | 1.2709 | 65.2 | 14.4248
[6 empty] | 4.53 | 4.53 | 10 | 0 | 18.6809 | 0.93405 | 431.4178 | -0.86809 | -52.0854 | 0 | 0 | -0.8516 | 1.2624 | 65.3 | 14.415
[6 empty] | 4.54 | 4.54 | 10 | 0 | 18.5073 | 0.92536 | 422.7369 | -0.85073 | -51.0437 | 0 | 0 | -0.83457 | 1.254 | 65.4 | 14.4053
[6 empty] | 4.55 | 4.55 | 10 | 0 | 18.3371 | 0.91686 | 414.2296 | -0.83371 | -50.0229 | 0 | 0 | -0.81787 | 1.2459 | 65.5 | 14.3956
[6 empty] | 4.56 | 4.56 | 10 | 0 | 18.1704 | 0.90852 | 405.8925 | -0.81704 | -49.0224 | 0 | 0 | -0.80152 | 1.2378 | 65.6 | 14.386
[6 empty] | 4.57 | 4.57 | 10 | 0 | 18.007 | 0.90035 | 397.7221 | -0.8007 | -48.042 | 0 | 0 | -0.78549 | 1.23 | 65.7 | 14.3764
[6 empty] | 4.58 | 4.58 | 10 | 0 | 17.8469 | 0.89234 | 389.7151 | -0.78469 | -47.0811 | 0 | 0 | -0.76978 | 1.2223 | 65.8 | 14.3668
[6 empty] | 4.59 | 4.59 | 10 | 0 | 17.6899 | 0.8845 | 381.8682 | -0.76899 | -46.1395 | 0 | 0 | -0.75438 | 1.2147 | 65.9 | 14.3573
[6 empty] | 4.6 | 4.6 | 10 | 0 | 17.5361 | 0.87681 | 374.1783 | -0.75361 | -45.2167 | 0 | 0 | -0.73929 | 1.2074 | 66 | 14.3478
[6 empty] | 4.61 | 4.61 | 10 | 0 | 17.3854 | 0.86927 | 366.6422 | -0.73854 | -44.3124 | 0 | 0 | -0.72451 | 1.2001 | 66.1 | 14.3384
[6 empty] | 4.62 | 4.62 | 10 | 0 | 17.2377 | 0.86188 | 359.2568 | -0.72377 | -43.4261 | 0 | 0 | -0.71002 | 1.193 | 66.2 | 14.329
[6 empty] | 4.63 | 4.63 | 10 | 0 | 17.0929 | 0.85465 | 352.0191 | -0.70929 | -42.5576 | 0 | 0 | -0.69582 | 1.1861 | 66.3 | 14.3197
[6 empty] | 4.64 | 4.64 | 10 | 0 | 16.9511 | 0.84755 | 344.9262 | -0.69511 | -41.7064 | 0 | 0 | -0.6819 | 1.1792 | 66.4 | 14.3103
[6 empty] | 4.65 | 4.65 | 10 | 0 | 16.8121 | 0.8406 | 337.9751 | -0.68121 | -40.8723 | 0 | 0 | -0.66826 | 1.1726 | 66.5 | 14.3011
[6 empty] | 4.66 | 4.66 | 10 | 0 | 16.6758 | 0.83379 | 331.163 | -0.66758 | -40.0549 | 0 | 0 | -0.6549 | 1.166 | 66.6 | 14.2918
[6 empty] | 4.67 | 4.67 | 10 | 0 | 16.5423 | 0.82711 | 324.4872 | -0.65423 | -39.2538 | 0 | 0 | -0.6418 | 1.1596 | 66.7 | 14.2827
[6 empty] | 4.68 | 4.68 | 10 | 0 | 16.4115 | 0.82057 | 317.9449 | -0.64115 | -38.4687 | 0 | 0 | -0.62896 | 1.1533 | 66.8 | 14.2735
[6 empty] | 4.69 | 4.69 | 10 | 0 | 16.2832 | 0.81416 | 311.5335 | -0.62832 | -37.6993 | 0 | 0 | -0.61638 | 1.1471 | 66.9 | 14.2644
[6 empty] | 4.7 | 4.7 | 10 | 0 | 16.1576 | 0.80788 | 305.2503 | -0.61576 | -36.9453 | 0 | 0 | -0.60406 | 1.1411 | 67 | 14.2553
[6 empty] | 4.71 | 4.71 | 10 | 0 | 16.0344 | 0.80172 | 299.0927 | -0.60344 | -36.2064 | 0 | 0 | -0.59198 | 1.1352 | 67.1 | 14.2463
[6 empty] | 4.72 | 4.72 | 10 | 0 | 15.9137 | 0.79569 | 293.0583 | -0.59137 | -35.4823 | 0 | 0 | -0.58014 | 1.1294 | 67.2 | 14.2373
[6 empty] | 4.73 | 4.73 | 10 | 0 | 15.7954 | 0.78977 | 287.1446 | -0.57954 | -34.7727 | 0 | 0 | -0.56853 | 1.1237 | 67.3 | 14.2283
[6 empty] | 4.74 | 4.74 | 10 | 0 | 15.6795 | 0.78398 | 281.3491 | -0.56795 | -34.0772 | 0 | 0 | -0.55716 | 1.1181 | 67.4 | 14.2194
[6 empty] | 4.75 | 4.75 | 10 | 0 | 15.5659 | 0.7783 | 275.6696 | -0.55659 | -33.3957 | 0 | 0 | -0.54602 | 1.1127 | 67.5 | 14.2105
[6 empty] | 4.76 | 4.76 | 10 | 0 | 15.4546 | 0.77273 | 270.1037 | -0.54546 | -32.7277 | 0 | 0 | -0.5351 | 1.1073 | 67.6 | 14.2017
[6 empty] | 4.77 | 4.77 | 10 | 0 | 15.3455 | 0.76728 | 264.649 | -0.53455 | -32.0732 | 0 | 0 | -0.5244 | 1.1021 | 67.7 | 14.1929
[6 empty] | 4.78 | 4.78 | 10 | 0 | 15.2386 | 0.76193 | 259.3035 | -0.52386 | -31.4317 | 0 | 0 | -0.51391 | 1.0969 | 67.8 | 14.1841
[6 empty] | 4.79 | 4.79 | 10 | 0 | 15.1338 | 0.75669 | 254.0649 | -0.51338 | -30.8031 | 0 | 0 | -0.50363 | 1.0919 | 67.9 | 14.1754
[6 empty] | 4.8 | 4.8 | 10 | 0 | 15.0312 | 0.75156 | 248.931 | -0.50312 | -30.187 | 0 | 0 | -0.49356 | 1.0869 | 68 | 14.1667
[6 empty] | 4.81 | 4.81 | 10 | 0 | 14.9305 | 0.74653 | 243.8999 | -0.49305 | -29.5833 | 0 | 0 | -0.48369 | 1.0821 | 68.1 | 14.158
[6 empty] | 4.82 | 4.82 | 10 | 0 | 14.8319 | 0.7416 | 238.9693 | -0.48319 | -28.9916 | 0 | 0 | -0.47401 | 1.0774 | 68.2 | 14.1494
[6 empty] | 4.83 | 4.83 | 10 | 0 | 14.7353 | 0.73676 | 234.1374 | -0.47353 | -28.4118 | 0 | 0 | -0.46453 | 1.0727 | 68.3 | 14.1408
[6 empty] | 4.84 | 4.84 | 10 | 0 | 14.6406 | 0.73203 | 229.4021 | -0.46406 | -27.8436 | 0 | 0 | -0.45524 | 1.0682 | 68.4 | 14.1322
[6 empty] | 4.85 | 4.85 | 10 | 0 | 14.5478 | 0.72739 | 224.7615 | -0.45478 | -27.2867 | 0 | 0 | -0.44614 | 1.0637 | 68.5 | 14.1237
[6 empty] | 4.86 | 4.86 | 10 | 0 | 14.4568 | 0.72284 | 220.2137 | -0.44568 | -26.741 | 0 | 0 | -0.43721 | 1.0593 | 68.6 | 14.1152
[6 empty] | 4.87 | 4.87 | 10 | 0 | 14.3677 | 0.71838 | 215.7569 | -0.43677 | -26.2061 | 0 | 0 | -0.42847 | 1.0551 | 68.7 | 14.1068
[6 empty] | 4.88 | 4.88 | 10 | 0 | 14.2803 | 0.71402 | 211.3892 | -0.42803 | -25.682 | 0 | 0 | -0.4199 | 1.0509 | 68.8 | 14.0984
[6 empty] | 4.89 | 4.89 | 10 | 0 | 14.1947 | 0.70974 | 207.1088 | -0.41947 | -25.1684 | 0 | 0 | -0.4115 | 1.0467 | 68.9 | 14.09
[6 empty] | 4.9 | 4.9 | 10 | 0 | 14.1108 | 0.70554 | 202.9141 | -0.41108 | -24.665 | 0 | 0 | -0.40327 | 1.0427 | 69 | 14.0816
[6 empty] | 4.91 | 4.91 | 10 | 0 | 14.0286 | 0.70143 | 198.8033 | -0.40286 | -24.1717 | 0 | 0 | -0.39521 | 1.0388 | 69.1 | 14.0733
[6 empty] | 4.92 | 4.92 | 10 | 0 | 13.948 | 0.6974 | 194.7747 | -0.3948 | -23.6883 | 0 | 0 | -0.3873 | 1.0349 | 69.2 | 14.065
[6 empty] | 4.93 | 4.93 | 10 | 0 | 13.8691 | 0.69345 | 190.8266 | -0.38691 | -23.2145 | 0 | 0 | -0.37956 | 1.0311 | 69.3 | 14.0568
[6 empty] | 4.94 | 4.94 | 10 | 0 | 13.7917 | 0.68959 | 186.9575 | -0.37917 | -22.7502 | 0 | 0 | -0.37197 | 1.0274 | 69.4 | 14.0486
[6 empty] | 4.95 | 4.95 | 10 | 0 | 13.7159 | 0.68579 | 183.1658 | -0.37159 | -22.2952 | 0 | 0 | -0.36453 | 1.0237 | 69.5 | 14.0404
[6 empty] | 4.96 | 4.96 | 10 | 0 | 13.6416 | 0.68208 | 179.45 | -0.36416 | -21.8493 | 0 | 0 | -0.35724 | 1.0201 | 69.6 | 14.0323
[6 empty] | 4.97 | 4.97 | 10 | 0 | 13.5687 | 0.67844 | 175.8084 | -0.35687 | -21.4123 | 0 | 0 | -0.35009 | 1.0166 | 69.7 | 14.0241
[6 empty] | 4.98 | 4.98 | 10 | 0 | 13.4973 | 0.67487 | 172.2397 | -0.34973 | -20.9841 | 0 | 0 | -0.34309 | 1.0132 | 69.8 | 14.0161
[6 empty] | 4.99 | 4.99 | 10 | 0 | 13.4274 | 0.67137 | 168.7424 | -0.34274 | -20.5644 | 0 | 0 | -0.33623 | 1.0099 | 69.9 | 14.008
[6 empty] | 5 | 5 | 10 | 0 | 13.3589 | 0.66794 | 165.315 | -0.33589 | -20.1531 | 0 | 0 | -0.3295 | 1.0066 | 70 | 14
[6 empty] | 5.01 | 5.01 | 10 | 0 | 13.2917 | 0.66458 | 161.9561 | -0.32917 | -19.75 | 0 | 0 | -0.32291 | 1.0033 | 70.1 | 13.992
[6 empty] | 5.02 | 5.02 | 10 | 0 | 13.2258 | 0.66129 | 158.6644 | -0.32258 | -19.355 | 0 | 0 | -0.31645 | 1.0002 | 70.2 | 13.9841
[6 empty] | 5.03 | 5.03 | 10 | 0 | 13.1613 | 0.65807 | 155.4386 | -0.31613 | -18.9679 | 0 | 0 | -0.31013 | 0.99706 | 70.3 | 13.9761
[6 empty] | 5.04 | 5.04 | 10 | 0 | 13.0981 | 0.6549 | 152.2773 | -0.30981 | -18.5886 | 0 | 0 | -0.30392 | 0.99402 | 70.4 | 13.9683
[6 empty] | 5.05 | 5.05 | 10 | 0 | 13.0361 | 0.65181 | 149.1792 | -0.30361 | -18.2168 | 0 | 0 | -0.29784 | 0.99105 | 70.5 | 13.9604
[6 empty] | 5.06 | 5.06 | 10 | 0 | 12.9754 | 0.64877 | 146.143 | -0.29754 | -17.8525 | 0 | 0 | -0.29189 | 0.98813 | 70.6 | 13.9526
[6 empty] | 5.07 | 5.07 | 10 | 0 | 12.9159 | 0.6458 | 143.1676 | -0.29159 | -17.4954 | 0 | 0 | -0.28605 | 0.98527 | 70.7 | 13.9448
[6 empty] | 5.08 | 5.08 | 10 | 0 | 12.8576 | 0.64288 | 140.2517 | -0.28576 | -17.1455 | 0 | 0 | -0.28033 | 0.98246 | 70.8 | 13.937
[6 empty] | 5.09 | 5.09 | 10 | 0 | 12.8004 | 0.64002 | 137.3941 | -0.28004 | -16.8026 | 0 | 0 | -0.27472 | 0.97972 | 70.9 | 13.9293
[6 empty] | 5.1 | 5.1 | 10 | 0 | 12.7444 | 0.63722 | 134.5937 | -0.27444 | -16.4666 | 0 | 0 | -0.26923 | 0.97702 | 71 | 13.9216
[6 empty] | 5.11 | 5.11 | 10 | 0 | 12.6895 | 0.63448 | 131.8493 | -0.26895 | -16.1372 | 0 | 0 | -0.26384 | 0.97439 | 71.1 | 13.9139
[6 empty] | 5.12 | 5.12 | 10 | 0 | 12.6357 | 0.63179 | 129.1597 | -0.26357 | -15.8145 | 0 | 0 | -0.25857 | 0.9718 | 71.2 | 13.9063
[6 empty] | 5.13 | 5.13 | 10 | 0 | 12.583 | 0.62915 | 126.524 | -0.2583 | -15.4982 | 0 | 0 | -0.2534 | 0.96927 | 71.3 | 13.8986
[6 empty] | 5.14 | 5.14 | 10 | 0 | 12.5314 | 0.62657 | 123.941 | -0.25314 | -15.1882 | 0 | 0 | -0.24833 | 0.96678 | 71.4 | 13.8911
[6 empty] | 5.15 | 5.15 | 10 | 0 | 12.4807 | 0.62404 | 121.4096 | -0.24807 | -14.8845 | 0 | 0 | -0.24336 | 0.96435 | 71.5 | 13.8835
[6 empty] | 5.16 | 5.16 | 10 | 0 | 12.4311 | 0.62156 | 118.9288 | -0.24311 | -14.5868 | 0 | 0 | -0.23849 | 0.96196 | 71.6 | 13.876
[6 empty] | 5.17 | 5.17 | 10 | 0 | 12.3825 | 0.61913 | 116.4977 | -0.23825 | -14.295 | 0 | 0 | -0.23372 | 0.95963 | 71.7 | 13.8685
[6 empty] | 5.18 | 5.18 | 10 | 0 | 12.3349 | 0.61674 | 114.1152 | -0.23349 | -14.0091 | 0 | 0 | -0.22905 | 0.95734 | 71.8 | 13.861
[6 empty] | 5.19 | 5.19 | 10 | 0 | 12.2882 | 0.61441 | 111.7804 | -0.22882 | -13.729 | 0 | 0 | -0.22447 | 0.95509 | 71.9 | 13.8536
[6 empty] | 5.2 | 5.2 | 10 | 0 | 12.2424 | 0.61212 | 109.4922 | -0.22424 | -13.4544 | 0 | 0 | -0.21998 | 0.95289 | 72 | 13.8462
[6 empty] | 5.21 | 5.21 | 10 | 0 | 12.1975 | 0.60988 | 107.2498 | -0.21975 | -13.1853 | 0 | 0 | -0.21558 | 0.95074 | 72.1 | 13.8388
[6 empty] | 5.22 | 5.22 | 10 | 0 | 12.1536 | 0.60768 | 105.0523 | -0.21536 | -12.9216 | 0 | 0 | -0.21127 | 0.94862 | 72.2 | 13.8314
[6 empty] | 5.23 | 5.23 | 10 | 0 | 12.1105 | 0.60553 | 102.8987 | -0.21105 | -12.6631 | 0 | 0 | -0.20704 | 0.94655 | 72.3 | 13.8241
[6 empty] | 5.24 | 5.24 | 10 | 0 | 12.0683 | 0.60342 | 100.7881 | -0.20683 | -12.4099 | 0 | 0 | -0.2029 | 0.94452 | 72.4 | 13.8168
[6 empty] | 5.25 | 5.25 | 10 | 0 | 12.0269 | 0.60135 | 98.7198 | -0.20269 | -12.1617 | 0 | 0 | -0.19884 | 0.94254 | 72.5 | 13.8095
[6 empty] | 5.26 | 5.26 | 10 | 0 | 11.9864 | 0.59932 | 96.6929 | -0.19864 | -11.9185 | 0 | 0 | -0.19487 | 0.94059 | 72.6 | 13.8023
[6 empty] | 5.27 | 5.27 | 10 | 0 | 11.9467 | 0.59733 | 94.7065 | -0.19467 | -11.6801 | 0 | 0 | -0.19097 | 0.93868 | 72.7 | 13.7951
[6 empty] | 5.28 | 5.28 | 10 | 0 | 11.9077 | 0.59539 | 92.7598 | -0.19077 | -11.4465 | 0 | 0 | -0.18715 | 0.93681 | 72.8 | 13.7879
[6 empty] | 5.29 | 5.29 | 10 | 0 | 11.8696 | 0.59348 | 90.852 | -0.18696 | -11.2176 | 0 | 0 | -0.18341 | 0.93497 | 72.9 | 13.7807
[6 empty] | 5.3 | 5.3 | 10 | 0 | 11.8322 | 0.59161 | 88.9824 | -0.18322 | -10.9932 | 0 | 0 | -0.17974 | 0.93317 | 73 | 13.7736
[6 empty] | 5.31 | 5.31 | 10 | 0 | 11.7956 | 0.58978 | 87.1502 | -0.17956 | -10.7733 | 0 | 0 | -0.17614 | 0.93141 | 73.1 | 13.7665
[6 empty] | 5.32 | 5.32 | 10 | 0 | 11.7596 | 0.58798 | 85.3547 | -0.17596 | -10.5579 | 0 | 0 | -0.17262 | 0.92969 | 73.2 | 13.7594
[6 empty] | 5.33 | 5.33 | 10 | 0 | 11.7245 | 0.58622 | 83.595 | -0.17245 | -10.3467 | 0 | 0 | -0.16917 | 0.92799 | 73.3 | 13.7523
[6 empty] | 5.34 | 5.34 | 10 | 0 | 11.69 | 0.5845 | 81.8706 | -0.169 | -10.1398 | 0 | 0 | -0.16579 | 0.92634 | 73.4 | 13.7453
[6 empty] | 5.35 | 5.35 | 10 | 0 | 11.6562 | 0.58281 | 80.1806 | -0.16562 | -9.937 | 0 | 0 | -0.16247 | 0.92471 | 73.5 | 13.7383
[6 empty] | 5.36 | 5.36 | 10 | 0 | 11.623 | 0.58115 | 78.5245 | -0.1623 | -9.7382 | 0 | 0 | -0.15922 | 0.92312 | 73.6 | 13.7313
[6 empty] | 5.37 | 5.37 | 10 | 0 | 11.5906 | 0.57953 | 76.9014 | -0.15906 | -9.5435 | 0 | 0 | -0.15604 | 0.92156 | 73.7 | 13.7244
[6 empty] | 5.38 | 5.38 | 10 | 0 | 11.5588 | 0.57794 | 75.3108 | -0.15588 | -9.3526 | 0 | 0 | -0.15292 | 0.92003 | 73.8 | 13.7175
[6 empty] | 5.39 | 5.39 | 10 | 0 | 11.5276 | 0.57638 | 73.7521 | -0.15276 | -9.1656 | 0 | 0 | -0.14986 | 0.91853 | 73.9 | 13.7106
[6 empty] | 5.4 | 5.4 | 10 | 0 | 11.497 | 0.57485 | 72.2245 | -0.1497 | -8.9822 | 0 | 0 | -0.14686 | 0.91706 | 74 | 13.7037
[6 empty] | 5.41 | 5.41 | 10 | 0 | 11.4671 | 0.57336 | 70.7274 | -0.14671 | -8.8026 | 0 | 0 | -0.14392 | 0.91562 | 74.1 | 13.6969
[6 empty] | 5.42 | 5.42 | 10 | 0 | 11.4378 | 0.57189 | 69.2603 | -0.14378 | -8.6265 | 0 | 0 | -0.14104 | 0.91421 | 74.2 | 13.69
[6 empty] | 5.43 | 5.43 | 10 | 0 | 11.409 | 0.57045 | 67.8226 | -0.1409 | -8.454 | 0 | 0 | -0.13822 | 0.91283 | 74.3 | 13.6832
[6 empty] | 5.44 | 5.44 | 10 | 0 | 11.3808 | 0.56904 | 66.4136 | -0.13808 | -8.2849 | 0 | 0 | -0.13546 | 0.91148 | 74.4 | 13.6765
[6 empty] | 5.45 | 5.45 | 10 | 0 | 11.3532 | 0.56766 | 65.0328 | -0.13532 | -8.1192 | 0 | 0 | -0.13275 | 0.91015 | 74.5 | 13.6697
[6 empty] | 5.46 | 5.46 | 10 | 0 | 11.3261 | 0.56631 | 63.6795 | -0.13261 | -7.9569 | 0 | 0 | -0.13009 | 0.90885 | 74.6 | 13.663
[6 empty] | 5.47 | 5.47 | 10 | 0 | 11.2996 | 0.56498 | 62.3534 | -0.12996 | -7.7977 | 0 | 0 | -0.12749 | 0.90757 | 74.7 | 13.6563
[6 empty] | 5.48 | 5.48 | 10 | 0 | 11.2736 | 0.56368 | 61.0538 | -0.12736 | -7.6418 | 0 | 0 | -0.12494 | 0.90632 | 74.8 | 13.6496
[6 empty] | 5.49 | 5.49 | 10 | 0 | 11.2482 | 0.56241 | 59.7802 | -0.12482 | -7.4889 | 0 | 0 | -0.12244 | 0.9051 | 74.9 | 13.643
[6 empty] | 5.5 | 5.5 | 10 | 0 | 11.2232 | 0.56116 | 58.532 | -0.12232 | -7.3391 | 0 | 0 | -0.12 | 0.9039 | 75 | 13.6364
[6 empty] | 5.51 | 5.51 | 10 | 0 | 11.1987 | 0.55994 | 57.3088 | -0.11987 | -7.1924 | 0 | 0 | -0.1176 | 0.90272 | 75.1 | 13.6298
[6 empty] | 5.52 | 5.52 | 10 | 0 | 11.1748 | 0.55874 | 56.1101 | -0.11748 | -7.0485 | 0 | 0 | -0.11524 | 0.90157 | 75.2 | 13.6232
[6 empty] | 5.53 | 5.53 | 10 | 0 | 11.1513 | 0.55756 | 54.9353 | -0.11513 | -6.9075 | 0 | 0 | -0.11294 | 0.90044 | 75.3 | 13.6166
[6 empty] | 5.54 | 5.54 | 10 | 0 | 11.1282 | 0.55641 | 53.7841 | -0.11282 | -6.7694 | 0 | 0 | -0.11068 | 0.89934 | 75.4 | 13.6101
[6 empty] | 5.55 | 5.55 | 10 | 0 | 11.1057 | 0.55528 | 52.6558 | -0.11057 | -6.634 | 0 | 0 | -0.10847 | 0.89825 | 75.5 | 13.6036
[6 empty] | 5.56 | 5.56 | 10 | 0 | 11.0836 | 0.55418 | 51.5502 | -0.10836 | -6.5013 | 0 | 0 | -0.1063 | 0.89719 | 75.6 | 13.5971
[6 empty] | 5.57 | 5.57 | 10 | 0 | 11.0619 | 0.55309 | 50.4666 | -0.10619 | -6.3713 | 0 | 0 | -0.10417 | 0.89615 | 75.7 | 13.5907
[6 empty] | 5.58 | 5.58 | 10 | 0 | 11.0406 | 0.55203 | 49.4047 | -0.10406 | -6.2439 | 0 | 0 | -0.10209 | 0.89512 | 75.8 | 13.5842
[6 empty] | 5.59 | 5.59 | 10 | 0 | 11.0198 | 0.55099 | 48.3641 | -0.10198 | -6.119 | 0 | 0 | -0.10005 | 0.89412 | 75.9 | 13.5778
[6 empty] | 5.6 | 5.6 | 10 | 0 | 10.9994 | 0.54997 | 47.3442 | -0.099944 | -5.9966 | 0 | 0 | -0.098045 | 0.89314 | 76 | 13.5714
[6 empty] | 5.61 | 5.61 | 10 | 0 | 10.9794 | 0.54897 | 46.3448 | -0.097945 | -5.8767 | 0 | 0 | -0.096084 | 0.89218 | 76.1 | 13.5651
[6 empty] | 5.62 | 5.62 | 10 | 0 | 10.9599 | 0.54799 | 45.3654 | -0.095986 | -5.7592 | 0 | 0 | -0.094162 | 0.89124 | 76.2 | 13.5587
[6 empty] | 5.63 | 5.63 | 10 | 0 | 10.9407 | 0.54703 | 44.4055 | -0.094066 | -5.644 | 0 | 0 | -0.092279 | 0.89032 | 76.3 | 13.5524
[6 empty] | 5.64 | 5.64 | 10 | 0 | 10.9218 | 0.54609 | 43.4648 | -0.092185 | -5.5311 | 0 | 0 | -0.090433 | 0.88941 | 76.4 | 13.5461
[6 empty] | 5.65 | 5.65 | 10 | 0 | 10.9034 | 0.54517 | 42.543 | -0.090341 | -5.4205 | 0 | 0 | -0.088625 | 0.88853 | 76.5 | 13.5398
[6 empty] | 5.66 | 5.66 | 10 | 0 | 10.8853 | 0.54427 | 41.6396 | -0.088534 | -5.3121 | 0 | 0 | -0.086852 | 0.88766 | 76.6 | 13.5336
[6 empty] | 5.67 | 5.67 | 10 | 0 | 10.8676 | 0.54338 | 40.7542 | -0.086764 | -5.2058 | 0 | 0 | -0.085115 | 0.88681 | 76.7 | 13.5273
[6 empty] | 5.68 | 5.68 | 10 | 0 | 10.8503 | 0.54251 | 39.8866 | -0.085028 | -5.1017 | 0 | 0 | -0.083413 | 0.88597 | 76.8 | 13.5211
[6 empty] | 5.69 | 5.69 | 10 | 0 | 10.8333 | 0.54166 | 39.0363 | -0.083328 | -4.9997 | 0 | 0 | -0.081745 | 0.88516 | 76.9 | 13.5149
[6 empty] | 5.7 | 5.7 | 10 | 0 | 10.8166 | 0.54083 | 38.203 | -0.081661 | -4.8997 | 0 | 0 | -0.08011 | 0.88436 | 77 | 13.5088
[6 empty] | 5.71 | 5.71 | 10 | 0 | 10.8003 | 0.54001 | 37.3864 | -0.080028 | -4.8017 | 0 | 0 | -0.078507 | 0.88357 | 77.1 | 13.5026
[6 empty] | 5.72 | 5.72 | 10 | 0 | 10.7843 | 0.53921 | 36.5862 | -0.078427 | -4.7056 | 0 | 0 | -0.076937 | 0.8828 | 77.2 | 13.4965
[6 empty] | 5.73 | 5.73 | 10 | 0 | 10.7686 | 0.53843 | 35.8019 | -0.076859 | -4.6115 | 0 | 0 | -0.075399 | 0.88205 | 77.3 | 13.4904
[6 empty] | 5.74 | 5.74 | 10 | 0 | 10.7532 | 0.53766 | 35.0333 | -0.075322 | -4.5193 | 0 | 0 | -0.073891 | 0.88131 | 77.4 | 13.4843
[6 empty] | 5.75 | 5.75 | 10 | 0 | 10.7382 | 0.53691 | 34.2801 | -0.073815 | -4.4289 | 0 | 0 | -0.072413 | 0.88058 | 77.5 | 13.4783
[6 empty] | 5.76 | 5.76 | 10 | 0 | 10.7234 | 0.53617 | 33.5419 | -0.072339 | -4.3403 | 0 | 0 | -0.070965 | 0.87987 | 77.6 | 13.4722
[6 empty] | 5.77 | 5.77 | 10 | 0 | 10.7089 | 0.53545 | 32.8185 | -0.070892 | -4.2535 | 0 | 0 | -0.069545 | 0.87918 | 77.7 | 13.4662
[6 empty] | 5.78 | 5.78 | 10 | 0 | 10.6947 | 0.53474 | 32.1096 | -0.069474 | -4.1685 | 0 | 0 | -0.068154 | 0.8785 | 77.8 | 13.4602
[6 empty] | 5.79 | 5.79 | 10 | 0 | 10.6808 | 0.53404 | 31.4149 | -0.068085 | -4.0851 | 0 | 0 | -0.066791 | 0.87783 | 77.9 | 13.4542
[6 empty] | 5.8 | 5.8 | 10 | 0 | 10.6672 | 0.53336 | 30.734 | -0.066723 | -4.0034 | 0 | 0 | -0.065455 | 0.87718 | 78 | 13.4483
[6 empty] | 5.81 | 5.81 | 10 | 0 | 10.6539 | 0.53269 | 30.0668 | -0.065389 | -3.9233 | 0 | 0 | -0.064146 | 0.87653 | 78.1 | 13.4423
[6 empty] | 5.82 | 5.82 | 10 | 0 | 10.6408 | 0.53204 | 29.4129 | -0.064081 | -3.8449 | 0 | 0 | -0.062863 | 0.87591 | 78.2 | 13.4364
[6 empty] | 5.83 | 5.83 | 10 | 0 | 10.628 | 0.5314 | 28.7721 | -0.062799 | -3.768 | 0 | 0 | -0.061606 | 0.87529 | 78.3 | 13.4305
[6 empty] | 5.84 | 5.84 | 10 | 0 | 10.6154 | 0.53077 | 28.1441 | -0.061543 | -3.6926 | 0 | 0 | -0.060374 | 0.87469 | 78.4 | 13.4247
[6 empty] | 5.85 | 5.85 | 10 | 0 | 10.6031 | 0.53016 | 27.5287 | -0.060312 | -3.6187 | 0 | 0 | -0.059167 | 0.87409 | 78.5 | 13.4188
[6 empty] | 5.86 | 5.86 | 10 | 0 | 10.5911 | 0.52955 | 26.9255 | -0.059106 | -3.5464 | 0 | 0 | -0.057983 | 0.87351 | 78.6 | 13.413
[6 empty] | 5.87 | 5.87 | 10 | 0 | 10.5792 | 0.52896 | 26.3345 | -0.057924 | -3.4754 | 0 | 0 | -0.056824 | 0.87295 | 78.7 | 13.4072
[6 empty] | 5.88 | 5.88 | 10 | 0 | 10.5677 | 0.52838 | 25.7552 | -0.056766 | -3.4059 | 0 | 0 | -0.055687 | 0.87239 | 78.8 | 13.4014
[6 empty] | 5.89 | 5.89 | 10 | 0 | 10.5563 | 0.52782 | 25.1876 | -0.05563 | -3.3378 | 0 | 0 | -0.054573 | 0.87184 | 78.9 | 13.3956
[6 empty] | 5.9 | 5.9 | 10 | 0 | 10.5452 | 0.52726 | 24.6313 | -0.054518 | -3.2711 | 0 | 0 | -0.053482 | 0.87131 | 79 | 13.3898 | bar: y
[6 empty] | 5.91 | 5.91 | 10 | 0 | 10.5343 | 0.52671 | 24.0861 | -0.053427 | -3.2056 | 0 | 0 | -0.052412 | 0.87078 | 79.1 | 13.3841 | bar: y
[6 empty] | 5.92 | 5.92 | 10 | 0 | 10.5236 | 0.52618 | 23.5518 | -0.052359 | -3.1415 | 0 | 0 | -0.051364 | 0.87027 | 79.2 | 13.3784 | bar: y
[6 empty] | 5.93 | 5.93 | 10 | 0 | 10.5131 | 0.52566 | 23.0282 | -0.051312 | -3.0787 | 0 | 0 | -0.050337 | 0.86977 | 79.3 | 13.3727 | bar: y
[6 empty] | 5.94 | 5.94 | 10 | 0 | 10.5029 | 0.52514 | 22.5151 | -0.050285 | -3.0171 | 0 | 0 | -0.04933 | 0.86927 | 79.4 | 13.367 | bar: y
[6 empty] | 5.95 | 5.95 | 10 | 0 | 10.4928 | 0.52464 | 22.0123 | -0.04928 | -2.9568 | 0 | 0 | -0.048343 | 0.86879 | 79.5 | 13.3613 | bar: y
[6 empty] | 5.96 | 5.96 | 10 | 0 | 10.4829 | 0.52415 | 21.5195 | -0.048294 | -2.8976 | 0 | 0 | -0.047376 | 0.86832 | 79.6 | 13.3557 | bar: y
[6 empty] | 5.97 | 5.97 | 10 | 0 | 10.4733 | 0.52366 | 21.0365 | -0.047328 | -2.8397 | 0 | 0 | -0.046429 | 0.86785 | 79.7 | 13.3501 | bar: y
[6 empty] | 5.98 | 5.98 | 10 | 0 | 10.4638 | 0.52319 | 20.5632 | -0.046382 | -2.7829 | 0 | 0 | -0.0455 | 0.8674 | 79.8 | 13.3445
[6 empty] | 5.99 | 5.99 | 10 | 0 | 10.4545 | 0.52273 | 20.0994 | -0.045454 | -2.7272 | 0 | 0 | -0.04459 | 0.86695 | 79.9 | 13.3389
[6 empty] | 6 | 6 | 10 | 0 | 10.4454 | 0.52227 | 19.6449 | -0.044545 | -2.6727 | 0 | 0 | -0.043699 | 0.86651 | 80 | 13.3333
Menu
Passive Pressure -Volume Curve
Blue = Static Green = Dynamic

## PS-old
Pressure Support
Lung Model
| Elastance (cm H2O/L) | 25
| Compliance (L/cm H2O) | 0.04 | 40
| Resistance (cm H2O/L/s) | 10 | 10
| Time Constant (s) | 0.4
| Tidal Volume (mL) | 465.9486
Ventilator Settings
| Inspiratory Pressure (cm H2O) | 15 | 15
| Peak Inspiratory Pressure (cm H2O) | 19
| Peak Inspiratory Flow (L/min) | 76.8241
| Pressure Rise Time (s) | 0.1 | 1
| Inspiratory Flow Time (s) | 0.62
| Flow Cycle Threshold (L/min) | 19.9743
| Flow Cycle (%) | 26 | 26
| PEEP (cm H2O) | 4 | 4
| Peak Inspiratory Flow (L/min) | 76.8241
| Peak Expiratory Flow (L/min) | -68.1667
| Mean Inspiratory Pressure (cm H2O) | 18.3949
[5 empty] | Pressure | Volume | Flow
|  |  |  | DP | P rise-t | R | C | RC | RC-t | t/(RC-t) | [5 empty] | t-peak | 60EV/R |  |  | Mean Inflation Pressure
|  |  |  | 15 | 0.02 | 10 | 0.04 | 0.4 | 0.38 | 0.052632 | [5 empty] | 0.06 | 69.8923 |  |  | 18.3949
[13 empty] | Insp | Total | Total | Insp | Total | Cum |  | Mean
[11 empty] | Insp | Pressure | Volume | Volume | Volume | Flow | Flow | total |  | total
[5 empty] | Time | e-t/t | e-t/RC | 1+t/(RC-t)*e-t/t | RC/(RC-t)*e-t/RC | 1+t/(RC-t)*e-t/t - RC/(RC-t)*e-t/RC | Time | (cm H2O) | (mL) | (mL) | (mL x 0.1) | (L/m) | (L/m) | pressure | interval | Pressure
[5 empty] | 0 | 1 | 1 | 1.0526 | 1.0526 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 4 | 1 | 4
[5 empty] | 0.01 | 0.60653 | 0.97531 | 1.0319 | 1.0266 | 0.0052807 | 0.01 | 9.902 | 3.1684 | 3.1684 | 0.31684 | 34.937 | 34.937 | 13.902 | 2 | 6.951
[5 empty] | 0.02 | 0.36788 | 0.95123 | 1.0194 | 1.0013 | 0.018068 | 0.02 | 13.4818 | 10.8408 | 10.8408 | 1.0841 | 55.2647 | 55.2647 | 27.3838 | 3 | 9.1279
[5 empty] | 0.03 | 0.22313 | 0.92774 | 1.0117 | 0.97657 | 0.035172 | 0.03 | 15.653 | 21.103 | 21.103 | 2.1103 | 66.7528 | 66.7528 | 43.0369 | 4 | 10.7592
[5 empty] | 0.04 | 0.13534 | 0.90484 | 1.0071 | 0.95246 | 0.054662 | 0.04 | 16.97 | 32.7975 | 32.7975 | 3.2797 | 72.9002 | 72.9002 | 60.0069 | 5 | 12.0014
[5 empty] | 0.05 | 0.082085 | 0.8825 | 1.0043 | 0.92894 | 0.075376 | 0.05 | 17.7687 | 45.2257 | 45.2257 | 4.5226 | 75.8285 | 75.8285 | 77.7756 | 6 | 12.9626
[5 empty] | 0.06 | 0.049787 | 0.86071 | 1.0026 | 0.90601 | 0.096612 | 0.06 | 18.2532 | 57.9672 | 57.9672 | 5.7967 | 76.8241 | 76.8241 | 96.0288 | 7 | 13.7184
[5 empty] | 0.07 | 0.030197 | 0.83946 | 1.0016 | 0.88364 | 0.11795 | 0.07 | 18.547 | 70.7702 | 70.7702 | 7.077 | 76.6667 | 76.6667 | 114.5758 | 8 | 14.322
[5 empty] | 0.08 | 0.018316 | 0.81873 | 1.001 | 0.86182 | 0.13914 | 0.08 | 18.7253 | 83.4853 | 83.4853 | 8.3485 | 75.8288 | 75.8288 | 133.3011 | 9 | 14.8112
[5 empty] | 0.09 | 0.011109 | 0.79852 | 1.0006 | 0.84054 | 0.16004 | 0.09 | 18.8334 | 96.0248 | 96.0248 | 9.6025 | 74.5965 | 74.5965 | 152.1345 | 10 | 15.2134
[5 empty] | 0.1 | 0.0067379 | 0.7788 | 1.0004 | 0.81979 | 0.18056 | 0.1 | 18.8989 | 108.3386 | 108.3386 | 10.8339 | 73.1428 | 73.1428 | 171.0334 | 11 | 15.5485
[5 empty] | 0.11 | 0.0040868 | 0.75957 | 1.0002 | 0.79955 | 0.20067 | 0.11 | 18.9387 | 120.3993 | 120.3993 | 12.0399 | 71.5723 | 71.5723 | 189.9721 | 12 | 15.831
[5 empty] | 0.12 | 0.0024788 | 0.74082 | 1.0001 | 0.77981 | 0.22032 | 0.12 | 18.9628 | 132.1931 | 132.1931 | 13.2193 | 69.9479 | 69.9479 | 208.9349 | 13 | 16.0719
[5 empty] | 0.13 | 0.0015034 | 0.72253 | 1.0001 | 0.76056 | 0.23952 | 0.13 | 18.9774 | 143.7144 | 143.7144 | 14.3714 | 68.3075 | 68.3075 | 227.9124 | 14 | 16.2795
[5 empty] | 0.14 | 0.00091188 | 0.70469 | 1 | 0.74178 | 0.25827 | 0.14 | 18.9863 | 154.9626 | 154.9626 | 15.4963 | 66.6735 | 66.6735 | 246.8987 | 15 | 16.4599
[5 empty] | 0.15 | 0.00055308 | 0.68729 | 1 | 0.72346 | 0.27657 | 0.15 | 18.9917 | 165.94 | 165.94 | 16.594 | 65.0592 | 65.0592 | 265.8904 | 16 | 16.6181
[5 empty] | 0.16 | 0.00033546 | 0.67032 | 1 | 0.7056 | 0.29442 | 0.16 | 18.995 | 176.6506 | 176.6506 | 17.6651 | 63.4722 | 63.4722 | 284.8853 | 17 | 16.758
[5 empty] | 0.17 | 0.00020347 | 0.65377 | 1 | 0.68818 | 0.31183 | 0.17 | 18.9969 | 187.0992 | 187.0992 | 18.7099 | 61.9168 | 61.9168 | 303.8823 | 18 | 16.8823
[5 empty] | 0.18 | 0.00012341 | 0.63763 | 1 | 0.67119 | 0.32882 | 0.18 | 18.9981 | 197.2914 | 197.2914 | 19.7291 | 60.3952 | 60.3952 | 322.8804 | 19 | 16.9937
[5 empty] | 0.19 | 7.4852e-05 | 0.62189 | 1 | 0.65462 | 0.34539 | 0.19 | 18.9989 | 207.2329 | 207.2329 | 20.7233 | 58.9083 | 58.9083 | 341.8793 | 20 | 17.094
[5 empty] | 0.2 | 4.54e-05 | 0.60653 | 1 | 0.63845 | 0.36155 | 0.2 | 18.9993 | 216.9294 | 216.9294 | 21.6929 | 57.4565 | 57.4565 | 360.8786 | 21 | 17.1847
[5 empty] | 0.21 | 2.7536e-05 | 0.59156 | 1 | 0.62269 | 0.37731 | 0.21 | 18.9996 | 226.387 | 226.387 | 22.6387 | 56.0395 | 56.0395 | 379.8782 | 22 | 17.2672
[5 empty] | 0.22 | 1.6702e-05 | 0.57695 | 1 | 0.60732 | 0.39269 | 0.22 | 18.9997 | 235.6112 | 235.6112 | 23.5611 | 54.6568 | 54.6568 | 398.878 | 23 | 17.3425
[5 empty] | 0.23 | 1.013e-05 | 0.5627 | 1 | 0.59232 | 0.40768 | 0.23 | 18.9998 | 244.6078 | 244.6078 | 24.4608 | 53.3079 | 53.3079 | 417.8778 | 24 | 17.4116
[5 empty] | 0.24 | 6.1442e-06 | 0.54881 | 1 | 0.5777 | 0.4223 | 0.24 | 18.9999 | 253.3823 | 253.3823 | 25.3382 | 51.9921 | 51.9921 | 436.8777 | 25 | 17.4751
[5 empty] | 0.25 | 3.7267e-06 | 0.53526 | 1 | 0.56343 | 0.43657 | 0.25 | 18.9999 | 261.9403 | 261.9403 | 26.194 | 50.7086 | 50.7086 | 455.8777 | 26 | 17.5338
[5 empty] | 0.26 | 2.2603e-06 | 0.52205 | 1 | 0.54952 | 0.45048 | 0.26 | 19 | 270.2869 | 270.2869 | 27.0287 | 49.4568 | 49.4568 | 474.8776 | 27 | 17.5881
[5 empty] | 0.27 | 1.371e-06 | 0.50916 | 1 | 0.53595 | 0.46405 | 0.27 | 19 | 278.4276 | 278.4276 | 27.8428 | 48.2357 | 48.2357 | 493.8776 | 28 | 17.6385
[5 empty] | 0.28 | 8.3153e-07 | 0.49659 | 1 | 0.52272 | 0.47728 | 0.28 | 19 | 286.3672 | 286.3672 | 28.6367 | 47.0448 | 47.0448 | 512.8776 | 29 | 17.6854
[5 empty] | 0.29 | 5.0435e-07 | 0.48432 | 1 | 0.50982 | 0.49018 | 0.29 | 19 | 294.1108 | 294.1108 | 29.4111 | 45.8833 | 45.8833 | 531.8776 | 30 | 17.7293
[5 empty] | 0.3 | 3.059e-07 | 0.47237 | 1 | 0.49723 | 0.50277 | 0.3 | 19 | 301.6632 | 301.6632 | 30.1663 | 44.7505 | 44.7505 | 550.8776 | 31 | 17.7702
[5 empty] | 0.31 | 1.8554e-07 | 0.4607 | 1 | 0.48495 | 0.51505 | 0.31 | 19 | 309.0292 | 309.0292 | 30.9029 | 43.6456 | 43.6456 | 569.8776 | 32 | 17.8087
[5 empty] | 0.32 | 1.1254e-07 | 0.44933 | 1 | 0.47298 | 0.52702 | 0.32 | 19 | 316.2133 | 316.2133 | 31.6213 | 42.568 | 42.568 | 588.8776 | 33 | 17.8448
[5 empty] | 0.33 | 6.8256e-08 | 0.43823 | 1 | 0.4613 | 0.5387 | 0.33 | 19 | 323.22 | 323.22 | 32.322 | 41.517 | 41.517 | 607.8776 | 34 | 17.8788
[5 empty] | 0.34 | 4.1399e-08 | 0.42741 | 1 | 0.44991 | 0.55009 | 0.34 | 19 | 330.0537 | 330.0537 | 33.0054 | 40.4919 | 40.4919 | 626.8776 | 35 | 17.9108
[5 empty] | 0.35 | 2.511e-08 | 0.41686 | 1 | 0.4388 | 0.5612 | 0.35 | 19 | 336.7187 | 336.7187 | 33.6719 | 39.4922 | 39.4922 | 645.8776 | 36 | 17.941
[5 empty] | 0.36 | 1.523e-08 | 0.40657 | 1 | 0.42797 | 0.57203 | 0.36 | 19 | 343.2192 | 343.2192 | 34.3219 | 38.5171 | 38.5171 | 664.8776 | 37 | 17.9697
[5 empty] | 0.37 | 9.2374e-09 | 0.39653 | 1 | 0.4174 | 0.5826 | 0.37 | 19 | 349.5591 | 349.5591 | 34.9559 | 37.5661 | 37.5661 | 683.8776 | 38 | 17.9968
[5 empty] | 0.38 | 5.6028e-09 | 0.38674 | 1 | 0.4071 | 0.5929 | 0.38 | 19 | 355.7425 | 355.7425 | 35.5743 | 36.6386 | 36.6386 | 702.8776 | 39 | 18.0225
[5 empty] | 0.39 | 3.3983e-09 | 0.37719 | 1 | 0.39704 | 0.60296 | 0.39 | 19 | 361.7733 | 361.7733 | 36.1773 | 35.734 | 35.734 | 721.8776 | 40 | 18.0469
[5 empty] | 0.4 | 2.0612e-09 | 0.36788 | 1 | 0.38724 | 0.61276 | 0.4 | 19 | 367.6551 | 367.6551 | 36.7655 | 34.8517 | 34.8517 | 740.8776 | 41 | 18.0702
[5 empty] | 0.41 | 1.2502e-09 | 0.3588 | 1 | 0.37768 | 0.62232 | 0.41 | 19 | 373.3917 | 373.3917 | 37.3392 | 33.9912 | 33.9912 | 759.8776 | 42 | 18.0923
[5 empty] | 0.42 | 7.5826e-10 | 0.34994 | 1 | 0.36836 | 0.63164 | 0.42 | 19 | 378.9867 | 378.9867 | 37.8987 | 33.152 | 33.152 | 778.8776 | 43 | 18.1134
[5 empty] | 0.43 | 4.5991e-10 | 0.3413 | 1 | 0.35926 | 0.64074 | 0.43 | 19 | 384.4435 | 384.4435 | 38.4444 | 32.3335 | 32.3335 | 797.8776 | 44 | 18.1336
[5 empty] | 0.44 | 2.7895e-10 | 0.33287 | 1 | 0.35039 | 0.64961 | 0.44 | 19 | 389.7656 | 389.7656 | 38.9766 | 31.5352 | 31.5352 | 816.8776 | 45 | 18.1528
[5 empty] | 0.45 | 1.6919e-10 | 0.32465 | 1 | 0.34174 | 0.65826 | 0.45 | 19 | 394.9563 | 394.9563 | 39.4956 | 30.7565 | 30.7565 | 835.8776 | 46 | 18.1713
[5 empty] | 0.46 | 1.0262e-10 | 0.31664 | 1 | 0.3333 | 0.6667 | 0.46 | 19 | 400.0189 | 400.0189 | 40.0019 | 29.9972 | 29.9972 | 854.8776 | 47 | 18.1889
[5 empty] | 0.47 | 6.2241e-11 | 0.30882 | 1 | 0.32507 | 0.67493 | 0.47 | 19 | 404.9564 | 404.9564 | 40.4956 | 29.2565 | 29.2565 | 873.8776 | 48 | 18.2058
[5 empty] | 0.48 | 3.7751e-11 | 0.30119 | 1 | 0.31705 | 0.68295 | 0.48 | 19 | 409.7721 | 409.7721 | 40.9772 | 28.5342 | 28.5342 | 892.8776 | 49 | 18.222
[5 empty] | 0.49 | 2.2897e-11 | 0.29376 | 1 | 0.30922 | 0.69078 | 0.49 | 19 | 414.4688 | 414.4688 | 41.4469 | 27.8297 | 27.8297 | 911.8776 | 50 | 18.2376
[5 empty] | 0.5 | 1.3888e-11 | 0.2865 | 1 | 0.30158 | 0.69842 | 0.5 | 19 | 419.0496 | 419.0496 | 41.905 | 27.1426 | 27.1426 | 930.8776 | 51 | 18.2525
[5 empty] | 0.51 | 8.4235e-12 | 0.27943 | 1 | 0.29414 | 0.70586 | 0.51 | 19 | 423.5173 | 423.5173 | 42.3517 | 26.4724 | 26.4724 | 949.8776 | 52 | 18.2669
[5 empty] | 0.52 | 5.1091e-12 | 0.27253 | 1 | 0.28688 | 0.71312 | 0.52 | 19 | 427.8747 | 427.8747 | 42.7875 | 25.8188 | 25.8188 | 968.8776 | 53 | 18.2807
[5 empty] | 0.53 | 3.0988e-12 | 0.2658 | 1 | 0.27979 | 0.72021 | 0.53 | 19 | 432.1244 | 432.1244 | 43.2124 | 25.1813 | 25.1813 | 987.8776 | 54 | 18.294
[5 empty] | 0.54 | 1.8795e-12 | 0.25924 | 1 | 0.27288 | 0.72712 | 0.54 | 19 | 436.2693 | 436.2693 | 43.6269 | 24.5596 | 24.5596 | 1006.8776 | 55 | 18.3069
[5 empty] | 0.55 | 1.14e-12 | 0.25284 | 1 | 0.26615 | 0.73385 | 0.55 | 19 | 440.3118 | 440.3118 | 44.0312 | 23.9532 | 23.9532 | 1025.8776 | 56 | 18.3192
[5 empty] | 0.56 | 6.9144e-13 | 0.2466 | 1 | 0.25958 | 0.74042 | 0.56 | 19 | 444.2545 | 444.2545 | 44.4255 | 23.3618 | 23.3618 | 1044.8776 | 57 | 18.3312
[5 empty] | 0.57 | 4.1938e-13 | 0.24051 | 1 | 0.25317 | 0.74683 | 0.57 | 19 | 448.0999 | 448.0999 | 44.81 | 22.785 | 22.785 | 1063.8776 | 58 | 18.3427
[5 empty] | 0.58 | 2.5437e-13 | 0.23457 | 1 | 0.24692 | 0.75308 | 0.58 | 19 | 451.8503 | 451.8503 | 45.185 | 22.2224 | 22.2224 | 1082.8776 | 59 | 18.3539
[5 empty] | 0.59 | 1.5428e-13 | 0.22878 | 1 | 0.24082 | 0.75918 | 0.59 | 19 | 455.5082 | 455.5082 | 45.5508 | 21.6738 | 21.6738 | 1101.8776 | 60 | 18.3646
[5 empty] | 0.6 | 9.3576e-14 | 0.22313 | 1 | 0.23487 | 0.76513 | 0.6 | 19 | 459.0757 | 459.0757 | 45.9076 | 21.1386 | 21.1386 | 1120.8776 | 61 | 18.375
[5 empty] | 0.61 | 5.6757e-14 | 0.21762 | 1 | 0.22907 | 0.77093 | 0.61 | 19 | 462.5551 | 462.5551 | 46.2555 | 20.6167 | 20.6167 | 1139.8776 | 62 | 18.3851
[5 empty] | 0.62 | 3.4425e-14 | 0.21225 | 1 | 0.22342 | 0.77658 | 0.62 | 19 | 465.9486 | 465.9486 | 46.5949 | 20.1077 | 20.1077 | 1158.8776 | 63 | 18.3949
[5 empty] | 0.63 | 2.088e-14 | 0.20701 | 1 | 0.2179 | 0.7821 | 0 | 4 | 469.2584 | 454.4443 | 45.4444 | 19.6112 | -68.1667 | 1162.8776 | 64 | 18.17
[5 empty] | 0.64 | 1.2664e-14 | 0.2019 | 1 | 0.21252 | 0.78748 | 0 | 4 | 472.4864 | 443.2241 | 44.3224 | 19.127 | -66.4836 | 1166.8776 | 65 | 17.952
[5 empty] | 0.65 | 7.6812e-15 | 0.19691 | 1 | 0.20728 | 0.79272 | 0 | 4 | 475.6347 | 432.2808 | 43.2281 | 18.6548 | -64.8421 | 1170.8776 | 66 | 17.7406
[5 empty] | 0.66 | 4.6589e-15 | 0.19205 | 1 | 0.20216 | 0.79784 | 0 | 4 | 478.7053 | 421.6078 | 42.1608 | 18.1942 | -63.2412 | 1174.8776 | 67 | 17.5355
[5 empty] | 0.67 | 2.8258e-15 | 0.18731 | 1 | 0.19717 | 0.80283 | 0 | 4 | 481.7001 | 411.1982 | 41.1198 | 17.745 | -61.6797 | 1178.8776 | 68 | 17.3364
[5 empty] | 0.68 | 1.7139e-15 | 0.18268 | 1 | 0.1923 | 0.8077 | 0 | 4 | 484.6209 | 401.0457 | 40.1046 | 17.3069 | -60.1569 | 1182.8776 | 69 | 17.1432
[5 empty] | 0.69 | 1.0395e-15 | 0.17817 | 1 | 0.18755 | 0.81245 | 0 | 4 | 487.4697 | 391.1439 | 39.1144 | 16.8796 | -58.6716 | 1186.8776 | 70 | 16.9554
[5 empty] | 0.7 | 6.3051e-16 | 0.17377 | 1 | 0.18292 | 0.81708 | 0 | 4 | 490.248 | 381.4865 | 38.1486 | 16.4628 | -57.223 | 1190.8776 | 71 | 16.7729
[5 empty] | 0.71 | 3.8242e-16 | 0.16948 | 1 | 0.1784 | 0.8216 | 0 | 4 | 492.9578 | 372.0676 | 37.2068 | 16.0563 | -55.8101 | 1194.8776 | 72 | 16.5955
[5 empty] | 0.72 | 2.3195e-16 | 0.1653 | 1 | 0.174 | 0.826 | 0 | 4 | 495.6007 | 362.8812 | 36.2881 | 15.6599 | -54.4322 | 1198.8776 | 73 | 16.423
[5 empty] | 0.73 | 1.4069e-16 | 0.16122 | 1 | 0.1697 | 0.8303 | 0 | 4 | 498.1783 | 353.9216 | 35.3922 | 15.2733 | -53.0882 | 1202.8776 | 74 | 16.2551
[5 empty] | 0.74 | 8.533e-17 | 0.15724 | 1 | 0.16551 | 0.83449 | 0 | 4 | 500.6923 | 345.1832 | 34.5183 | 14.8962 | -51.7775 | 1206.8776 | 75 | 16.0917
[5 empty] | 0.75 | 5.1756e-17 | 0.15335 | 1 | 0.16143 | 0.83857 | 0 | 4 | 503.1442 | 336.6606 | 33.6661 | 14.5284 | -50.4991 | 1210.8776 | 76 | 15.9326
[5 empty] | 0.76 | 3.1391e-17 | 0.14957 | 1 | 0.15744 | 0.84256 | 0 | 4 | 505.5356 | 328.3485 | 32.8348 | 14.1697 | -49.2523 | 1214.8776 | 77 | 15.7776
[5 empty] | 0.77 | 1.904e-17 | 0.14588 | 1 | 0.15355 | 0.84645 | 0 | 4 | 507.8679 | 320.2415 | 32.0242 | 13.8198 | -48.0362 | 1218.8776 | 78 | 15.6266
[5 empty] | 0.78 | 1.1548e-17 | 0.14227 | 1 | 0.14976 | 0.85024 | 0 | 4 | 510.1427 | 312.3347 | 31.2335 | 13.4786 | -46.8502 | 1222.8776 | 79 | 15.4795
[5 empty] | 0.79 | 7.0044e-18 | 0.13876 | 1 | 0.14606 | 0.85394 | 0 | 4 | 512.3613 | 304.6231 | 30.4623 | 13.1458 | -45.6935 | 1226.8776 | 80 | 15.336
[5 empty] | 0.8 | 4.2484e-18 | 0.13534 | 1 | 0.14246 | 0.85754 | 0 | 4 | 514.5251 | 297.102 | 29.7102 | 12.8212 | -44.5653 | 1230.8776 | 81 | 15.196
[5 empty] | 0.81 | 2.5768e-18 | 0.13199 | 1 | 0.13894 | 0.86106 | 0 | 4 | 516.6355 | 289.7665 | 28.9767 | 12.5047 | -43.465 | 1234.8776 | 82 | 15.0595
[5 empty] | 0.82 | 1.5629e-18 | 0.12873 | 1 | 0.13551 | 0.86449 | 0 | 4 | 518.6937 | 282.6121 | 28.2612 | 12.1959 | -42.3918 | 1238.8776 | 83 | 14.9262
[5 empty] | 0.83 | 9.4794e-19 | 0.12556 | 1 | 0.13216 | 0.86784 | 0 | 4 | 520.7012 | 275.6344 | 27.5634 | 11.8948 | -41.3452 | 1242.8776 | 84 | 14.7962
[5 empty] | 0.84 | 5.7495e-19 | 0.12246 | 1 | 0.1289 | 0.8711 | 0 | 4 | 522.6591 | 268.829 | 26.8829 | 11.6011 | -40.3243 | 1246.8776 | 85 | 14.6691
[5 empty] | 0.85 | 3.4873e-19 | 0.11943 | 1 | 0.12572 | 0.87428 | 0 | 4 | 524.5687 | 262.1916 | 26.2192 | 11.3147 | -39.3287 | 1250.8776 | 86 | 14.5451
[5 empty] | 0.86 | 2.1151e-19 | 0.11648 | 1 | 0.12261 | 0.87739 | 0 | 4 | 526.4311 | 255.718 | 25.5718 | 11.0353 | -38.3577 | 1254.8776 | 87 | 14.4239
[5 empty] | 0.87 | 1.2829e-19 | 0.11361 | 1 | 0.11959 | 0.88041 | 0 | 4 | 528.2475 | 249.4043 | 24.9404 | 10.7629 | -37.4107 | 1258.8776 | 88 | 14.3054
[5 empty] | 0.88 | 7.7811e-20 | 0.1108 | 1 | 0.11663 | 0.88337 | 0 | 4 | 530.0191 | 243.2465 | 24.3247 | 10.4971 | -36.487 | 1262.8776 | 89 | 14.1896
[5 empty] | 0.89 | 4.7195e-20 | 0.10807 | 1 | 0.11376 | 0.88624 | 0 | 4 | 531.7469 | 237.2407 | 23.7241 | 10.238 | -35.5861 | 1266.8776 | 90 | 14.0764
[5 empty] | 0.9 | 2.8625e-20 | 0.1054 | 1 | 0.11095 | 0.88905 | 0 | 4 | 533.4321 | 231.3833 | 23.1383 | 9.9852 | -34.7075 | 1270.8776 | 91 | 13.9657
[5 empty] | 0.91 | 1.7362e-20 | 0.1028 | 1 | 0.10821 | 0.89179 | 0 | 4 | 535.0756 | 225.6704 | 22.567 | 9.7387 | -33.8506 | 1274.8776 | 92 | 13.8574
[5 empty] | 0.92 | 1.0531e-20 | 0.10026 | 1 | 0.10554 | 0.89446 | 0 | 4 | 536.6786 | 220.0986 | 22.0099 | 9.4982 | -33.0148 | 1278.8776 | 93 | 13.7514
[5 empty] | 0.93 | 6.3871e-21 | 0.097783 | 1 | 0.10293 | 0.89707 | 0 | 4 | 538.242 | 214.6643 | 21.4664 | 9.2637 | -32.1996 | 1282.8776 | 94 | 13.6476
[5 empty] | 0.94 | 3.874e-21 | 0.095369 | 1 | 0.10039 | 0.89961 | 0 | 4 | 539.7668 | 209.3642 | 20.9364 | 9.035 | -31.4046 | 1286.8776 | 95 | 13.5461
[5 empty] | 0.95 | 2.3497e-21 | 0.093014 | 1 | 0.09791 | 0.90209 | 0 | 4 | 541.254 | 204.195 | 20.4195 | 8.8119 | -30.6293 | 1290.8776 | 96 | 13.4466
[5 empty] | 0.96 | 1.4252e-21 | 0.090718 | 1 | 0.095493 | 0.90451 | 0 | 4 | 542.7045 | 199.1534 | 19.9153 | 8.5943 | -29.873 | 1294.8776 | 97 | 13.3493
[5 empty] | 0.97 | 8.6441e-22 | 0.088478 | 1 | 0.093135 | 0.90687 | 0 | 4 | 544.1191 | 194.2363 | 19.4236 | 8.3821 | -29.1354 | 1298.8776 | 98 | 13.2539
[5 empty] | 0.98 | 5.2429e-22 | 0.086294 | 1 | 0.090835 | 0.90916 | 0 | 4 | 545.4988 | 189.4406 | 18.9441 | 8.1752 | -28.4161 | 1302.8776 | 99 | 13.1604
[5 empty] | 0.99 | 3.18e-22 | 0.084163 | 1 | 0.088593 | 0.91141 | 0 | 4 | 546.8444 | 184.7633 | 18.4763 | 7.9733 | -27.7145 | 1306.8776 | 100 | 13.0688
[5 empty] | 1 | 1.9287e-22 | 0.082085 | 1 | 0.086405 | 0.91359 | 0 | 4 | 548.1568 | 180.2015 | 18.0201 | 7.7765 | -27.0302 | 1310.8776 | 101 | 12.979
[5 empty] | 1.01 | 1.1698e-22 | 0.080058 | 1 | 0.084272 | 0.91573 | 0 | 4 | 549.4369 | 175.7523 | 17.5752 | 7.5845 | -26.3628 | 1314.8776 | 102 | 12.891
[5 empty] | 1.02 | 7.0955e-23 | 0.078082 | 1 | 0.082191 | 0.91781 | 0 | 4 | 550.6853 | 171.4129 | 17.1413 | 7.3972 | -25.7119 | 1318.8776 | 103 | 12.8046
[5 empty] | 1.03 | 4.3036e-23 | 0.076154 | 1 | 0.080162 | 0.91984 | 0 | 4 | 551.9028 | 167.1807 | 16.7181 | 7.2146 | -25.0771 | 1322.8776 | 104 | 12.72
[5 empty] | 1.04 | 2.6103e-23 | 0.074274 | 1 | 0.078183 | 0.92182 | 0 | 4 | 553.0904 | 163.053 | 16.3053 | 7.0364 | -24.458 | 1326.8776 | 105 | 12.6369
[5 empty] | 1.05 | 1.5832e-23 | 0.07244 | 1 | 0.076252 | 0.92375 | 0 | 4 | 554.2486 | 159.0272 | 15.9027 | 6.8627 | -23.8541 | 1330.8776 | 106 | 12.5554
[5 empty] | 1.06 | 9.6027e-24 | 0.070651 | 1 | 0.07437 | 0.92563 | 0 | 4 | 555.3782 | 155.1008 | 15.5101 | 6.6933 | -23.2651 | 1334.8776 | 107 | 12.4755
[5 empty] | 1.07 | 5.8243e-24 | 0.068907 | 1 | 0.072534 | 0.92747 | 0 | 4 | 556.4799 | 151.2714 | 15.1271 | 6.528 | -22.6907 | 1338.8776 | 108 | 12.397
[5 empty] | 1.08 | 3.5326e-24 | 0.067206 | 1 | 0.070743 | 0.92926 | 0 | 4 | 557.5544 | 147.5365 | 14.7536 | 6.3668 | -22.1305 | 1342.8776 | 109 | 12.32
[5 empty] | 1.09 | 2.1426e-24 | 0.065546 | 1 | 0.068996 | 0.931 | 0 | 4 | 558.6024 | 143.8938 | 14.3894 | 6.2096 | -21.5841 | 1346.8776 | 110 | 12.2443
[5 empty] | 1.1 | 1.2996e-24 | 0.063928 | 1 | 0.067292 | 0.93271 | 0 | 4 | 559.6245 | 140.341 | 14.0341 | 6.0563 | -21.0512 | 1350.8776 | 111 | 12.1701
[5 empty] | 1.11 | 7.8824e-25 | 0.062349 | 1 | 0.065631 | 0.93437 | 0 | 4 | 560.6214 | 136.876 | 13.6876 | 5.9068 | -20.5314 | 1354.8776 | 112 | 12.0971
[5 empty] | 1.12 | 4.7809e-25 | 0.06081 | 1 | 0.064011 | 0.93599 | 0 | 4 | 561.5936 | 133.4965 | 13.3497 | 5.761 | -20.0245 | 1358.8776 | 113 | 12.0255
[5 empty] | 1.13 | 2.8998e-25 | 0.059309 | 1 | 0.06243 | 0.93757 | 0 | 4 | 562.5419 | 130.2005 | 13.02 | 5.6187 | -19.5301 | 1362.8776 | 114 | 11.9551
[5 empty] | 1.14 | 1.7588e-25 | 0.057844 | 1 | 0.060889 | 0.93911 | 0 | 4 | 563.4667 | 126.9858 | 12.6986 | 5.48 | -19.0479 | 1366.8776 | 115 | 11.8859
[5 empty] | 1.15 | 1.0668e-25 | 0.056416 | 1 | 0.059385 | 0.94061 | 0 | 4 | 564.3688 | 123.8505 | 12.3851 | 5.3447 | -18.5776 | 1370.8776 | 116 | 11.8179
[5 empty] | 1.16 | 6.4702e-26 | 0.055023 | 1 | 0.057919 | 0.94208 | 0 | 4 | 565.2485 | 120.7926 | 12.0793 | 5.2127 | -18.1189 | 1374.8776 | 117 | 11.7511
[5 empty] | 1.17 | 3.9244e-26 | 0.053665 | 1 | 0.056489 | 0.94351 | 0 | 4 | 566.1065 | 117.8103 | 11.781 | 5.084 | -17.6715 | 1378.8776 | 118 | 11.6854
[5 empty] | 1.18 | 2.3803e-26 | 0.05234 | 1 | 0.055094 | 0.94491 | 0 | 4 | 566.9433 | 114.9015 | 11.4902 | 4.9585 | -17.2352 | 1382.8776 | 119 | 11.6208
[5 empty] | 1.19 | 1.4437e-26 | 0.051047 | 1 | 0.053734 | 0.94627 | 0 | 4 | 567.7595 | 112.0646 | 11.2065 | 4.8361 | -16.8097 | 1386.8776 | 120 | 11.5573
[5 empty] | 1.2 | 8.7565e-27 | 0.049787 | 1 | 0.052407 | 0.94759 | 0 | 4 | 568.5555 | 109.2977 | 10.9298 | 4.7167 | -16.3947 | 1390.8776 | 121 | 11.4949
[5 empty] | 1.21 | 5.3111e-27 | 0.048558 | 1 | 0.051113 | 0.94889 | 0 | 4 | 569.3319 | 106.5991 | 10.6599 | 4.6002 | -15.9899 | 1394.8776 | 122 | 11.4334
[5 empty] | 1.22 | 3.2213e-27 | 0.047359 | 1 | 0.049851 | 0.95015 | 0 | 4 | 570.0891 | 103.9672 | 10.3967 | 4.4866 | -15.5951 | 1398.8776 | 123 | 11.373
[5 empty] | 1.23 | 1.9538e-27 | 0.04619 | 1 | 0.048621 | 0.95138 | 0 | 4 | 570.8276 | 101.4002 | 10.14 | 4.3759 | -15.21 | 1402.8776 | 124 | 11.3135
[5 empty] | 1.24 | 1.1851e-27 | 0.045049 | 1 | 0.04742 | 0.95258 | 0 | 4 | 571.5479 | 98.8967 | 9.8897 | 4.2678 | -14.8345 | 1406.8776 | 125 | 11.255
[5 empty] | 1.25 | 7.1878e-28 | 0.043937 | 1 | 0.046249 | 0.95375 | 0 | 4 | 572.2504 | 96.4549 | 9.6455 | 4.1624 | -14.4682 | 1410.8776 | 126 | 11.1974
[5 empty] | 1.26 | 4.3596e-28 | 0.042852 | 1 | 0.045108 | 0.95489 | 0 | 4 | 572.9355 | 94.0734 | 9.4073 | 4.0597 | -14.111 | 1414.8776 | 127 | 11.1408
[5 empty] | 1.27 | 2.6442e-28 | 0.041794 | 1 | 0.043994 | 0.95601 | 0 | 4 | 573.6037 | 91.7507 | 9.1751 | 3.9594 | -13.7626 | 1418.8776 | 128 | 11.085
[5 empty] | 1.28 | 1.6038e-28 | 0.040762 | 1 | 0.042908 | 0.95709 | 0 | 4 | 574.2555 | 89.4854 | 8.9485 | 3.8617 | -13.4228 | 1422.8776 | 129 | 11.0301
[5 empty] | 1.29 | 9.7276e-29 | 0.039756 | 1 | 0.041848 | 0.95815 | 0 | 4 | 574.8911 | 87.276 | 8.7276 | 3.7663 | -13.0914 | 1426.8776 | 130 | 10.976
[5 empty] | 1.3 | 5.9001e-29 | 0.038774 | 1 | 0.040815 | 0.95919 | 0 | 4 | 575.511 | 85.1211 | 8.5121 | 3.6733 | -12.7682 | 1430.8776 | 131 | 10.9227
[5 empty] | 1.31 | 3.5786e-29 | 0.037817 | 1 | 0.039807 | 0.96019 | 0 | 4 | 576.1157 | 83.0195 | 8.3019 | 3.5827 | -12.4529 | 1434.8776 | 132 | 10.8703
[5 empty] | 1.32 | 2.1705e-29 | 0.036883 | 1 | 0.038824 | 0.96118 | 0 | 4 | 576.7054 | 80.9697 | 8.097 | 3.4942 | -12.1455 | 1438.8776 | 133 | 10.8186
[5 empty] | 1.33 | 1.3165e-29 | 0.035973 | 1 | 0.037866 | 0.96213 | 0 | 4 | 577.2805 | 78.9706 | 7.8971 | 3.4079 | -11.8456 | 1442.8776 | 134 | 10.7677
[5 empty] | 1.34 | 7.9849e-30 | 0.035084 | 1 | 0.036931 | 0.96307 | 0 | 4 | 577.8415 | 77.0208 | 7.7021 | 3.3238 | -11.5531 | 1446.8776 | 135 | 10.7176
[5 empty] | 1.35 | 4.8431e-30 | 0.034218 | 1 | 0.036019 | 0.96398 | 0 | 4 | 578.3886 | 75.1191 | 7.5119 | 3.2417 | -11.2679 | 1450.8776 | 136 | 10.6682
[5 empty] | 1.36 | 2.9375e-30 | 0.033373 | 1 | 0.03513 | 0.96487 | 0 | 4 | 578.9221 | 73.2644 | 7.3264 | 3.1617 | -10.9897 | 1454.8776 | 137 | 10.6195
[5 empty] | 1.37 | 1.7817e-30 | 0.032549 | 1 | 0.034262 | 0.96574 | 0 | 4 | 579.4426 | 71.4555 | 7.1456 | 3.0836 | -10.7183 | 1458.8776 | 138 | 10.5716
[5 empty] | 1.38 | 1.0806e-30 | 0.031746 | 1 | 0.033416 | 0.96658 | 0 | 4 | 579.9501 | 69.6913 | 6.9691 | 3.0075 | -10.4537 | 1462.8776 | 139 | 10.5243
[5 empty] | 1.39 | 6.5544e-31 | 0.030962 | 1 | 0.032591 | 0.96741 | 0 | 4 | 580.4452 | 67.9706 | 6.7971 | 2.9332 | -10.1956 | 1466.8776 | 140 | 10.4777
[5 empty] | 1.4 | 3.9754e-31 | 0.030197 | 1 | 0.031787 | 0.96821 | 0 | 4 | 580.928 | 66.2924 | 6.6292 | 2.8608 | -9.9439 | 1470.8776 | 141 | 10.4318
[5 empty] | 1.41 | 2.4112e-31 | 0.029452 | 1 | 0.031002 | 0.969 | 0 | 4 | 581.3989 | 64.6556 | 6.4656 | 2.7902 | -9.6983 | 1474.8776 | 142 | 10.3865
[5 empty] | 1.42 | 1.4625e-31 | 0.028725 | 1 | 0.030236 | 0.96976 | 0 | 4 | 581.8581 | 63.0593 | 6.3059 | 2.7213 | -9.4589 | 1478.8776 | 143 | 10.3418
[5 empty] | 1.43 | 8.8704e-32 | 0.028015 | 1 | 0.02949 | 0.97051 | 0 | 4 | 582.306 | 61.5024 | 6.1502 | 2.6541 | -9.2254 | 1482.8776 | 144 | 10.2978
[5 empty] | 1.44 | 5.3802e-32 | 0.027324 | 1 | 0.028762 | 0.97124 | 0 | 4 | 582.7429 | 59.9839 | 5.9984 | 2.5886 | -8.9976 | 1486.8776 | 145 | 10.2543
[5 empty] | 1.45 | 3.2632e-32 | 0.026649 | 1 | 0.028052 | 0.97195 | 0 | 4 | 583.169 | 58.5028 | 5.8503 | 2.5247 | -8.7754 | 1490.8776 | 146 | 10.2115
[5 empty] | 1.46 | 1.9793e-32 | 0.025991 | 1 | 0.027359 | 0.97264 | 0 | 4 | 583.5846 | 57.0584 | 5.7058 | 2.4623 | -8.5588 | 1494.8776 | 147 | 10.1692
[5 empty] | 1.47 | 1.2005e-32 | 0.025349 | 1 | 0.026684 | 0.97332 | 0 | 4 | 583.9898 | 55.6496 | 5.565 | 2.4015 | -8.3474 | 1498.8776 | 148 | 10.1276
[5 empty] | 1.48 | 7.2813e-33 | 0.024724 | 1 | 0.026025 | 0.97398 | 0 | 4 | 584.3851 | 54.2756 | 5.4276 | 2.3422 | -8.1413 | 1502.8776 | 149 | 10.0864
[5 empty] | 1.49 | 4.4163e-33 | 0.024113 | 1 | 0.025382 | 0.97462 | 0 | 4 | 584.7707 | 52.9356 | 5.2936 | 2.2844 | -7.9403 | 1506.8776 | 150 | 10.0459
[5 empty] | 1.5 | 2.6786e-33 | 0.023518 | 1 | 0.024756 | 0.97524 | 0 | 4 | 585.1467 | 51.6286 | 5.1629 | 2.228 | -7.7443 | 1510.8776 | 151 | 10.0058
[5 empty] | 1.51 | 1.6247e-33 | 0.022937 | 1 | 0.024144 | 0.97586 | 0 | 4 | 585.5134 | 50.3539 | 5.0354 | 2.173 | -7.5531 | 1514.8776 | 152 | 9.9663
[5 empty] | 1.52 | 9.8542e-34 | 0.022371 | 1 | 0.023548 | 0.97645 | 0 | 4 | 585.8711 | 49.1106 | 4.9111 | 2.1193 | -7.3666 | 1518.8776 | 153 | 9.9273
[5 empty] | 1.53 | 5.9768e-34 | 0.021818 | 1 | 0.022967 | 0.97703 | 0 | 4 | 586.2199 | 47.8981 | 4.7898 | 2.067 | -7.1847 | 1522.8776 | 154 | 9.8888
[5 empty] | 1.54 | 3.6251e-34 | 0.02128 | 1 | 0.0224 | 0.9776 | 0 | 4 | 586.5602 | 46.7155 | 4.6715 | 2.016 | -7.0073 | 1526.8776 | 155 | 9.8508
[5 empty] | 1.55 | 2.1988e-34 | 0.020754 | 1 | 0.021847 | 0.97815 | 0 | 4 | 586.892 | 45.5621 | 4.5562 | 1.9662 | -6.8343 | 1530.8776 | 156 | 9.8133
[5 empty] | 1.56 | 1.3336e-34 | 0.020242 | 1 | 0.021307 | 0.97869 | 0 | 4 | 587.2156 | 44.4371 | 4.4437 | 1.9177 | -6.6656 | 1534.8776 | 157 | 9.7763
[5 empty] | 1.57 | 8.0888e-35 | 0.019742 | 1 | 0.020781 | 0.97922 | 0 | 4 | 587.5313 | 43.34 | 4.334 | 1.8703 | -6.501 | 1538.8776 | 158 | 9.7397
[5 empty] | 1.58 | 4.9061e-35 | 0.019255 | 1 | 0.020268 | 0.97973 | 0 | 4 | 587.8391 | 42.2699 | 4.227 | 1.8241 | -6.3405 | 1542.8776 | 159 | 9.7036
[5 empty] | 1.59 | 2.9757e-35 | 0.018779 | 1 | 0.019768 | 0.98023 | 0 | 4 | 588.1394 | 41.2263 | 4.1226 | 1.7791 | -6.1839 | 1546.8776 | 160 | 9.668
[5 empty] | 1.6 | 1.8049e-35 | 0.018316 | 1 | 0.01928 | 0.98072 | 0 | 4 | 588.4322 | 40.2084 | 4.0208 | 1.7352 | -6.0313 | 1550.8776 | 161 | 9.6328
[5 empty] | 1.61 | 1.0947e-35 | 0.017863 | 1 | 0.018804 | 0.9812 | 0 | 4 | 588.7178 | 39.2156 | 3.9216 | 1.6923 | -5.8823 | 1554.8776 | 162 | 9.598
[5 empty] | 1.62 | 6.6397e-36 | 0.017422 | 1 | 0.018339 | 0.98166 | 0 | 4 | 588.9964 | 38.2474 | 3.8247 | 1.6505 | -5.7371 | 1558.8776 | 163 | 9.5637
[5 empty] | 1.63 | 4.0272e-36 | 0.016992 | 1 | 0.017887 | 0.98211 | 0 | 4 | 589.2681 | 37.3031 | 3.7303 | 1.6098 | -5.5955 | 1562.8776 | 164 | 9.5297
[5 empty] | 1.64 | 2.4426e-36 | 0.016573 | 1 | 0.017445 | 0.98256 | 0 | 4 | 589.533 | 36.382 | 3.6382 | 1.57 | -5.4573 | 1566.8776 | 165 | 9.4962
[5 empty] | 1.65 | 1.4815e-36 | 0.016163 | 1 | 0.017014 | 0.98299 | 0 | 4 | 589.7915 | 35.4838 | 3.5484 | 1.5313 | -5.3226 | 1570.8776 | 166 | 9.4631
[5 empty] | 1.66 | 8.9858e-37 | 0.015764 | 1 | 0.016594 | 0.98341 | 0 | 4 | 590.0435 | 34.6077 | 3.4608 | 1.4935 | -5.1912 | 1574.8776 | 167 | 9.4304
[5 empty] | 1.67 | 5.4502e-37 | 0.015375 | 1 | 0.016184 | 0.98382 | 0 | 4 | 590.2894 | 33.7532 | 3.3753 | 1.4566 | -5.063 | 1578.8776 | 168 | 9.3981
[5 empty] | 1.68 | 3.3057e-37 | 0.014996 | 1 | 0.015785 | 0.98422 | 0 | 4 | 590.5291 | 32.9198 | 3.292 | 1.4206 | -4.938 | 1582.8776 | 169 | 9.3661
[5 empty] | 1.69 | 2.005e-37 | 0.014625 | 1 | 0.015395 | 0.9846 | 0 | 4 | 590.7629 | 32.107 | 3.2107 | 1.3856 | -4.8161 | 1586.8776 | 170 | 9.3346
[5 empty] | 1.7 | 1.2161e-37 | 0.014264 | 1 | 0.015015 | 0.98499 | 0 | 4 | 590.991 | 31.3143 | 3.1314 | 1.3513 | -4.6971 | 1590.8776 | 171 | 9.3034
[5 empty] | 1.71 | 7.376e-38 | 0.013912 | 1 | 0.014644 | 0.98536 | 0 | 4 | 591.2134 | 30.5412 | 3.0541 | 1.318 | -4.5812 | 1594.8776 | 172 | 9.2725
[5 empty] | 1.72 | 4.4738e-38 | 0.013569 | 1 | 0.014283 | 0.98572 | 0 | 4 | 591.4304 | 29.7871 | 2.9787 | 1.2854 | -4.4681 | 1598.8776 | 173 | 9.2421
[5 empty] | 1.73 | 2.7135e-38 | 0.013234 | 1 | 0.01393 | 0.98607 | 0 | 4 | 591.642 | 29.0517 | 2.9052 | 1.2537 | -4.3577 | 1602.8776 | 174 | 9.2119
[5 empty] | 1.74 | 1.6458e-38 | 0.012907 | 1 | 0.013586 | 0.98641 | 0 | 4 | 591.8483 | 28.3344 | 2.8334 | 1.2228 | -4.2502 | 1606.8776 | 175 | 9.1822
[5 empty] | 1.75 | 9.9824e-39 | 0.012588 | 1 | 0.013251 | 0.98675 | 0 | 4 | 592.0496 | 27.6348 | 2.7635 | 1.1926 | -4.1452 | 1610.8776 | 176 | 9.1527
[5 empty] | 1.76 | 6.0546e-39 | 0.012277 | 1 | 0.012924 | 0.98708 | 0 | 4 | 592.2459 | 26.9525 | 2.6952 | 1.1631 | -4.0429 | 1614.8776 | 177 | 9.1236
[5 empty] | 1.77 | 3.6723e-39 | 0.011974 | 1 | 0.012604 | 0.9874 | 0 | 4 | 592.4373 | 26.287 | 2.6287 | 1.1344 | -3.9431 | 1618.8776 | 178 | 9.0948
[5 empty] | 1.78 | 2.2274e-39 | 0.011679 | 1 | 0.012293 | 0.98771 | 0 | 4 | 592.6241 | 25.638 | 2.5638 | 1.1064 | -3.8457 | 1622.8776 | 179 | 9.0664
[5 empty] | 1.79 | 1.351e-39 | 0.01139 | 1 | 0.01199 | 0.98801 | 0 | 4 | 592.8062 | 25.005 | 2.5005 | 1.0791 | -3.7507 | 1626.8776 | 180 | 9.0382
[5 empty] | 1.8 | 8.194e-40 | 0.011109 | 1 | 0.011694 | 0.98831 | 0 | 4 | 592.9838 | 24.3876 | 2.4388 | 1.0524 | -3.6581 | 1630.8776 | 181 | 9.0104
[5 empty] | 1.81 | 4.9699e-40 | 0.010835 | 1 | 0.011405 | 0.9886 | 0 | 4 | 593.157 | 23.7855 | 2.3785 | 1.0264 | -3.5678 | 1634.8776 | 182 | 8.9828
[5 empty] | 1.82 | 3.0144e-40 | 0.010567 | 1 | 0.011123 | 0.98888 | 0 | 4 | 593.326 | 23.1982 | 2.3198 | 1.0011 | -3.4797 | 1638.8776 | 183 | 8.9556
[5 empty] | 1.83 | 1.8283e-40 | 0.010306 | 1 | 0.010849 | 0.98915 | 0 | 4 | 593.4908 | 22.6255 | 2.2625 | 0.97639 | -3.3938 | 1642.8776 | 184 | 8.9287
[5 empty] | 1.84 | 1.1089e-40 | 0.010052 | 1 | 0.010581 | 0.98942 | 0 | 4 | 593.6515 | 22.0668 | 2.2067 | 0.95228 | -3.31 | 1646.8776 | 185 | 8.902
[5 empty] | 1.85 | 6.7261e-41 | 0.0098037 | 1 | 0.01032 | 0.98968 | 0 | 4 | 593.8082 | 21.522 | 2.1522 | 0.92877 | -3.2283 | 1650.8776 | 186 | 8.8757
[5 empty] | 1.86 | 4.0796e-41 | 0.0095616 | 1 | 0.010065 | 0.98994 | 0 | 4 | 593.9611 | 20.9906 | 2.0991 | 0.90584 | -3.1486 | 1654.8776 | 187 | 8.8496
[5 empty] | 1.87 | 2.4744e-41 | 0.0093255 | 1 | 0.0098163 | 0.99018 | 0 | 4 | 594.1102 | 20.4724 | 2.0472 | 0.88347 | -3.0709 | 1658.8776 | 188 | 8.8238
[5 empty] | 1.88 | 1.5008e-41 | 0.0090953 | 1 | 0.009574 | 0.99043 | 0 | 4 | 594.2556 | 19.9669 | 1.9967 | 0.86166 | -2.995 | 1662.8776 | 189 | 8.7983
[5 empty] | 1.89 | 9.1027e-42 | 0.0088707 | 1 | 0.0093376 | 0.99066 | 0 | 4 | 594.3974 | 19.4739 | 1.9474 | 0.84038 | -2.9211 | 1666.8776 | 190 | 8.773
[5 empty] | 1.9 | 5.5211e-42 | 0.0086517 | 1 | 0.009107 | 0.99089 | 0 | 4 | 594.5358 | 18.9931 | 1.8993 | 0.81963 | -2.849 | 1670.8776 | 191 | 8.7481
[5 empty] | 1.91 | 3.3487e-42 | 0.0084381 | 1 | 0.0088822 | 0.99112 | 0 | 4 | 594.6707 | 18.5242 | 1.8524 | 0.7994 | -2.7786 | 1674.8776 | 192 | 8.7233
[5 empty] | 1.92 | 2.0311e-42 | 0.0082297 | 1 | 0.0086629 | 0.99134 | 0 | 4 | 594.8023 | 18.0668 | 1.8067 | 0.77966 | -2.71 | 1678.8776 | 193 | 8.6988
[5 empty] | 1.93 | 1.2319e-42 | 0.0080266 | 1 | 0.008449 | 0.99155 | 0 | 4 | 594.9306 | 17.6207 | 1.7621 | 0.76041 | -2.6431 | 1682.8776 | 194 | 8.6746
[5 empty] | 1.94 | 7.472e-43 | 0.0078284 | 1 | 0.0082404 | 0.99176 | 0 | 4 | 595.0558 | 17.1857 | 1.7186 | 0.74164 | -2.5778 | 1686.8776 | 195 | 8.6507
[5 empty] | 1.95 | 4.532e-43 | 0.0076351 | 1 | 0.0080369 | 0.99196 | 0 | 4 | 595.1778 | 16.7613 | 1.6761 | 0.72332 | -2.5142 | 1690.8776 | 196 | 8.6269
[5 empty] | 1.96 | 2.7488e-43 | 0.0074466 | 1 | 0.0078385 | 0.99216 | 0 | 4 | 595.2969 | 16.3475 | 1.6348 | 0.70547 | -2.4521 | 1694.8776 | 197 | 8.6034
[5 empty] | 1.97 | 1.6672e-43 | 0.0072627 | 1 | 0.007645 | 0.99236 | 0 | 4 | 595.413 | 15.9439 | 1.5944 | 0.68805 | -2.3916 | 1698.8776 | 198 | 8.5802
[5 empty] | 1.98 | 1.0112e-43 | 0.0070834 | 1 | 0.0074562 | 0.99254 | 0 | 4 | 595.5263 | 15.5502 | 1.555 | 0.67106 | -2.3325 | 1702.8776 | 199 | 8.5572
[5 empty] | 1.99 | 6.1334e-44 | 0.0069085 | 1 | 0.0072721 | 0.99273 | 0 | 4 | 595.6367 | 15.1663 | 1.5166 | 0.65449 | -2.2749 | 1706.8776 | 200 | 8.5344
[5 empty] | 2 | 3.7201e-44 | 0.0067379 | 1 | 0.0070926 | 0.99291 | 0 | 4 | 595.7445 | 14.7918 | 1.4792 | 0.63833 | -2.2188 | 1710.8776 | 201 | 8.5118
[5 empty] | 2.01 | 2.2563e-44 | 0.0065716 | 1 | 0.0069175 | 0.99308 | 0 | 4 | 595.8495 | 14.4266 | 1.4427 | 0.62257 | -2.164 | 1714.8776 | 202 | 8.4895
[5 empty] | 2.02 | 1.3685e-44 | 0.0064093 | 1 | 0.0067467 | 0.99325 | 0 | 4 | 595.952 | 14.0704 | 1.407 | 0.6072 | -2.1106 | 1718.8776 | 203 | 8.4674
[5 empty] | 2.03 | 8.3006e-45 | 0.0062511 | 1 | 0.0065801 | 0.99342 | 0 | 4 | 596.0519 | 13.723 | 1.3723 | 0.59221 | -2.0585 | 1722.8776 | 204 | 8.4455
[5 empty] | 2.04 | 5.0346e-45 | 0.0060967 | 1 | 0.0064176 | 0.99358 | 0 | 4 | 596.1494 | 13.3842 | 1.3384 | 0.57759 | -2.0076 | 1726.8776 | 205 | 8.4238
[5 empty] | 2.05 | 3.0536e-45 | 0.0059462 | 1 | 0.0062592 | 0.99374 | 0 | 4 | 596.2445 | 13.0537 | 1.3054 | 0.56333 | -1.9581 | 1730.8776 | 206 | 8.4023
[5 empty] | 2.06 | 1.8521e-45 | 0.0057994 | 1 | 0.0061046 | 0.9939 | 0 | 4 | 596.3372 | 12.7315 | 1.2731 | 0.54942 | -1.9097 | 1734.8776 | 207 | 8.3811
[5 empty] | 2.07 | 1.1234e-45 | 0.0056562 | 1 | 0.0059539 | 0.99405 | 0 | 4 | 596.4277 | 12.4171 | 1.2417 | 0.53585 | -1.8626 | 1738.8776 | 208 | 8.36
[5 empty] | 2.08 | 6.8136e-46 | 0.0055166 | 1 | 0.0058069 | 0.99419 | 0 | 4 | 596.5159 | 12.1105 | 1.2111 | 0.52262 | -1.8166 | 1742.8776 | 209 | 8.3391
[5 empty] | 2.09 | 4.1326e-46 | 0.0053804 | 1 | 0.0056635 | 0.99434 | 0 | 4 | 596.6019 | 11.8115 | 1.1812 | 0.50972 | -1.7717 | 1746.8776 | 210 | 8.3185
[5 empty] | 2.1 | 2.5066e-46 | 0.0052475 | 1 | 0.0055237 | 0.99448 | 0 | 4 | 596.6858 | 11.5199 | 1.152 | 0.49713 | -1.728 | 1750.8776 | 211 | 8.298
[5 empty] | 2.11 | 1.5203e-46 | 0.005118 | 1 | 0.0053873 | 0.99461 | 0 | 4 | 596.7676 | 11.2355 | 1.1235 | 0.48486 | -1.6853 | 1754.8776 | 212 | 8.2777
[5 empty] | 2.12 | 9.2211e-47 | 0.0049916 | 1 | 0.0052543 | 0.99475 | 0 | 4 | 596.8474 | 10.9581 | 1.0958 | 0.47289 | -1.6437 | 1758.8776 | 213 | 8.2576
[5 empty] | 2.13 | 5.5929e-47 | 0.0048684 | 1 | 0.0051246 | 0.99488 | 0 | 4 | 596.9253 | 10.6875 | 1.0688 | 0.46121 | -1.6031 | 1762.8776 | 214 | 8.2377
[5 empty] | 2.14 | 3.3923e-47 | 0.0047482 | 1 | 0.0049981 | 0.995 | 0 | 4 | 597.0012 | 10.4236 | 1.0424 | 0.44982 | -1.5635 | 1766.8776 | 215 | 8.218
[5 empty] | 2.15 | 2.0575e-47 | 0.0046309 | 1 | 0.0048747 | 0.99513 | 0 | 4 | 597.0752 | 10.1663 | 1.0166 | 0.43872 | -1.5249 | 1770.8776 | 216 | 8.1985
[5 empty] | 2.16 | 1.2479e-47 | 0.0045166 | 1 | 0.0047543 | 0.99525 | 0 | 4 | 597.1474 | 9.9153 | 0.99153 | 0.42789 | -1.4873 | 1774.8776 | 217 | 8.1792
[5 empty] | 2.17 | 7.5692e-48 | 0.0044051 | 1 | 0.0046369 | 0.99536 | 0 | 4 | 597.2179 | 9.6705 | 0.96705 | 0.41732 | -1.4506 | 1778.8776 | 218 | 8.16
[5 empty] | 2.18 | 4.5909e-48 | 0.0042963 | 1 | 0.0045224 | 0.99548 | 0 | 4 | 597.2865 | 9.4317 | 0.94317 | 0.40702 | -1.4148 | 1782.8776 | 219 | 8.141
[5 empty] | 2.19 | 2.7845e-48 | 0.0041902 | 1 | 0.0044108 | 0.99559 | 0 | 4 | 597.3535 | 9.1988 | 0.91988 | 0.39697 | -1.3798 | 1786.8776 | 220 | 8.1222
[5 empty] | 2.2 | 1.6889e-48 | 0.0040868 | 1 | 0.0043019 | 0.9957 | 0 | 4 | 597.4189 | 8.9717 | 0.89717 | 0.38717 | -1.3458 | 1790.8776 | 221 | 8.1035
[5 empty] | 2.21 | 1.0244e-48 | 0.0039859 | 1 | 0.0041957 | 0.9958 | 0 | 4 | 597.4826 | 8.7502 | 0.87502 | 0.37761 | -1.3125 | 1794.8776 | 222 | 8.085
[5 empty] | 2.22 | 6.2132e-49 | 0.0038875 | 1 | 0.0040921 | 0.99591 | 0 | 4 | 597.5448 | 8.5341 | 0.85341 | 0.36829 | -1.2801 | 1798.8776 | 223 | 8.0667
[5 empty] | 2.23 | 3.7685e-49 | 0.0037915 | 1 | 0.003991 | 0.99601 | 0 | 4 | 597.6054 | 8.3234 | 0.83234 | 0.35919 | -1.2485 | 1802.8776 | 224 | 8.0486
[5 empty] | 2.24 | 2.2857e-49 | 0.0036979 | 1 | 0.0038925 | 0.99611 | 0 | 4 | 597.6645 | 8.1179 | 0.81179 | 0.35032 | -1.2177 | 1806.8776 | 225 | 8.0306
[5 empty] | 2.25 | 1.3863e-49 | 0.0036066 | 1 | 0.0037964 | 0.9962 | 0 | 4 | 597.7222 | 7.9175 | 0.79175 | 0.34167 | -1.1876 | 1810.8776 | 226 | 8.0127
[5 empty] | 2.26 | 8.4086e-50 | 0.0035175 | 1 | 0.0037026 | 0.9963 | 0 | 4 | 597.7784 | 7.722 | 0.7722 | 0.33324 | -1.1583 | 1814.8776 | 227 | 7.9951
[5 empty] | 2.27 | 5.1001e-50 | 0.0034307 | 1 | 0.0036112 | 0.99639 | 0 | 4 | 597.8333 | 7.5314 | 0.75314 | 0.32501 | -1.1297 | 1818.8776 | 228 | 7.9775
[5 empty] | 2.28 | 3.0934e-50 | 0.003346 | 1 | 0.0035221 | 0.99648 | 0 | 4 | 597.8868 | 7.3454 | 0.73454 | 0.31699 | -1.1018 | 1822.8776 | 229 | 7.9602
[5 empty] | 2.29 | 1.8762e-50 | 0.0032634 | 1 | 0.0034351 | 0.99656 | 0 | 4 | 597.9389 | 7.164 | 0.7164 | 0.30916 | -1.0746 | 1826.8776 | 230 | 7.9429
[5 empty] | 2.3 | 1.138e-50 | 0.0031828 | 1 | 0.0033503 | 0.99665 | 0 | 4 | 597.9898 | 6.9872 | 0.69872 | 0.30153 | -1.0481 | 1830.8776 | 231 | 7.9259
[5 empty] | 2.31 | 6.9022e-51 | 0.0031042 | 1 | 0.0032676 | 0.99673 | 0 | 4 | 598.0395 | 6.8147 | 0.68147 | 0.29408 | -1.0222 | 1834.8776 | 232 | 7.909
[5 empty] | 2.32 | 4.1864e-51 | 0.0030276 | 1 | 0.0031869 | 0.99681 | 0 | 4 | 598.0879 | 6.6464 | 0.66464 | 0.28682 | -0.99696 | 1838.8776 | 233 | 7.8922
[5 empty] | 2.33 | 2.5392e-51 | 0.0029528 | 1 | 0.0031082 | 0.99689 | 0 | 4 | 598.1351 | 6.4823 | 0.64823 | 0.27974 | -0.97235 | 1842.8776 | 234 | 7.8755
[5 empty] | 2.34 | 1.5401e-51 | 0.0028799 | 1 | 0.0030315 | 0.99697 | 0 | 4 | 598.1811 | 6.3223 | 0.63223 | 0.27283 | -0.94834 | 1846.8776 | 235 | 7.8591
[5 empty] | 2.35 | 9.3411e-52 | 0.0028088 | 1 | 0.0029566 | 0.99704 | 0 | 4 | 598.226 | 6.1662 | 0.61662 | 0.2661 | -0.92492 | 1850.8776 | 236 | 7.8427
[5 empty] | 2.36 | 5.6657e-52 | 0.0027394 | 1 | 0.0028836 | 0.99712 | 0 | 4 | 598.2698 | 6.0139 | 0.60139 | 0.25953 | -0.90209 | 1854.8776 | 237 | 7.8265
[5 empty] | 2.37 | 3.4364e-52 | 0.0026718 | 1 | 0.0028124 | 0.99719 | 0 | 4 | 598.3125 | 5.8654 | 0.58654 | 0.25312 | -0.87981 | 1858.8776 | 238 | 7.8104
[5 empty] | 2.38 | 2.0843e-52 | 0.0026058 | 1 | 0.002743 | 0.99726 | 0 | 4 | 598.3542 | 5.7206 | 0.57206 | 0.24687 | -0.85809 | 1862.8776 | 239 | 7.7945
[5 empty] | 2.39 | 1.2642e-52 | 0.0025415 | 1 | 0.0026753 | 0.99732 | 0 | 4 | 598.3948 | 5.5794 | 0.55794 | 0.24077 | -0.83691 | 1866.8776 | 240 | 7.7787
[5 empty] | 2.4 | 7.6676e-53 | 0.0024788 | 1 | 0.0026092 | 0.99739 | 0 | 4 | 598.4345 | 5.4416 | 0.54416 | 0.23483 | -0.81624 | 1870.8776 | 241 | 7.763
[5 empty] | 2.41 | 4.6507e-53 | 0.0024176 | 1 | 0.0025448 | 0.99746 | 0 | 4 | 598.4731 | 5.3073 | 0.53073 | 0.22903 | -0.79609 | 1874.8776 | 242 | 7.7474
[5 empty] | 2.42 | 2.8208e-53 | 0.0023579 | 1 | 0.002482 | 0.99752 | 0 | 4 | 598.5108 | 5.1762 | 0.51762 | 0.22338 | -0.77643 | 1878.8776 | 243 | 7.732
[5 empty] | 2.43 | 1.7109e-53 | 0.0022996 | 1 | 0.0024207 | 0.99758 | 0 | 4 | 598.5476 | 5.0484 | 0.50484 | 0.21786 | -0.75726 | 1882.8776 | 244 | 7.7167
[5 empty] | 2.44 | 1.0377e-53 | 0.0022429 | 1 | 0.0023609 | 0.99764 | 0 | 4 | 598.5835 | 4.9238 | 0.49238 | 0.21248 | -0.73857 | 1886.8776 | 245 | 7.7015
[5 empty] | 2.45 | 6.294e-54 | 0.0021875 | 1 | 0.0023026 | 0.9977 | 0 | 4 | 598.6184 | 4.8022 | 0.48022 | 0.20724 | -0.72033 | 1890.8776 | 246 | 7.6865
[5 empty] | 2.46 | 3.8175e-54 | 0.0021335 | 1 | 0.0022458 | 0.99775 | 0 | 4 | 598.6525 | 4.6836 | 0.46836 | 0.20212 | -0.70255 | 1894.8776 | 247 | 7.6716
[5 empty] | 2.47 | 2.3154e-54 | 0.0020808 | 1 | 0.0021903 | 0.99781 | 0 | 4 | 598.6858 | 4.568 | 0.4568 | 0.19713 | -0.6852 | 1898.8776 | 248 | 7.6568
[5 empty] | 2.48 | 1.4044e-54 | 0.0020294 | 1 | 0.0021362 | 0.99786 | 0 | 4 | 598.7183 | 4.4552 | 0.44552 | 0.19226 | -0.66828 | 1902.8776 | 249 | 7.6421
[5 empty] | 2.49 | 8.518e-55 | 0.0019793 | 1 | 0.0020835 | 0.99792 | 0 | 4 | 598.7499 | 4.3452 | 0.43452 | 0.18751 | -0.65178 | 1906.8776 | 250 | 7.6275
[5 empty] | 2.5 | 5.1664e-55 | 0.0019305 | 1 | 0.0020321 | 0.99797 | 0 | 4 | 598.7808 | 4.2379 | 0.42379 | 0.18289 | -0.63569 | 1910.8776 | 251 | 7.6131
[5 empty] | 2.51 | 3.1336e-55 | 0.0018828 | 1 | 0.0019819 | 0.99802 | 0 | 4 | 598.8109 | 4.1333 | 0.41333 | 0.17837 | -0.61999 | 1914.8776 | 252 | 7.5987
[5 empty] | 2.52 | 1.9006e-55 | 0.0018363 | 1 | 0.001933 | 0.99807 | 0 | 4 | 598.8402 | 4.0312 | 0.40312 | 0.17397 | -0.60469 | 1918.8776 | 253 | 7.5845
[5 empty] | 2.53 | 1.1528e-55 | 0.001791 | 1 | 0.0018852 | 0.99811 | 0 | 4 | 598.8689 | 3.9317 | 0.39317 | 0.16967 | -0.58976 | 1922.8776 | 254 | 7.5704
[5 empty] | 2.54 | 6.992e-56 | 0.0017467 | 1 | 0.0018387 | 0.99816 | 0 | 4 | 598.8968 | 3.8346 | 0.38346 | 0.16548 | -0.5752 | 1926.8776 | 255 | 7.5564
[5 empty] | 2.55 | 4.2409e-56 | 0.0017036 | 1 | 0.0017933 | 0.99821 | 0 | 4 | 598.924 | 3.74 | 0.374 | 0.1614 | -0.56099 | 1930.8776 | 256 | 7.5425
[5 empty] | 2.56 | 2.5722e-56 | 0.0016616 | 1 | 0.001749 | 0.99825 | 0 | 4 | 598.9506 | 3.6476 | 0.36476 | 0.15741 | -0.54714 | 1934.8776 | 257 | 7.5287
[5 empty] | 2.57 | 1.5601e-56 | 0.0016205 | 1 | 0.0017058 | 0.99829 | 0 | 4 | 598.9765 | 3.5576 | 0.35576 | 0.15352 | -0.53363 | 1938.8776 | 258 | 7.515
[5 empty] | 2.58 | 9.4626e-57 | 0.0015805 | 1 | 0.0016637 | 0.99834 | 0 | 4 | 599.0018 | 3.4697 | 0.34697 | 0.14973 | -0.52046 | 1942.8776 | 259 | 7.5015
[5 empty] | 2.59 | 5.7394e-57 | 0.0015415 | 1 | 0.0016226 | 0.99838 | 0 | 4 | 599.0264 | 3.3841 | 0.33841 | 0.14604 | -0.50761 | 1946.8776 | 260 | 7.488
[5 empty] | 2.6 | 3.4811e-57 | 0.0015034 | 1 | 0.0015826 | 0.99842 | 0 | 4 | 599.0505 | 3.3005 | 0.33005 | 0.14243 | -0.49508 | 1950.8776 | 261 | 7.4746
[5 empty] | 2.61 | 2.1114e-57 | 0.0014663 | 1 | 0.0015435 | 0.99846 | 0 | 4 | 599.0739 | 3.219 | 0.3219 | 0.13891 | -0.48285 | 1954.8776 | 262 | 7.4614
[5 empty] | 2.62 | 1.2806e-57 | 0.0014301 | 1 | 0.0015054 | 0.99849 | 0 | 4 | 599.0968 | 3.1395 | 0.31395 | 0.13548 | -0.47093 | 1958.8776 | 263 | 7.4482
[5 empty] | 2.63 | 7.7674e-58 | 0.0013948 | 1 | 0.0014682 | 0.99853 | 0 | 4 | 599.1191 | 3.062 | 0.3062 | 0.13214 | -0.4593 | 1962.8776 | 264 | 7.4351
[5 empty] | 2.64 | 4.7112e-58 | 0.0013604 | 1 | 0.001432 | 0.99857 | 0 | 4 | 599.1408 | 2.9864 | 0.29864 | 0.12888 | -0.44796 | 1966.8776 | 265 | 7.4222
[5 empty] | 2.65 | 2.8575e-58 | 0.0013268 | 1 | 0.0013966 | 0.9986 | 0 | 4 | 599.162 | 2.9127 | 0.29127 | 0.12569 | -0.4369 | 1970.8776 | 266 | 7.4093
[5 empty] | 2.66 | 1.7331e-58 | 0.001294 | 1 | 0.0013621 | 0.99864 | 0 | 4 | 599.1827 | 2.8408 | 0.28408 | 0.12259 | -0.42612 | 1974.8776 | 267 | 7.3965
[5 empty] | 2.67 | 1.0512e-58 | 0.0012621 | 1 | 0.0013285 | 0.99867 | 0 | 4 | 599.2029 | 2.7706 | 0.27706 | 0.11956 | -0.41559 | 1978.8776 | 268 | 7.3839
[5 empty] | 2.68 | 6.3759e-59 | 0.0012309 | 1 | 0.0012957 | 0.9987 | 0 | 4 | 599.2226 | 2.7022 | 0.27022 | 0.11661 | -0.40533 | 1982.8776 | 269 | 7.3713
[5 empty] | 2.69 | 3.8672e-59 | 0.0012005 | 1 | 0.0012637 | 0.99874 | 0 | 4 | 599.2418 | 2.6355 | 0.26355 | 0.11373 | -0.39533 | 1986.8776 | 270 | 7.3588
[5 empty] | 2.7 | 2.3456e-59 | 0.0011709 | 1 | 0.0012325 | 0.99877 | 0 | 4 | 599.2605 | 2.5704 | 0.25704 | 0.11093 | -0.38557 | 1990.8776 | 271 | 7.3464
[5 empty] | 2.71 | 1.4226e-59 | 0.001142 | 1 | 0.0012021 | 0.9988 | 0 | 4 | 599.2788 | 2.507 | 0.2507 | 0.10819 | -0.37605 | 1994.8776 | 272 | 7.3341
[5 empty] | 2.72 | 8.6288e-60 | 0.0011138 | 1 | 0.0011724 | 0.99883 | 0 | 4 | 599.2966 | 2.4451 | 0.24451 | 0.10552 | -0.36676 | 1998.8776 | 273 | 7.3219
[5 empty] | 2.73 | 5.2336e-60 | 0.0010863 | 1 | 0.0011434 | 0.99886 | 0 | 4 | 599.3139 | 2.3847 | 0.23847 | 0.10291 | -0.35771 | 2002.8776 | 274 | 7.3098
[5 empty] | 2.74 | 3.1744e-60 | 0.0010595 | 1 | 0.0011152 | 0.99888 | 0 | 4 | 599.3309 | 2.3258 | 0.23258 | 0.10037 | -0.34887 | 2006.8776 | 275 | 7.2977
[5 empty] | 2.75 | 1.9253e-60 | 0.0010333 | 1 | 0.0010877 | 0.99891 | 0 | 4 | 599.3474 | 2.2684 | 0.22684 | 0.097891 | -0.34026 | 2010.8776 | 276 | 7.2858
[5 empty] | 2.76 | 1.1678e-60 | 0.0010078 | 1 | 0.0010608 | 0.99894 | 0 | 4 | 599.3635 | 2.2124 | 0.22124 | 0.095474 | -0.33186 | 2014.8776 | 277 | 7.2739
[5 empty] | 2.77 | 7.083e-61 | 0.0009829 | 1 | 0.0010346 | 0.99897 | 0 | 4 | 599.3792 | 2.1578 | 0.21578 | 0.093117 | -0.32367 | 2018.8776 | 278 | 7.2621
[5 empty] | 2.78 | 4.296e-61 | 0.00095864 | 1 | 0.0010091 | 0.99899 | 0 | 4 | 599.3945 | 2.1045 | 0.21045 | 0.090818 | -0.31567 | 2022.8776 | 279 | 7.2505
[5 empty] | 2.79 | 2.6057e-61 | 0.00093497 | 1 | 0.00098418 | 0.99902 | 0 | 4 | 599.4095 | 2.0525 | 0.20525 | 0.088576 | -0.30788 | 2026.8776 | 280 | 7.2388
[5 empty] | 2.8 | 1.5804e-61 | 0.00091188 | 1 | 0.00095988 | 0.99904 | 0 | 4 | 599.4241 | 2.0019 | 0.20019 | 0.086389 | -0.30028 | 2030.8776 | 281 | 7.2273
[5 empty] | 2.81 | 9.5857e-62 | 0.00088937 | 1 | 0.00093618 | 0.99906 | 0 | 4 | 599.4383 | 1.9524 | 0.19524 | 0.084256 | -0.29286 | 2034.8776 | 282 | 7.2159
[5 empty] | 2.82 | 5.814e-62 | 0.00086741 | 1 | 0.00091306 | 0.99909 | 0 | 4 | 599.4522 | 1.9042 | 0.19042 | 0.082176 | -0.28563 | 2038.8776 | 283 | 7.2045
[5 empty] | 2.83 | 3.5264e-62 | 0.00084599 | 1 | 0.00089052 | 0.99911 | 0 | 4 | 599.4657 | 1.8572 | 0.18572 | 0.080147 | -0.27858 | 2042.8776 | 284 | 7.1932
[5 empty] | 2.84 | 2.1389e-62 | 0.0008251 | 1 | 0.00086853 | 0.99913 | 0 | 4 | 599.4789 | 1.8114 | 0.18114 | 0.078168 | -0.2717 | 2046.8776 | 285 | 7.182
[5 empty] | 2.85 | 1.2973e-62 | 0.00080473 | 1 | 0.00084709 | 0.99915 | 0 | 4 | 599.4917 | 1.7666 | 0.17666 | 0.076238 | -0.26499 | 2050.8776 | 286 | 7.1709
[5 empty] | 2.86 | 7.8684e-63 | 0.00078486 | 1 | 0.00082617 | 0.99917 | 0 | 4 | 599.5043 | 1.723 | 0.1723 | 0.074356 | -0.25845 | 2054.8776 | 287 | 7.1599
[5 empty] | 2.87 | 4.7725e-63 | 0.00076549 | 1 | 0.00080577 | 0.99919 | 0 | 4 | 599.5165 | 1.6805 | 0.16805 | 0.07252 | -0.25207 | 2058.8776 | 288 | 7.1489
[5 empty] | 2.88 | 2.8946e-63 | 0.00074659 | 1 | 0.00078588 | 0.99921 | 0 | 4 | 599.5285 | 1.639 | 0.1639 | 0.070729 | -0.24585 | 2062.8776 | 289 | 7.138
[5 empty] | 2.89 | 1.7557e-63 | 0.00072815 | 1 | 0.00076648 | 0.99923 | 0 | 4 | 599.5401 | 1.5985 | 0.15985 | 0.068983 | -0.23978 | 2066.8776 | 290 | 7.1272
[5 empty] | 2.9 | 1.0649e-63 | 0.00071017 | 1 | 0.00074755 | 0.99925 | 0 | 4 | 599.5515 | 1.559 | 0.1559 | 0.06728 | -0.23386 | 2070.8776 | 291 | 7.1164
[5 empty] | 2.91 | 6.4588e-64 | 0.00069264 | 1 | 0.00072909 | 0.99927 | 0 | 4 | 599.5625 | 1.5206 | 0.15206 | 0.065619 | -0.22808 | 2074.8776 | 292 | 7.1057
[5 empty] | 2.92 | 3.9175e-64 | 0.00067554 | 1 | 0.00071109 | 0.99929 | 0 | 4 | 599.5733 | 1.483 | 0.1483 | 0.063998 | -0.22245 | 2078.8776 | 293 | 7.0951
[5 empty] | 2.93 | 2.3761e-64 | 0.00065886 | 1 | 0.00069354 | 0.99931 | 0 | 4 | 599.5839 | 1.4464 | 0.14464 | 0.062418 | -0.21696 | 2082.8776 | 294 | 7.0846
[5 empty] | 2.94 | 1.4412e-64 | 0.00064259 | 1 | 0.00067641 | 0.99932 | 0 | 4 | 599.5942 | 1.4107 | 0.14107 | 0.060877 | -0.2116 | 2086.8776 | 295 | 7.0742
[5 empty] | 2.95 | 8.7411e-65 | 0.00062673 | 1 | 0.00065971 | 0.99934 | 0 | 4 | 599.6042 | 1.3759 | 0.13759 | 0.059374 | -0.20638 | 2090.8776 | 296 | 7.0638
[5 empty] | 2.96 | 5.3017e-65 | 0.00061125 | 1 | 0.00064342 | 0.99936 | 0 | 4 | 599.6139 | 1.3419 | 0.13419 | 0.057908 | -0.20128 | 2094.8776 | 297 | 7.0535
[5 empty] | 2.97 | 3.2157e-65 | 0.00059616 | 1 | 0.00062754 | 0.99937 | 0 | 4 | 599.6235 | 1.3088 | 0.13088 | 0.056478 | -0.19631 | 2098.8776 | 298 | 7.0432
[5 empty] | 2.98 | 1.9504e-65 | 0.00058144 | 1 | 0.00061204 | 0.99939 | 0 | 4 | 599.6328 | 1.2764 | 0.12764 | 0.055084 | -0.19147 | 2102.8776 | 299 | 7.033
[5 empty] | 2.99 | 1.183e-65 | 0.00056709 | 1 | 0.00059693 | 0.9994 | 0 | 4 | 599.6418 | 1.2449 | 0.12449 | 0.053724 | -0.18674 | 2106.8776 | 300 | 7.0229
[5 empty] | 3 | 7.1751e-66 | 0.00055308 | 1 | 0.00058219 | 0.99942 | 0 | 4 | 599.6507 | 1.2142 | 0.12142 | 0.052397 | -0.18213 | 2110.8776 | 301 | 7.0129
Menu
time
time pk flow
P
time
P rise 
PEEP
flow at
time
flow cycle L/min
P
time
P rise 
PEEP
PEEP

## Eq
Equation of Motion for the Respiratory System
| Pvent | = pressure (above PEEP) generated by ventiltor during inspiration (varies with time)
| Pmus | = pressure generated by patient's ventilatory muscles (varies with time)
| E | = respiratory system elastance, DP/DV, (constant)
| R | = respiratory system resistenc, (DP/DV', (constant)
| V | = volume (varies with time)
| V' | = flow (varies with time)
Equations Derived from the Equation of Motion
Menu
Volume Control
Pressure Control

## Gloss
STANDARDIZED VOCABULARY FOR MECHANICAL VENTILATION
Ó Robert L. Chatburn 2015
https://app.box.com/s/sbnajwyhdo53yf6wqcv50orxypumb16j
Key Definitions for Taxonomy | Maxim
BREATH INSPIRATORY TIME EXPIRATORY TIME TOTAL CYCLE TIME I:E RATIO DUTY CYCLE TIDAL VOLUME | 1. A breath is one cycle of positive flow (inspiration) and negative flow (expiration) defined in terms of the flow-time curve
ASSISTED BREATH WORK OF BREATHING TRANSRESPIRATORY PRESSURE LOADED BREATH | 2. A breath is assisted if the ventilator provides some or all of the work of breathing
EQUATION OF MOTION VOLUME CONTROL PRESSURE CONTROL | 3. A ventilator assists breathing using either pressure control or volume control based on the equation of motion for the respiratory system
TRIGGER TRIGGER VARIABLE CYCLE CYCLE VARIABLE | 4. Breaths are classified according to the criteria that trigger (start) and cycle (stop) inspiration
PATIENT TRIGGERING PATIENT CYCLING MACHINE TRIGGERING MACHINE CYCLING TRIGGER WINDOW SYNCHRONIZATION WINDOW PRESSURE TRIGGERING PRESSURE CYCLING VOLUME TRIGGERING VOLUME CYCLING FLOW TRIGGERING FLOW CYCLING | 5. Trigger and cycle events can be either patient-initiated or ventilator-initiated
MANDATORY BRATH SPONTANEOUS BREATH | 6. Breaths are classified as spontaneous or mandatory based on both the trigger and cycle events
CONTINUOUS MANDATORY VENTILATION INTERMITTENT MANDATORY VENTILATION CONTINUOUS SPONTANEOUS VENTILATION | 7. Ventilators deliver 3 basic breath sequences: CMV, IMV, and continuous spontaneous ventilation CSV
VC-CMV VC-IMV PC-CMV PC-IMV PC-CSV TC-IMV | 8. Ventilator deliver 5 basic ventilatory patterns: VC-CMV, VC-IMV, PC-CMV, PC-IMV, PC-CSV
TARGET TARGETING SCHEME SET-POINT TARGETING SCHEME DUAL TARGETING SCHEME BIOVARIABLE TARGETING SCHEME SERVO TARGETING SCHEME ADAPTIVE TARGETING SCHEME OPTIMAL TARGETING SCHEME INTELLIGENT TARGETING SCHEME | 9. Within each ventilatory pattern, there are several types that can be distinguished by their targeting schemes (set-point, dual, bio-variable, servo, adaptive, optimal, and intelligent)
MODE OF VENTILATION TAXONOMY MODE VARIETY | 10. A mode of ventilation is classified according to its control variable, breath sequence, and targeting scheme(s)
ACTIVE EXHALATION VALVE | A mechanism for holding pressure in the breathing circuit by delivering the flow required to allow the patient breathe spontaneously. This feature is especially prominent in modes like Airway Pressure Release Ventilation that are intended to allow unrestricted spontaneous breathing during a prolonged mandatory (i.e., time triggered and time cycled) pressure controlled breath.
ACTIVE VENTILATION | Inspiratory muscles contribute to the generation of inspiratory flow. Active ventilation may be assisted by the ventilator (ie, the ventilator does some of the work of breathing) or unassisted. See ASSISTED BREATH
ADAPTIVE TARGETING SCHEME | A control system that allows the ventilator to automatically set some (or conceivably all) of the targets between breaths to achieve other preset targets. One common example is adaptive pressure targeting (e.g., Pressure Regulated Volume Control mode on the Maquet Servo-i ventilator) where a static inspiratory pressure is targeted within a breath (i.e., pressure controlled inspiration) but this target is automatically adjusted by the ventilator between breaths to achieve an operator set tidal volume target (aka, volume-targeted pressure control). Another common example is Mandatory Minute Ventilation; the ventilator monitors the total minute ventilation (sum of mandatory and spontaneous breath volumes) and if it falls below the preset threshold for minimum minute ventilation, the mandatory breath frequency is increased to compensate. See MANDATORY MINUTE VENTILATION
AIRWAY PRESSURE | The pressure at the airway opening measured relative to atmospheric pressure during mechanical ventilation.
AIRWAY PRESSURE RELEASE VENTILATION (APRV) | A form of pressure control intermittent mandatory ventilation that is designed to allow unrestricted spontaneous breathing throughout the breath cycle. APRV is applied using I:E ratios much greater than 1:1 and usually relying on short expiratory times and gas trapping to maintain end expiratory lung volume rather than a preset PEEP. This is in contrast to Bilevel Positive Airway Pressure (BIPAP) which is also pressure control intermittent mandatory ventilation but with I:E ratios closer to 1:1, expiratory times that do not create significant gas tapping and preset PEEP levels above zero.
ASYNCHRONY | Asynchrony during mechanical ventilation is defined as the absence of one of the two reference waveforms generated by either the patient or the ventilator. See SYNCHRONY, DYSSYNCHRONY
ASSISTED BREATH | A breath during which all or part of inspiratory (or expiratory) flow is generated by the ventilator doing work on the patient. In simple terms, if the airway pressure rises above end expiratory pressure during inspiration, the inspiration is assisted (as in the Pressure Support mode). It is also possible to assist expiration by dropping airway pressure below end expiratory pressure (such as Automatic Tube Compensation on the Dräger Evita 4 ventilator). In contrast, spontaneous breaths during CPAP are unassisted because the ventilator attempts to maintain a constant airway pressure during inspiration.
ATELECTRAUMA | Ventilator induced lung injury caused by cyclic opening and closing of airway and pulmonary units due to insufficient PEEP and subsequent atelectasis on expiration. See VOLUTRAUMA
AUTO-PEEP | The positive difference between end-expiratory alveolar pressure (total or intrinsic PEEP) and the end-expiratory airway pressure (set or extrinsic PEEP). When autoPEEP exists, a positive pressure difference drives flow throughout exhalation until the subsequent breath interrupts deflation. AutoPEEP is caused when expiratory time (either set by the patient’s brain or a ventilator) is short relative to the expiratory time constant of the respiratory system (possibly including the expiratory resistance of the breathing circuit).
AUTOMATIC TUBE COMPENSATION | A feature that allows the operator to enter the size of the patient’s endotracheal tube and have the ventilator calculate the tube’s resistance and then generate just enough pressure (in proportion to inspiratory or expiratory flow) to compensate for the added resistive load. See SERVO CONTROL
AUTO-TRIGGER | A condition in which the ventilator repeatedly triggers itself because the sensitivity is set too high(sometimes called “autocycling”). The ventilator may auto-trigger due to a leak in the system dropping airway pressure below a pressure trigger threshold. When sensitivity is set too high, even the heartbeat can cause inadvertent triggering. Auto-triggering is a form of patient-ventilator asynchrony.
BAROTRAUMA | An obsolete term (from the 1970s) indicating ventilator induced lung damage from excessive inspiratory pressure. Current understanding is that the causative agent is not airway pressure but transalveolar or transpulmonary pressure difference. See VOLUTRAUMA
BIOTRAUMA | Ventilator induced lung damage caused by the release of proinflammatory cytokines and the recruitment of white cells, with lung inflammation occurring as a result applied stess and the resulting lung strain
BIO-VARIABLE TARGETING SCHEME | A control system that allows the ventilator to automatically set the inspiratory pressure or tidal volume randomly to mimic the variability observed during normal breathing. Currently this “biologically variable” targeting scheme is only available in one mode, Variable Pressure Support, on the Dräger V500 ventilator. The operator sets a target inspiratory pressure and a percent variability from 0% to 100%. A setting 0% means the preset inspiratory pressure will be delivered for every breath. A 100 % variability setting means that the actual inspiratory pressure varies randomly from PEEP/CPAP level to double the preset pressure support level.
BLOWER | A blower is a machine for generating relatively large flows of gas as the direct ventilator output with a relatively moderate increase of pressure (e.g., 2 psi). Blowers are used on home care and transport ventilators. See COMPRESSOR
BREATH | A positive change in airway flow (inspiration) paired with a negative change in airway flow (expiration), associated with ventilation of the lungs. This definition excludes flow changes caused by hiccups or cardiogenic oscillations. However, it allows the superimposition of, for example, a spontaneous breath on a mandatory breath or vice versa. The flows are paired by size, not necessarily by timing. For example, in Airway Pressure Release Ventilation there is a large mandatory inspiration (transition from low pressure to high pressure) possibly followed by a few small spontaneous inspirations and expirations, followed finally by a large expiration (transition from high pressure to low pressure). These comprise several small spontaneous breaths superimposed on one large mandatory breath. In contrast, during High Frequency Oscillatory Ventilation, small mandatory breaths are superimposed on larger spontaneous breaths.
BREATH SEQUENCE | A particular pattern of spontaneous and/or mandatory breaths. The 3 possible breath sequences are: continuous mandatory ventilation, (CMV), intermittent mandatory ventilation (IMV), and continuous spontaneous ventilation (CSV).
COMPLIANCE | defined as the ratio of the change in volume to the associated change in the pressure difference across the system. A mechanical property of a structure such as the respiratory system; a parameter of a lung model, or setting of a lung simulator. Compliance is the reciprocal of elastance.
COMPRESSOR | A compressor is a machine for moving a relatively low flow of gas to a storage container at a higher level of pressure (e.g, 20 psi). Compressors are generally found on intensive care ventilators whereas blowers are used on home care and transport ventilators. Compressors are typically larger and consume more electrical power than blowers, hence the use of the latter on small, portable devices. (see blower)
CMV | See CONTINUOUS MANDATORY VENTILAITON
CONTINUOUS MANDATORY VENTILATION | Commonly known as “Assist/Control”; CMV is a breath sequence for which spontaneous breaths are not permitted between mandatory breaths because every patient trigger signal in the trigger window produces a machine cycled inspiration (ie, a mandatory breath). Machine triggered mandatory breaths may be delivered at a preset rate. Therefore, in contrast to IMV, the mandatory breath frequency may be higher than the set frequency but never below it. In some pressure controlled modes on ventilators with an active exhalation valve, spontaneous breaths may occur during mandatory breaths, but the defining characteristic of CMV is that spontaneous breaths are not permitted between mandatory breaths. See MANDATORY BREATH, INTERMITTENT MANDATORY VENTILATION, TRIGGER WINDOW
CONTINUOUS SPONTANEOUS VENTILATION | A breath sequence for which all breaths are spontaneous.
CONTROL VARIABLE | The variable (ie, pressure or volume in the equation of motion) that the ventilator uses as a feedback signal to manipulate inspiration. For simple set-point targeting, the control variable can be identified as follows: If the peak inspiratory pressure remains constant as the load experienced by the ventilator changes, then the control variable is pressure. If the peak pressure changes as the load changes but tidal volume remains constant, then the control variable is volume. Volume control implies flow control and vice versa, but it is possible to distinguish the two on the basis of which signal is used for feedback control. Some modes (eg, High Frequency Oscillation and Intrapulmonary Percussive Ventilation) do not maintain either constant peak pressure or tidal volume by any feedback mechanism. However, in both these cases there is open loop control of pressure with the operator making the adjustments. Hence we can still classify them as pressure control modes, its just that the "control" is very crude.
CPAP | Continuous (ie, constant, unchanging) positive airway pressure; the set or measured mean value of transrespiratory system pressure during unassisted breathing or between assisted breaths. While this term is sometimes used synonymously for PEEP, historically, PEEP came first. PEEP mechanisms originally required the patient to drop transrespiratory system pressure to below atmospheric pressure to inhale, imposing a load and causing an increased work of breathing. CPAP mechanisms were developed so that the patient only had to drop pressure below the set CPAP level to inhale, thus decreasing the imposed load. See PEEP.
CSV | See CONTINUOUS SPONTANEOUS VENTILATION, SPONTANEOUS BREATH
CYCLE | To end the inspiratory time (and begin expiratory flow)
CYCLE VARIABLE | The variable (usually pressure, volume, flow, or time) that is used to end inspiratory time (and begin expiratory flow).
DRIVING PRESSURE | Driving pressure (also called tidal pressure) is most often defined as tidal volume divided by compliance, equivalent to plateau pressure minus PEEP. Newton's 3rd law of motion says that for every action, there is an equal and opposite reaction. Driving pressure can be understood within this framework as existing in two aspects (depending on the frame of reference) according to the equation of motion for the respiratory system: During mechanical ventilation, from the point of view of the lungs and chest wall, there is the experience of an active driving force attempting to inflate the lungs. Total driving pressure has two components, dynamic and static. Dynamic driving pressure is the inflating force (generated by the ventilator) that is required to drive flow through the airways. During inflation with preset flow, dynamic driving pressure is measured as airway pressure above end expiratory pressure (usually PEEP) at any moment in the inspiratory time. Static driving pressure is the force required to maintain the inflation volume when flow is zero throughout the system, and is measured as plateau pressure minus total PEEP (ie, from end inspiratory pause minus end expiratory pause pressures). From the point of view of the ventilator, there is a reactive force opposing inflation (the right hand side of the equation of motion). The dynamic component of this force is calculated as resistance times flow (resistive load). The static component is calculated as elastance times volume (elastic load) plus autoPEEP. See EQUATION OF MOTION
DUAL TARGETING SCHEME | A control system that allows the ventilator to switch between volume control and pressure control during a single inspiration. Dual targeting is a more advanced version of set-point targeting. It gives the ventilator the decision of whether the breath will be volume or pressure controlled according to the operator set priorities. The breath may start out in pressure control and automatically switch to volume control, as in the Bird “VAPS” mode or, the reverse, as in the Dräger “Pressure Limited” mode feature. The Maquet Servo-i ventilator has a mode called “Volume Control” and the operator presets both inspiratory time and tidal volume as would be expected with any conventional volume control mode. However, if the patient makes an inspiratory effort that decreases inspiratory pressure by 3 cm H2O, the ventilator switches to pressure control and, if the effort lasts long enough, flow cycles the breath. Indeed, if the tidal volume and inspiratory time are set relatively low and the inspiratory effort is relatively large, the resultant breath delivery is indistinguishable from Pressure Support. As a result, the tidal volume may be much larger than the expected, preset value. This highlights the need to understand dual targeting. Because both pressure and volume are the control variables during dual targeting, we identify the control variable as the one with which the breath initiates. This is because the alternate control variable may never be implemented during the breath, depending on the other factors in the targeting scheme.
DUTY CYCLE | The ratio of inspiratory time to total cycle time, usually expressed as a percent.
DYNAMIC CHARACTERISTIC | A term coined to describe a metric calculated as tidal volume divided by (peak inspiratory pressure minus plateau pressure) for a volume controlled breath with constant flow. This metric is often confused with dynamic compliance. Dynamic characteristic is always less than dynamic compliance due to the resistive pressure drop it contains (ie, peak inspiratory pressure has components due to both compliance and resistance). See STATIC COMPLIANCE, DYNAMIC COMPLIANCE. EQUATION OF MOTION.
DYNAMIC COMPLIANCE | The slope of the pressure-volume curve drawn between two points of zero flow (eg, at the start and end of inspiration). Dynamic compliance is less than static compliance for a system comprised of different time constants because some of the pressure is due to pendelluft between lung units. See STATIC COMPLIANCE, DYNAMIC CHARACTERISTIC
DYNAMIC HYPERINFLATION | The increase in lung volume that occurs whenever insufficient exhalation time prevents the respiratory system from returning to its normal resting end-expiratory equilibrium volume between breath cycles; the increase in lung volume associated with autoPEEP. Inappropriate operator set expiratory time may lead to dynamic hyperinflation, inability of the patient to trigger breaths, and an increased work of breathing.
DYSSYNCHRONY | Dyssynchrony during mechanical ventilation is defined as a no-zero phase difference between the pressure waveform representing patient inspiratory effort, Pmus, and the pressure waveform generated by the ventilator during an assisted breath. See SYNCHRONY, ASYNCHRONY
ELASTANCE | Defined as the ratio of the change in the pressure difference across the system to the associated change in volume. A mechanical property of a structure such as the respiratory system; a parameter of a lung model, or setting of a lung simulator. Elastance is the reciprocal of compliance.
ELASTIC LOAD | The pressure difference applied across a system (e.g., a container) that sustains the system's volume relative to some reference volume, and/or the amount of its compressible contents relative to some reference amount. For a linear system: elastance ´ volume, or, volume/compliance; for a container, the overall effective elastance (compliance) includes the elastances (compliances) of its structural components and the compressibility of the fluid (gas or liquid) within it.
EQUATION OF MOTION FOR THE RESPIRATORY SYSTEM | A mathematical relation among pressure difference, volume, and flow (as variable functions of time) that describes the physical interaction between the ventilator and the respiratory system. The simplest form is a differential equation with constant coefficients describing the respiratory system as a single deformable compartment including the lungs and chest wall connected in series to a single flow conducting tube. P(t) = Elastance x Volume(t) + Resistance x Flow(t) In this equation, P is pressure difference across the system (eg, transrespiratory pressure or transpulmonary pressure). Pressure difference, volume, and flow are all functions of time (t). The left hand side may be expanded to distinguish pressure generated by the ventilator (Pvent) from the pressure generated by the inspiratory muscles (Pmus). The right hand side may be expanded to account for autoPEEP: Pvent(t) + Pmus(t) = Elastance x Volume(t) + Resistance x Flow(t) + autoPEEP See ELASTIC LOAD, RESISTIVE LOAD
EXPIRATORY FLOW TIME | The period from the start of expiratory flow to the instant when expiratory flow stops, usually expressed in seconds. By convention, expiratory flow is in the negative direction (below zero) in graphs.
EXPIRATORY PAUSE TIME | The period from cessation of expiratory flow to the start of inspiratory flow.
EXPIRATORY PRESSURE | General term for the pressure at the patient connection during the expiratory phase of mechanical ventilation. See EXPIRATORY PRESSURE TARGET
EXPIRATORY PRESSURE TARGET | Setting for maximum expiratory pressure during mechanical ventilation; also called PEEP. See EXPIRATORY PRESSURE, PEEP
EXPIRATORY TIME | The period from the start of expiratory flow to the start of inspiratory flow, usually expressed in seconds. Expiratory time equals expiratory flow time plus expiratory pause time.
FEEDBACK CONTROL | Closed loop control is accomplished by using the output as a signal that is fed back (compared) to the operator-set input such that the difference between the two is used to drive the system toward the desired output (ie, negative feedback control). For example, pressure controlled modes use airway pressure as the feedback signal to manipulate gas flow from the ventilator to maintain an inspiratory pressure set-point.
FIO2 | Fraction of inspired oxygen; concentration of oxygen (usually expressed as a decimal value but also in percent) in a mixture of inspired gas.
FLOW TRIGGERING | The starting of inspiratory flow due to a patient inspiratory effort that generates inspiratory flow above a preset threshold. Flow triggering is a form of patient triggering. See SENSITIVITY
FLOW TARGET | A preset value for inspiratory flow, usually specified as a peak flow and perhaps also as a particular waveform. See TARGET
FLOW CYCLING | The ending of inspiratory time due to inspiratory flow decay below a preset threshold. Flow cycling is a form of patient cycling.
FULL VENTILATORY SUPPORT | The ventilator provides all of the work of breathing; muscle pressure in the equation of motion is zero; achievable only with sedation and paralysis of patient. See PARTIAL VENTILATORY SUPPORT
IMV | See INTERMITTENT MANDATORY VENTILATION
I:E | The ratio of inspiratory time to expiratory time. For example, an I:E ratio of 1:2 is equivalent to 1/2 or 0.5. Hence an I:E ratio of 1:2 (0.5) is larger than a ratio of 1:4 (0.25).
INSPIRATORY FLOW | The flow into the airway opening during the inspiratory time. By convention, inspiratory flow is in the positive direction (above zero) in graphs.
INSPIRATORY FLOW TIME | The period from the start of inspiratory flow (into the airway opening) to the cessation of inspiratory flow, usually expressed in seconds.
INSPIRATORY HOLD (PAUSE) | An intentional maneuver during mechanical ventilation whereby the airway opening is occluded for a short period (inspiratory hold time) after an assisted inspiration. This maneuver is used to assess static respiratory system mechanics and also to increase mean airway pressure during volume control ventilation in an attempt to improve gas exchange.
INSPIRATORY HOLD (PAUSE) TIME | The period from the cessation of inspiratory flow (into the airway opening) to the start of expiratory flow during mechanical ventilation, during an inspiratory hold, usually expressed in seconds.
INSPIRATORY PRESSURE | General term for the pressure at the patient connection during the inspiratory phase of mechanical ventilation. See INSPIRATORY PRESSURE TARGET
INSPIRATORY PRESSURE TARGET | Setting for maximum inspiratory pressure during pressure control ventilation. The inspiratory pressure target may set relative to PEEP or to atmospheric pressure. See INSPIRATORY PRESSURE
INSPIRATORY PRESSURE TIME | The period during a pressure control breath equivalent to inspiratory time minus the sum of pressure rise time and inspiratory hold time
INSPIRATORY TIME | The period from the start of inspiratory flow to the start of expiratory flow. Alternative definitions for inspiratory time can be created that are relative to either the patient or the ventilator. Patient, or neural inspiratory time is defined as the duration of the neural signal to the diaphragm causing it to contract (eg, using the EAdi signal from a catheter designed for Neurally Adjusted Ventilatory Assist). Neural inspiratory time can be modeled as the duration of a positive change in muscle pressure on a breathing simulator. Ventilator inspiratory time is defined in two different ways based on ventilator settings: (1) in volume control modes, ventilator inspiratory time is the period from the start of inspiratory flow to the start of expiratory flow and equal to inspiratory flow time (tidal volume/flow) plus inspiratory pause time (if any), (2) in pressure control modes, ventilator inspiratory time for mandatory breaths is generally preset but for spontaneous breaths it is the period from the start of pressure rise above PEEP (ventilator trigger event) to point when inspiratory flow meets the cycle threshold (ventilator cycle event).
INTELLIGENT TARGETING SCHEME | A ventilator control system that uses artificial intelligence programs such as fuzzy logic, rule based expert systems, and artificial neural networks. Examples include the rule based system used by SmartCare (Dräger Evita XL ventilator) and IntelliVent-ASV (Hamilton S1 ventilator).
INTERMITTENT MANDATORY VENTILATION | Breath sequence for which spontaneous breaths are permitted between mandatory breaths. If a synchronization window is used to coordinate flow with inspiratory effort, it is called synchronized IMV (SIMV). Three forms of IMV are: (1) Mandatory breaths are always delivered at the set frequency; (2) Mandatory breaths are delivered only when the spontaneous breath frequency falls below the set frequency; (3) Mandatory breaths are delivered only when the spontaneous minute ventilation (ie, product of spontaneous breath frequency and spontaneous breath tidal volume) drops below a preset or computed threshold (aka Mandatory Minute Ventilation). In contrast to CMV, with IMV the mandatory breath frequency can never be higher than the set rate but it may be lower. See MACHINE TRIGGERING, PATIENT TRIGGERING, SYNCHRONIZATION WINDOW, TRIGGER WINDOW, CONTINUOUS M ANDATORY VENTILATION
LOAD | The pressure resisting the force applied by the ventilator to assist inspiration. See ELASTIC LOAD, RESISTIVE LOAD
LOADED BREATH | A breath during which all or part of inspiratory (or expiratory) flow is generated by the patient doing work on the ventilator. In simple terms, if the airway pressure falls below end expiratory pressure during inspiration, the inspiration is loaded. If pressure rises above baseline on expiration, then expiration is loaded.
MACHINCE CYCLING | Ending inspiratory time independent of signals representing the patient determined components of the equation of motion (Pmus, elastance, or resistance). Common examples are cycling due to a preset tidal volume or inspiratory time. If a patient signal (indicating expiration) occurs during an inspiratory time synchronization window, inspiration stops and is defined as a machine cycled event that ends a mandatory breath. See MACHINE TRIGGERING, PATIENT TRIGGERING, SYNCHRONICATION WINDOW, TRIGGER WINDOW, CONTINUOUS MANDATORY VENTILATION, INTERMITTENT MANDATORY VENTILATION
MACHINE TRIGGERING | Starting inspiratory flow based on a signal (usually time) from the ventilator, independent of a patient trigger signal. Examples include triggering based on a preset frequency (which sets the ventilatory period), or based on a preset minimum minute ventilation (determined by tidal volume divided by the ventilatory period). If a signal from the patient (indicating an inspiratory effort) occurs within a synchronization window, the start of inspiration is defined as a machine trigger event that begins a mandatory breath. See MACHINE CYCLING, PATIENT TRIGGERING, SYNCHRONIZATION WINDOW, TRIGGER WINDOW, CONTINUOUS MANDATORY VENTILATION, INTERMITTENT MANDATORY VENTILATION
MANDATORY BREATH | During mechanical ventilation, a mandatory breath is one for which inspiration is started or stopped by the ventilator, independent of the patient; inspiration is machine triggered or machine cycled. A mandatory breath can occur during a spontaneous breath (eg, High Frequency Jet Ventilation). A mandatory breath is, by definition, assisted. See ASSISTED BREATH, SPONTANEOUS BREATH
MANDATORY MINUTE VENTILATION | A form of intermittent mandatory ventilation in which the ventilator monitors the exhaled minute ventilation as a target variable. If the exhaled minute ventilation falls below the operator set value, the ventilator will trigger mandatory breaths or increase the inspiratory pressure until the target is reached.
MEAN AIRWAY PRESSURE | The average airway pressure over time, usually one ventilatory period.
MECHANICAL VENTILATOR | An automatic machine designed to provide all or part of the work required in generating enough breaths to satisfy the body’s respiratory needs.
MINUTE VENTILATION | The product of alveolar volume (tidal volume minus dead space volume) times ventilatory frequency, usually expressed in L/min.
MINUTE VENTILATION | The product of tidal volume times ventilatory frequency, usually expressed in L/min.
MODE VARIETY | A variation of a mode that requires description beyond control variable, breath sequence, and targeting schemes. For example, NAVA and PAV are two varieties of PC-CSVs that differ in the specific ways they implement servo targeting
MODE OF VENTILATION | A predetermined pattern of interaction between a patient and a ventilator, specified as a particular combination of control variable, breath sequence, and targeting schemes for primary and secondary breaths.
MUSCLE PRESSURE | The aggregate forces developed by the respiratory muscles to produce breathing movements. Pmus is not measurable directly but is estimated as the equivalent transrespiratory pressure required to inflate the passive respiratory system to the same tidal volume. See Pmus, TRANSRESPIRATORY PRESSURE, EQUATION OF MOTION
NEGATIVE PRESSURE VENTILATION | A type of assisted breathing for which transrespiratory pressure difference is generated by keeping airway pressure equal to atmospheric pressure and making body surface pressure less than atmospheric pressure. Examples would be ventilation with an “iron lung” or “chest cuirass”.
NEURALLY ADJUSTED VENTILATORY ASSIST | The name of a mode using a servo targeting scheme in which the controller sets airway pressure to be proportional to patient effort based on the voltage recorded from diaphragmatic activity from sensors embedded in an orogastric or nasogastric tube.
OPTIMAL TARGETING SCHEME | A ventilator control system that automatically adjusts the targets of the ventilatory pattern to either minimize or maximize some overall performance characteristic. This is an advanced form of adaptive targeting. One example is Adaptive Support Ventilation (Hamilton Medical G5 ventilator) in which the ventilator adjusts the mandatory tidal volume and frequency (for a passive patient) is such a way as to minimize the work rate of ventilation.
PARTIAL VENTILATORY SUPPORT | The ventilator and the respiratory muscles each provide some of the work of breathing; muscle pressure adds to ventilator pressure in the equation of motion. See FULL VENTILATORY SUPPORT, EQUATION OF MOTION
PASSIVE VENTILATION | Inspiratory muscles do not contribute to the generation of inspiratory flow. Passive ventilation is assisted by definition (ie, the ventilator does all of of the work of breathing). See ASSISTED BREATH
PATIENT CYCLING | Ending inspiratory time based on signals representing the patient determined components of the equation of motion, (Pmus , elastance, or resistance). Common examples of cycling variables are peak inspiratory pressure and percent inspiratory flow. See machine triggering, machine cycling, patient triggering, synchronization window, trigger window, continuous mandatory ventilation, intermittent mandatory ventilation
PATIENT TRIGGERING | Starting inspiration based on a patient signal occurring in a trigger window, independent of a machine trigger signal. The signal is related to one of the patient determined components of the equation of motion (Pmus, elastance, or resistance). Common examples of patient trigger variables are airway pressure drop below baseline and inspiratory flow due to patient effort. See MACHINE TRIGGERING, MACHINE CYCLING, SYNCHRONICATION WINDOW, TRIGGER WINDOW, CONTINUOUS MANDATORY VENTILATION, INTERMITTENT MANDATORY VENTILATION
PC-CMV | Pressure controlled continuous mandatory ventilation.
PC-IMV | Pressure controlled intermittent mandatory ventilation.
PC-CSV | Pressure controlled continuous spontaneous ventilation.
PEAK AIRWAY PRESSURE | The maximum airway pressure measured during a mechanically assisted inspiration, measured relative to atmospheric pressure.
PEAK INSPIRATORY PRESSURE | The maximum pressure (relative to atmospheric pressure) during inspiration. See INSPIRATORY PRESSURE TARGET
PEEP | Positive end expiratory pressure; the value of transrespiratory system pressure at end expiration. See CPAP, EXPIRATORY PRESSURE TARGET
PLATEAU PRESSURE | The pressure at the airway opening (relative to atmospheric pressure) during an inspiratory hold (ie, flow throughout the respiratory system is zero), ie, passive recoil pressure of the elastance of the respiratory system
Pmus | Muscle pressure. The aggregate forces developed by the respiratory muscles to produce breathing movements. Pmus is not measurable directly but is estimated as the equivalent transrespiratory pressure required to inflate the passive respiratory system to the same tidal volume. See TRANSRESPIRATORY PRESSURE, EQUATION OF MOTION
POSIVITVE PRESSURE VENTILATION | A type of assisted breathing for which transrespiratory pressure difference is generated by raising airway pressure above body surface pressure (usually equal to atmospheric pressure). Examples would be ventilation with intensive care or transport ventilators.
PRESSURE | A measure of force per unit of area at a particular point in space.
PRESSURE CHANGE | The difference between pressure (or pressure gradient) measured at one point in time and the same pressure measured at a previous point in time.
PRESSURE CONTROL | A general category of ventilator modes for which pressure delivery is either proportional to patient effort or has a particular waveform regardless of respiratory system mechanics. When inspiratory pressure is preset, we further specify that inspiration must start out with the preset pressure to avoid confusion with dual targeting that may switch from a preset flow to a preset pressure (eg, Pmax feature used with volume control modes on the Dräger Evita Infinity V500 ventilator). See dual targeting scheme. According to the equation of motion, pressure control means that inspiratory pressure is predetermined as the independent variable so that volume and flow become the dependent variables. See VOLUME CONTROL, EQUATION OF MOTION, DUAL TARGETING
PRESSURE GRADIENT | The difference between pressure measured at one point in space and another point in space. Examples include the pressure difference across a cell membrane causing gas diffusion into the cell and the pressure difference across the respiratory system causing flow into the lungs. See TRANSAIRWAY PRESSURE, TRANSALVEOLAR PRESSURE, TRANSCHESTWALL PRESSURE, TRANSPULMONARY PRESSURE, TRANSRESPIRATORY PRESSURE
PRESSURE CYCLING | Ending inspiratory time when airway pressure reaches a preset threshold.
PRESSURE RISE TIME | see RISE TIME
PRESSURE SUPPORT | The name of a mode that is classified as pressure control continuous spontaneous ventilation with set-point targeting; inspiration is pressure or flow triggered and flow cycled.
PRESSURE TRIGGERING | The starting of inspiratory flow due to a patient inspiratory effort that generates an airway pressure drop below end expiratory pressure larger than a preset threshold (ie, the trigger sensitivity setting). Pressure triggering is a form of patient triggering.
PRESSURE TARGET | Inspiratory pressure reaches a preset value before inspiration cycles off.
PRIMARY BREATHS | Mandatory breaths during CMV or IMV or spontaneous breaths during CSV.
PROPORTIONAL ASSIST VENTILATION (PAV) | The name of a mode that is classified as pressure control continuous spontaneous ventilation with servo targeting based on the equation of motion for the respiratory system and making inspiratory pressure proportional to inspiratory effort. See EQUATION OF MOTION
Pvent | Total driving pressure during mechanical ventilation; elastance x volume + resistance x flow. See DRIVING PRESSURE, EQUATION OF MOTION
RAMP | A mathematical function whose value rises or falls at a constant rate. Ascending (rising) or descending (falling) functions are sometimes used for inspiratory flow in volume control modes.
RESISTANCE | defined as the ratio of the change in the pressure difference across the system to the associated change in flow. A mechanical property of a structure such as the respiratory system; a parameter of a lung model, or setting of a lung simulator
RESISTIVE LOAD | The pressure difference applied across a system (e.g., a container) that is related to a rate of change of the system's volume and/or the flow of fluid within or through the system. For a linear system: resistance x flow, or, resistance x rate of change of volume; for a container, the effective resistance includes the mechanical (usually viscous) resistances of its structural components and the flow resistance of the fluid [gas or liquid] within it.
RISE TIME | The period from the start of increase of a waveform above baseline to (or near) its peak value. To avoid ambiguity caused by noise in the signal, rise times are often specified as the period from 10% to 90% of peak value.
SECONDARY BREATHS | Spontaneous breaths during intermittent mandatory ventilation
SENSITIVITY | The sensitivity setting of the ventilator is the amount the trigger variable must change to start inspiratory flow. Sensitivity is sometimes used to refer to cycling also. The trigger or cycle threshold is the particular value of the trigger or cycle variable that initiates the event. For example, if a pressure trigger sensitivity is set at -2 cm H2O below PEEP, and PEEP is set at 5 cm H2O, then the trigger threshold is +3 cm H2O.
SERVO TARGETING | A control system for which the output of the ventilator automatically follows a varying input. In practice, this usually means that inspiratory pressure is proportional to inspiratory effort. For example, the Automatic Tube Compensation feature on the Dräger Evita 4 ventilator tracks flow and forces pressure to be equal to flow squared and multiplied by a constant (representing endotracheal tube resistance). Other examples include Proportional Assist Ventilation (Covidien PB 840 ventilator; pressure is proportional to spontaneous volume and flow) and Neurally Adjusted Ventilatory Assist (Maquet Servo-i ventilator; pressure is proportional to diaphragmatic electrical activity). For all three of these example modes airway pressure is effectively proportional to the patient’s inspiratory effort.
SET-POINT TARGETING | A control system for which the operator sets the parameters of the pressure waveform (eg, inspiratory pressure, pressure rise time, and inspiratory time for pressure control modes) or volume and flow waveforms (eg, tidal volume and inspiratory flow for volume control modes). There are no targets set automatically by the ventilator. Advanced volume control modes actually allow the ventilator to make small adjustments to the set inspiratory flow to compensate for such factors as patient circuit compliance. From an engineering point of view, this is adaptive feedback control, but from a ventilator mode taxonomy point of view, such adjustments are better seen as a way of implementing operator preset values, and thus classified as set-point targeting.
SINUSOIDAL | A mathematical function having a magnitude that varies as the sine of an independent variable (eg, time). A sinusoidal function is sometimes used for inspiratory flow in volume control modes.
SPECIFIC COMPLIANC | Compliance normalize to lung volume so that lung mechanics can be compared across different body sizes. Calculated as compliance divided by functional residual volume, FRC. See COMPLIANCE
SPECIFIC ELASTANCE | The inverse of specific compliance, ie, FRC divided by compliance. See ELASTANCE, SPECIFIC COMPLIANCE
SPONTANEOUS BREATH | A breath for which both the start and end of inspiration are determined by the patient, independent of any machine settings for inspiratory time and expiratory time: inspiration is both patient triggered and patient cycled. As a result, the patient retains substantial control over timing. Note that use of this definition for determining the breath sequence (ie, CMV, IMV, CSV) assumes normal ventilator operation. For example, coughing during VC-CMV may result in patient cycling for a patient triggered breath due to the pressure alarm limit. While inspiration for that breath is both patient triggered and patient cycled, this is not normal operation and the sequence does not turn into IMV. A spontaneous breath may occur during a mandatory breath (eg Airway Pressure Release Ventilation).A spontaneous breath may be assisted or unassisted. See ASSISTED BREATH, MANDATORY BREATH
SPO2 | Transcutaneous blood oxygen saturation (in percent) obtained with a pulse oximeter.
STATIC COMPLIANCE | The slope of the pressure-volume curve drawn between two points of zero flow throughout the system (eg, at the start of inspiration and end of inspiratory hold). See DYNAMIC COMPLIANC, DYNAMIC CHARACTERISTIC.
STRAIN | The deformation of a structure associated with an applied stress. In terms of pulmonry physiology, strain is the lung volume change (ie, tidal volume) associated with the stress (change in transalveolar pressure) relative to the normal end expiratory lung volume (ie, functional residual capacity, FRC).
STRESS | The internal distribution of the counterforce per unit of area that balances and reacts to an external load. In terms of pulmonary physiology, stress is the tidal pressure associated with the delivery of a tidal volume. There is a linear relation between stress ( the change in transalveolar pressure, PTA or static transpulmonary pressure) and strain (the change in lung volume, V, above functional residual volume, FRC), with the constant of proportionality being specific lung elastance, Espec divided by FRC: DPTA = Espec x (DV/FRC) Rearranging this equation yields Espec = DPTA x (FRC/DV) Now, if the change in DPTA is large enough to make DV equal to FRC, then FRC/DV = 1 and specific elastance can be interpreted as the change in transalveolar pressure required to double the FRC.
SYNCHRONIZED IMV (SIMV) | A form of IMV in which mandatory breath delivery is coordinated with patient inspiratory effort. A synchronized breath is considered to be machine triggered. See INTERMITTENT MANDATORY VENTILATION
SYNCHRONIZATION WINDOW | A short period, at the end of a preset expiratory time or at the end of a preset inspiratory time (for pressure control inspirations) during which a patient signal may be used to synchronize a machine trigger or cycle event to a patient trigger signal. If the patient signal occurs during an expiratory time synchronization window, inspiration starts and is defined as a machine triggered event. This is because the mandatory breath would have been time triggered regardless of whether the patient signal had appeared or not and because the distinction is necessary to avoid logical inconsistencies in defining mandatory and spontaneous breaths which are the foundation of the mode taxonomy. If inspiration is triggered in a synchronization window, the actual ventilatory period for the previous breath will be shorter than the set ventilatory period (determined by the set mandatory breath frequency). Some ventilators add the lost time to the next mandatory breath period to maintain the set frequency. Sometimes a synchronization window is used at the end of the inspiratory time of a pressure controlled, time cycled breath (eg Airway Pressure Release Ventilation). If the patient signal occurs during such an inspiratory time synchronization window, expiration starts and is defined as a machine cycled event. See INTERMITTENT MANDATORY VENTILATION, MACHINE TRIGGERING, PATIENT TRIGGERING, TRIGGER WINDOW
SYNCHRONY | Synchrony during mechanical ventilation is defined as a zero phase difference between the pressure waveform representing patient inspiratory effort, Pmus, and the pressure waveform generated by the ventilator during an assisted breath. See DYSSYNCHRONY, ASYNCHRONY
TAG | A mode classification, or taxonomic attribute grouping. A tag can be an acronym. For example the mode named Volume A/C is classified as volume control (VC) continuous mandatory ventilation (CMV) with set-point targeting (s) and can be represented as VC-CMVs. Another example (using both primary and secondary breaths) would be PRVC SIMV classified as PC-IMVa,s, where the primary breath uses adaptive targeting (a) and the secondary breath uses set-point targeting (s). The mode named Adaptive Support Ventilation has multiple targeting for each type of breath (ie, both optimal, o, and intelligent, i). It is classified as PC-IMVoi,oi.
TARGET | A predetermined goal of ventilator output. Targets can be viewed as the goals of the targeting scheme. Within-breath targets are the parameters of the pressure, volume, or flow waveform. Examples of within-breath targets include inspiratory flow or pressure and rise time (set-point targeting), tidal volume (dual targeting) and constant of proportionality between inspiratory pressure and patient effort (servo targeting). Note that preset values within a breath that end inspiration, such as tidal volume, inspiratory time, or percent of peak flow, are also cycle variables. Between-breath targets serve to modify the within-breath targets and/or the overall ventilatory pattern. Between-breath targets are used with more advanced targeting schemes, where targets act over multiple breaths. Examples of between-breath targets and targeting schemes include average tidal volume (for adaptive targeting), percent minute ventilation (for optimal targeting) and combined PCO2, volume, and frequency values describing a “zone of comfort” (for intelligent targeting). Symbols used for targeting schemes: s = set-point, d = dual, b = bio-variable, r = servo, a = adaptive, o = optimal, i = intelligent.
TARGETING SCHEME | A model of the relationship between operator inputs and ventilator outputs to achieve a specific ventilatory pattern, usually in the form of a feedback control system. The targeting scheme is a key component of a mode description.
TAXONOMY | A hierarchical classification system. A taxonomy for modes of ventilation has four levels: 1) the control variable, 2) the breath sequence; 3) the targeting scheme for primary breaths and ; 4) the targeting scheme for secondary breaths.
TIDAL PRESSURE | The change in transalveolar pressure associated with the inhalation or exhalation of a tidal volume. Calculated as tidal volume divided by compliance or plateau pressure minus total PEEP. Sometimes called driving pressure.
TIDAL VOUME | The volume of gas, either inhaled or exhaled, during a breath. The maximum value of the volume vs time waveform, corresponding to the integral of flow from the start of inspiratory flow to the start of expiratory flow.
TIME CONSTANT | The time at which an exponential function attains 63% of its steady state value in response to a step input, calculated as the product of resistance and compliance (usually in units of seconds) The time necessary for inflated lungs to passively empty by 63%; the time necessary for the lungs to passively fill 63% during pressure control ventilation with a rectangular pressure waveform. Passive inhalation or exhalation is virtually complete after 5 time constants.
TIME CYCLING | Ending inspiratory time after a preset time interval has elapsed. The most common examples are a preset inspiratory time or a preset inspiratory pause time. Time cycling is a form of machine cycling.
TIME TRIGGERING | The starting of inspiratory time due to a preset time interval. The most common example is a preset ventilatory frequency.
TOTAL CYCLE TIME | Same as ventilatory period, the sum of inspiratory time and expiratory time, usually expressed in seconds.
TOTAL PEEP | The sum of autoPEEP and intentionally applied PEEP or CPAP.
TOTAL VENTILATORY SUPPORT | See FULL VENTILATORY SUPPORT
TRANSAIRWAY PRESSURE | Pressure at the airway opening minus pressure in the lungs (i.e., alveolar pressure).
TRANSALVEOLAR PRESSURE | Pressure in the lungs minus pressure in the pleural space. Equal to transpulmonary pressure only under static conditions.
TRANSCHESTWALL PRESSURE | Pressure in the pleural space minus pressure on the body surface.
TRANSPULMONARY PRESSURE | Pressure at the airway opening minus pressure in the pleural space.
TRANSRESPIRATORY PRESSURE | Pressure at the airway opening minus pressure on the body surface; equal to the sum of transairway pressure plus transalveolar pressure plus transchestwall pressure.
TRANSTHORACIC PRESSURE | Pressure in the lungs minus pressure on the body surface; equal to the sum of transalveolar pressure plus transchestwall pressure
TRIGGER | To start the inspiratory time. See MACHINE TRIGGERING, PATIENT TRIGGERING
TRIGGER VARIABLE | The variable (usually pressure, volume, flow, or time) that is used to start the inspiratory time.
TRIGGER WINDOW | The period comprised of the entire expiratory time minus a short “refractory” period at the end of the previous inspiration, required to reduce the risk of double triggering (eg, for CMV and CSV breath sequences) and minus any synchronization window (IMV breath sequences). If a signal from the patient (indicating an inspiratory effort) occurs within this trigger window, inspiration starts and is defined as a patient triggered event. See See MACHINE TRIGGERING, MACHINE CYCLING, SYNCHRONICATION WINDOW, CONTINUOUS MANDATORY VENTILATION, INTERMITTENT MANDATORY VENTILATION
VENTILATORY PATTERN | A sequence of breaths (CMV, IMV, or CSV) with a designated control variable (volume or pressure) for the mandatory breaths (or the spontaneous breaths for CSV).
VENTILATORY PERIOD | The time from the start of inspiratory flow of one breath to the start of inspiratory flow of the next breath; inspiratory time plus expiratory time; the reciprocal of ventilatory frequency. Also called total cycle time or total breath cycle.
VOLUME CONTROL | A general category of ventilator modes for which both inspiratory flow and tidal volume are preset and independent of respiratory system mechanics. Usually, flow and tidal volume may be set directly by the operator. Alternatively, the ventilator may determine tidal volume based on operator preset values for frequency and minute ventilation or the ventilator may determine inspiratory flow based on operator set tidal volume and inspiratory time. When inspiratory volume and flow are preset, we further specify that inspiration must start out with the preset flow to avoid confusion with dual targeting that may switch from a preset pressure to a preset flow and volume (eg. Volume Assured Pressure Support). See dual targeting scheme. Note that setting tidal volume is a necessary but not sufficient criterion for volume control. The reason is that some ventilators use pressure control with adaptive targeting and allow the operator to set a tidal volume but not an inspiratory flow. In this case, the tidal volume setting refers to the between-breath tidal volume target, not a within-breath target. See adaptive targeting scheme. Likewise, setting inspiratory flow is also a necessary but not sufficient criterion for volume control. For example, the Bird Mark 7 ventilator requires an inspiratory flow setting but has no tidal volume setting. Instead the operator sets the inspiratory pressure, which is also the cycle variable. Hence, breaths are pressure controlled, and changing lung mechanics change the rate of pressure rise, the inspiratory time, and hence the delivered tidal volume as in other examples of pressure control. According to the equation of motion, volume control means that both volume and flow are predetermined as the independent variables and pressure is thus the dependent variable. See PRESSURE CONTROL, EQUATION OF MOTION
VOLUME CYCLING | Ending inspiratorytime when inspiratory volume reaches a preset threshold (ie, tidal volume). Volume cycling is a form of machine cycling.
VC-CMV | Volume controlled continuous mandatory ventilation.
VC-IMV | Volume controlled intermittent mandatory ventilation.
VOLUME TARGET | A preset value for tidal volume that the ventilator is set to attain either within a breath or as an average over multiple breaths.
VOLUME TRIGGERING | The starting of inspiratory time due to a patient inspiratory effort that generates an inspiratory volume signal larger than a preset threshold (ie, the trigger sensitivity setting).
VOLUTRAUMA | Ventilator induced lung injury caused by excessively large tidal volumes. See BAROTRAUMA
WORK OF BREATHING | The general definition of work is the integral of pressure change with respect to volume change during ventilation. Pressure means the pressure difference across a struc ture, eg, transrepsiratory pressure or transpulmonary pressure or the equivalent muscle pressure. There are two general components of work related to mechanical ventilation. One kind is the work performed by the ventilator on the patient, which is reflected by a positive change in airway pressure above baseline during inspiration. The second component is the work the patient does on the ventilator to (eg, to trigger inspiration), which is reflected by a negative change in airway pressure below baseline during inspiration.
Menu

## Ex VC
Ventilator Simulator Exercise Workbook
Volume Control Exercises
| Setp 1: Set any values for Patient and Ventilator |  |  |  | Step 3: Change compliance to any value |  |  |  |  | Step 6: Change resistance to any value |  |  |  |  | Step 9: Change tidal volume to any value |  |  |  |  | Step 12: Change peak flow to any value
| Step 2: Record values below |  |  |  | Step 4: Redcord resultant values |  |  |  |  | Step 7: Redcord resultant values |  |  |  |  | Step 10: Redcord resultant values |  |  |  |  | Step 13: Redcord resultant values
[5 empty] | Step 5: Indicate by an x if the result was an |  |  |  |  | Step 8: Indicate by an x if the result was an |  |  |  |  | Step 11: Indicate by an x if the result was an |  |  |  |  | Step 14: Indicate by an x if the result was an
[6 empty] | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change
[7 empty] | Result |  |  |  |  | Result |  |  |  |  | Result |  |  |  |  | Result
|  |  | Patient Parameters |  | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | =
|  |  |  | Elastance (cm H2O/L) |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L)
|  |  |  | Compliance (L/cm H2O) |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O)
|  |  |  | Resistance (cm H2O/L/s) |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s)
|  |  |  | Time Constant (s) |  | Time Constant (s) |  |  |  |  | Time Constant (s) |  |  |  |  | Time Constant (s) |  |  |  |  | Time Constant (s)
|  |  |  | Tidal Lung Pressure (cm H2O) |  | Tidal Lung Pressure (cm H2O) |  |  |  |  | Tidal Lung Pressure (cm H2O) |  |  |  |  | Tidal Lung Pressure (cm H2O) |  |  |  |  | Tidal Lung Pressure (cm H2O)
|  |  | Ventilator Variables |  | Ventilator Variables |  |  |  |  | Ventilator Variables |  |  |  |  | Ventilator Variables |  |  |  |  | Ventilator Variables
|  |  |  | Flow (square wave, L/min) |  | Flow (square wave, L/min) |  |  |  |  | Flow (square wave, L/min) |  |  |  |  | Flow (square wave, L/min) |  |  |  |  | Flow (square wave, L/min)
|  |  |  | Flow (L/s) |  | Flow (L/s) |  |  |  |  | Flow (L/s) |  |  |  |  | Flow (L/s) |  |  |  |  | Flow (L/s)
|  |  |  | Tidal Volume (mL) |  | Tidal Volume (mL) |  |  |  |  | Tidal Volume (mL) |  |  |  |  | Tidal Volume (mL) |  |  |  |  | Tidal Volume (mL)
|  |  |  | Tidal Volume (L) |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L)
|  |  |  | Frequency (breaths/min) |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min)
|  |  |  | PEEP (cm H2O) |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O)
|  |  |  | Pause Time (s) |  | Pause Time (s) |  |  |  |  | Pause Time (s) |  |  |  |  | Pause Time (s) |  |  |  |  | Pause Time (s)
|  |  |  | Total Cycle Time (s) |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s)
|  |  |  | Inspirtory Time (s) |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s)
|  |  |  | I:E |  | I:E |  |  |  |  | I:E |  |  |  |  | I:E |  |  |  |  | I:E
|  |  |  | Duty Cycle (%) |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%)
|  |  |  | Plateau pressure (cm H2O) |  | Plateau pressure |  |  |  |  | Plateau pressure |  |  |  |  | Plateau pressure |  |  |  |  | Plateau pressure
|  |  |  | Driving Pressure (cm H2O) |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O)
|  |  |  | Peak Inspiratory Flow (L/min) |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min)
|  |  |  | Peak Expiratory Flow (L/min) |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min)
|  |  |  | Mean Inspiratory Pressure (cm H2O) |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O)
|  |  |  | Mean Airway Pressure (cm H2O) |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O)
Menu

## Ex PC
VentSet simulator Exercise Workbook
Pressure Control Exercises
| Setp 1: Set any values for Patient and Ventilator |  |  |  | Step 3: Change compliance to any value |  |  |  |  | Step 6: Change resistance to any value |  |  |  |  | Step 9: Change inspiratory pressure to any value |  |  |  |  | Step 12: Change inspiratory time to any value
| Step 2: Record values below |  |  |  | Step 4: Redcord resultant values |  |  |  |  | Step 7: Redcord resultant values |  |  |  |  | Step 10: Redcord resultant values |  |  |  |  | Step 13: Redcord resultant values
[5 empty] | Step 5: Indicate by an x if the result was an |  |  |  |  | Step 8: Indicate by an x if the result was an |  |  |  |  | Step 11: Indicate by an x if the result was an |  |  |  |  | Step 14: Indicate by an x if the result was an
[6 empty] | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change |  |  |  |  | increase, decrease, or no change
[7 empty] | Result |  |  |  |  | Result |  |  |  |  | Result |  |  |  |  | Result
|  |  | Patient Parameters |  | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | = | Patient Parameters |  | ­ | ¯ | =
|  |  |  | Elastance (cm H2O/L) |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L) |  |  |  |  | Elastance (cm H2O/L)
|  |  |  | Compliance (L/cm H2O) |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O) |  |  |  |  | Compliance (L/cm H2O)
|  |  |  | Resistance (cm H2O/L/s) |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s) |  |  |  |  | Resistance (cm H2O/L/s)
|  |  |  | Time Constant (s) |  | Time Constant (s) |  |  |  |  | Time Constant (s) |  |  |  |  | Time Constant (s) |  |  |  |  | Time Constant (s)
|  |  |  | Tidal Volume (L) |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L) |  |  |  |  | Tidal Volume (L)
Ventilator Variables
|  |  |  | Inspiratory Pressure (cm H2O) |  | Inspiratory Pressure (cm H2O) |  |  |  |  | Inspiratory Pressure (cm H2O) |  |  |  |  | Inspiratory Pressure (cm H2O) |  |  |  |  | Inspiratory Pressure (cm H2O)
|  |  |  | Peak Inspiratory Pressure (cm H2O) |  | Peak Inspiratory Pressure (cm H2O) |  |  |  |  | Peak Inspiratory Pressure (cm H2O) |  |  |  |  | Peak Inspiratory Pressure (cm H2O) |  |  |  |  | Peak Inspiratory Pressure (cm H2O)
|  |  |  | Frequency (breaths/min) |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min) |  |  |  |  | Frequency (breaths/min)
|  |  |  | Inspiratory Time (s) |  | Inspiratory Time (s) |  |  |  |  | Inspiratory Time (s) |  |  |  |  | Inspiratory Time (s) |  |  |  |  | Inspiratory Time (s)
|  |  |  | PEEP (cm H2O) |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O) |  |  |  |  | PEEP (cm H2O)
|  |  |  | Total Cycle Time (s) |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s) |  |  |  |  | Total Cycle Time (s)
|  |  |  | Inspirtory Time (s) |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s) |  |  |  |  | Inspirtory Time (s)
|  |  |  | I:E |  | I:E |  |  |  |  | I:E |  |  |  |  | I:E |  |  |  |  | I:E
|  |  |  | Duty Cycle (%) |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%) |  |  |  |  | Duty Cycle (%)
|  |  |  | Plateau pressure |  | Plateau pressure |  |  |  |  | Plateau pressure |  |  |  |  | Plateau pressure |  |  |  |  | Plateau pressure
|  |  |  | Driving Pressure (cm H2O) |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O) |  |  |  |  | Driving Pressure (cm H2O)
|  |  |  | Peak Inspiratory Flow (L/min) |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min) |  |  |  |  | Peak Inspiratory Flow (L/min)
|  |  |  | Peak Expiratory Flow (L/min) |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min) |  |  |  |  | Peak Expiratory Flow (L/min)
|  |  |  | Mean Inspiratory Pressure (cm H2O) |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O) |  |  |  |  | Mean Inspiratory Pressure (cm H2O)
|  |  |  | Mean Airway Pressure (cm H2O) |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O) |  |  |  |  | Mean Airway Pressure (cm H2O)
Menu

## Class
1245 |  |  | How to Classify Modes of Ventilation
https://app.box.com/s/sbnajwyhdo53yf6wqcv50orxypumb16j
1. Identify the Control Variable (Figure 1)
Volume Control
|  |  | Definition | Operator can preset target values for tidal volume and inspiratory flow
|  |  | Example | Mode called "VC Continuous Mandatory Ventilation" on the Dräger Evita Infinity V500 ventilator
Pressure Control
|  |  | Definition | Operator can preset target value for inspiratory pressure or inspiratory pressure is proportional to inspiratory effort
|  |  | Example | Mode called "Pressure Support" on most ventilators (preset inspiratory pressure target)
Mode called "Proportional Assist Ventilation (PAV)" on Medtronic PB 840 ventilator (inspiratory pressure proportional to effort)
Time Control
|  |  | Definition | Operator cannot preset target values for volume, flow, or pressure
|  |  | Example | Mode called "High Frequency Oscillation" on Sensormedics high frequency oscillatory ventilator
2. Identify Breath Sequence (Figure 3)
2A. Determine trigger and cycle variables (see ventilator operator's manual)
2B. Determine trigger and cycle events (Figure 2)
Continuous Mandatory Ventilation
|  |  | Definition | Spontaneous breaths are not permitted between mandatory breaths
|  |  | Example | Mode called "Assist Control" on most ventilators
Intermittent Mandatory Ventilation
|  |  | Definition | Spontaneous breaths are parmitted between mandatory breaths
|  |  | Example | Mode called "SIMV" on most ventilators
Continuous Spontaneous Ventilation
|  |  | Definition | Mandatory breaths are not parmitted between spontaneous breaths (all breaths are spontaneoust)
|  |  | Example | Mode called "SIMV" on most ventilators
3. Identify Targeting Schemes
Set-point (s)
|  |  | Definition | Operator sets all targets
|  |  | Example | Mode called "Pressure Control Continuous Mandatory Ventilation" on Hamilton G5 ventilator
Dual (d)
|  |  | Definition | Ventilator switches between volume control and pressure control
|  |  | Example | Mode called "Volume Control" on Maquet Servo-i ventilator (see Respir Care 2012;57(8):1297–1304)
Bio-variable (b)
|  |  | Definition | Ventilator randomly changes target inspiratory pressure
|  |  | Example | Mode called "Variable Pressure Support" on the Dräger Evita Infinity V500 ventilator
Servo (r)
|  |  | Definition | Ventilator automatically sets inspiratory pressure to be proportional to inspiratory effort
|  |  | Example | Mode called "Neurally Adjusted Ventilatory Assist" on the Maquet Servo-I ventilator
Adaptive (a)
|  |  | Definition | Ventilator adjusts targets to adapt to changing patient conditions
a. Ventilator adjusts inspiratory pressure to assure average tidal volume matches set target
b. Ventilator adjusts mandatory breath rate in response to patient trigger efforts to meet minute volume target
|  |  | Examples | a. Mode called "Pressure Regulated Volume Control" on Maquet Servo-i ventilator
b. Mode called "Mandatory Minute Volme" on the Dräger Evita Infinity V500 ventilator
Optimal (o)
|  |  | Definition | Ventilator adjusts targets to maximize or minimize some desired parameter
a. Ventilator adjusts frequency and inspiratory pressure of mandatory breaths to minimize power
transmission from ventilator to patient
|  |  | Example | Mode called "Adaptive Support Ventilation" on Hamilton G5 ventilator
Intelligent (i)
|  |  | Definition | Ventilator uses tools of artificial intelligence to set targets
|  |  | Example | a. Mode called "SmartCare PS" on the Dräger Evita Infinity V500 ventilator
b. Mode called "IntelliVent" on the Hamilton G5 ventilator
Example Mode Classification
Nome name "A/C Volume Control"
Get the Free App for Android: VENTILATOR MODE MAP
https://play.google.com/store/apps/details?id=mandupress.ventilatormodemap
Menu
Figure 1. Rubric for determining the control variable of a mode.
Figure 2. Rubric for determining trigger and cycle events
Figure 3. Rubric for determining breath sequence

## Ref
References
Key Articles
1. Chatburn RL, Khatib ME, Mireles-Cabodevila E. A Taxonomy for Mechanical Ventilation: 10 Fundamental Maxims. Respir Care 2014;59(11):1747–1763.
https://pubmed.ncbi.nlm.nih.gov/25118309/
2. Mireles-Cabodevila E, Hatipoglu U, Chatburn RL. A rational framework for selecting modes of ventilation. Respir Care. 2013 ;58(2):348-66.
https://pubmed.ncbi.nlm.nih.gov/22710796/
Articles Referencing Key Articles
· ECRI. New taxonomy for ventilation modes and features. Health Devices August 2013;42(8).
· Robertson TE. Ventilator Management: A Systematic Approach to Choosing and Using New Modes. Adv Surg 2016;50(1):173-186.
· Murias G, Montanya J, Chacon E, Estruga A, Subira C, Fernandez R, et al. Automatic detection of ventilatory modes during invasive mechanical ventilation. Crit Care 2016;20(1):258.
· Chatburn RL, Mireles-Cabodevila E, Sasidhar M. Tidal volume measurement error in pressure control modes of mechanical ventilation: A model study. Comput Biol Med 2016;75:235-242.
· Volsko TA, Hoffman J, Conger A, Chatburn RL. The effect of targeting scheme on tidal volume delivery during volume control mechanical ventilation. Respir Care 2012;57(8):1297-1304.
· Mayordomo-Colunga J, Pons-Odena M, Medina A, Rey C, Milesi C, Kallio M, et al. Non-invasive ventilation practices in children across Europe. Pediatr Pulmonol 201;1-8.
Book Chapters
· Chatburn RL. Classification of mechanical ventilators and modes of ventilation. In: Tobin MJ, ed. Principles and practice of mechanical ventilation, 3rd New York: McGraw-Hill, 2013:45-64.
· Chatburn RL, Mireles-Cabodevila E. Basic principles of ventilator design and operation. In: Tobin MJ, ed. Principles and practice of mechanical ventilation, 3rd New York: McGraw-Hill, 2013:65-97.
· Chatburn RL, Carlo WA. A taxonomy for mechanical ventilation. In: Berhardt LV, editor, Advances in biology and medicine, volume 78. Nova Scientific Publishers, Inc. 2014, ISBN: 978-1-63117-664-7 (eBook)
https://app.box.com/s/jgmx7o3btngwm2mdnbm51orni1m296p0
o password for pdf is: 344544
· Chatburn RL, Volsko TA. Mechanical ventilators: Classification and principles of operation. In: Hess DR, MacIntyre NR, Galvin WF, Mishoe SC. Respiratory Care: Principles and practice. 3rd Burlington: Jones & Bartlett Learning, 2016.
· Chatburn RL, Mireles-Cabodevila E. Design and function of mechanical ventilators. In: Oxford textbook of critical care, second edition. Edited by: Webb A, Angus DC, Finer S, Gattinoni L, Singer M. Oxford: Oxford University Press, 2016:419-429.
· Chatburn RL, Volsko TA. Mechanical ventilators. In: Kacmarek RM, Stoller JK and Heuer AH. Egan’s Fundamentals of Respiratory Care, 11th edition. St. Louis: Mosby Elsevier, 2017
· Keszler M, Chatburn RL. Overview of assisted ventilation. In: Assisted ventilation of the neonate. Edited by Goldsmith JP, Karotkin EH, Keszler M, Suresh GK. Philadelphia: Elsevier, 2017:140-152.
· Chatburn RL, Carlo WA. Description of available devices. In: Assisted ventilation of the neonate. Edited by Goldsmith JP, Karotkin EH, Keszler M, Suresh GK. Philadelphia: Elsevier, 2017:251-274.
Books
· Chatburn RL, Mireles-Cabodevila M. Handbook for Respiratory Care. 3rd edition. Sudbury: Jones & Bartlett Learning,2011. Chinese translation. Taiwan: Ho-Chi Book Publishing, 2015.
https://www.amazon.com/Handbook-Respiratory-Care-Robert-Chatburn/dp/0763784095/ref=sr_1_1?ie=UTF8&qid=1513098834&sr=8-1&keywords=chatburn+handbook
· Volsko TA, Chatburn RL, El-Khatib M. Respiratory Care Equipment. Burlington: Jones & Bartlett Learning, 2016.
o This book is unique in that it classifies all the modes on all the ventilators sold in the US using the taxonomy described in reference #1 in Key Articles above
https://www.amazon.com/Equipment-Respiratory-Care-Teresa-Volsko/dp/1449652832/ref=sr_1_1?s=books&ie=UTF8&qid=1513098878&sr=1-1&keywords=chatburn+volsko
Menu

## Params
Parmeters for Creating Passive Adult Lung Models
Reference: Arnal JM, Garnero A, Saoli M, Chatburn RL. Parameters for Simulation of Adult Patients During Mechanical Ventilation. Respir Care 2017.
https://app.box.com/s/cb24ylc2en7haqf0c957dvkerjbhg6ud
ARDS
| Normal | COPD | All | Mild | Mod | Severe
C (mL/cm H2O) | 50 | 60 | 40 | 45 | 40 | 35
HH - R (cm H20×s/L) | 10 | 20 | 10 | 10 | 10 | 10
HH - RC (s) | 0.5 | 1.2 | 0.35 | 0.45 | 0.35 | 0.25
HME - R (cm H20×s/L) | 15 | 25 | 15 | 15 | 15 | 15
HME - RC (s) | 0.75 | 1.5 | 0.53 | 0.68 | 0.53 | 0.38
C = compliance of respiratory system
R = resistance of airways (natural and artificial)
RC = time constant of respiratory system
HH = ventilation with heated humidifier
HME = ventilation with heat and moisture exchanger
L = liter
s = second
Parmeters for Creating Active Pediatric Lung Models
Reference: Dexter A, McNinch N, Kaznoch D, Volsko TA. Validating Lung Models Using the ASL 5000 Breathing Simulator. Simulation in healthcare : journal of the Society for Simulation in Healthcare 2018.
https://www.ncbi.nlm.nih.gov/pubmed/?term=dexter+volsko
| Infant | BPD | Child | NMD
C (mL/cm H2O) | 7.5 | 1.8 | 15 | 14.5
R (cm H20×s/L) | 52 | 69 | 46 | 46
Frequency (bpm) | 50 | 60 | 25 | 25
max Pmus (cm H2O) | 9.8 | 16.5 | 9 | 5
C = compliance of respiratory system
R = resistance of airways (natural and artificial)
max Pmus = inspiratory effort amplitude
Infant = 3.8 kg with normal lungs
BPD = 3.8 kg wih severe bronchopulmonary dysplasia
Child = 10 kg with normal lungs
NMD = 9 kg with neuromuscular disease
Menu

## About
SIVA
Simulator Interface for Ventilation Analysis
V11.1.21
About This Simulator
This is a simulated laboratory experience demonstrating the basic physical interactions between ventilator and patient
Desiger and copyright holder: Prof. Robert L. Chatburn, MHHS, RRT-NPS, FAARC
Lerner College of Medicine of Case Western University, USA
Reviewer: Matthias van der Staay, MSc in Electrical Engeneering
Control Engineer at imtmedical AG, Switzerland
This simulator may be freely used and distributed, but not modified
Questions, concerns, and bug reports can be submitted to
robert.chatburn@gmail.com
Educational Objectives
Provide practice in basic ventilator settings for volume and pressure control modes
Show how ventilator settings interact (eg, volume, time, flow, frequency, I:E)
Demonstrate the effects of ventilator settings and model parameters on important ventilator variables such as plateau pressure, driving pressure, and mean airway pressure
Illustrate the effects of changing lung mechanics on ventilator waveforms
Apply the three step process to classify any mode on any ventilator
Identify and describe ventilator waveforms for basic modes
Demonstrate how ventilator settings alter the power transfer from ventilator to patient
Explain how patient inspiratory effort shifts work from the ventilator to the patient
Define all the major terms related to mechanical ventilation
Show the effects of pressure rise time and flow cycle threshold on tidal volume and flow for Pressure Support mode
| Indicate the conditions under which VT display error occurs during PC on some ventilators |  | see Comput Biol Med. 2016 Aug 1;75:235-42
Simulator Features
Based on single-compartment model, as used in all ventilators
Adjustable parameters for lung model and effort model
Evidence-based lung model parameters suggested for experimentation
Can be used to model neonatal, pediatric, and adult patients
Provides wide range of calculated values related ventilator settings
Calculates indicators of risk for ventilator induced lung injury and synchrony
Work per breath
Work shifting from ventilator to patient
Power transfer from ventilator to patient
Demonstrates volume and pressure as control variables
Educational Features
Facilitates learning of interaction among ventilator settings
Adjustable flow in volume control (constant or descending ramp)
Adjustable pressure rise and cycle threshold in Pressure Support
Slide shows that can be downloaded and used for teaching
Worksheets for guided self-study
Handouts for learners
Over 100 self-study questions with answers
Simplified instructions for classifying any mode of ventilation
Table comparing modes on major ICU ventilators
Ventilator Mode Map application that classifies every mode on every major mechanical ventilator used in the United States
Complete glossary of standard terms related to mechanical ventilation
Complete set of reference papers with links
Limitations
Simulates mechanics only, no gas exchange or hemodynamics models
Constant resistance and compliance
Demonstrates single breaths only (no breath sequences)
Only set-point targeting - no advanced modes
Menu

## Me
Professor Robert L. Chatburn, MHHS, RRT-NPS, FAARC

## Res
Resources
Handouts
Introduction to mechanical ventilation (one page)
https://app.box.com/s/9v9aszgq9zaabsxlmiu5qxl7otz45y9p
How to identify modes
https://app.box.com/s/ztz9emw78uvb8kk6uxo97ftmxyu4skg1
How to compare modes
https://app.box.com/s/vkgf6irc59k6zyohts1fglzodb3bly6p
Slide Shows (with annotations)
Crash Course in Mechanical Ventilation
https://app.box.com/s/ygeab08r4kv6zajjvugc2e1wfvshrbiw
How to Teach Mechanical Ventilation (for clinicians)
https://app.box.com/s/dn3c0e3nugx6127mrvikz69na50tnohb
How to identify modes (taxonomy and 10 maxims)
https://app.box.com/s/k9507cunrtgdqofqxoreaoputd4xtk0e
How to compare modes (rational frameworkd for selecting the best mode for the situation)
https://app.box.com/s/bb44w3lt0b7psfb0ieam2zd3a3bhxlox
Understanding ventilator graphics
https://app.box.com/s/cmbhvg3jxbq53tg3busvqfk71vtvyrld
History of mechanical ventilation (big file - 40 MB)
https://app.box.com/s/60qaw2w1zjgx8bwl9kiwl3gmfjid4u59
Textbook: Fundamentals of Mechanical Ventilation (2003)
https://app.box.com/s/2sz4c5gsnvos018i15l1gn0ditrwqkeb
Ten Fundamental Maxims of Ventilator Technology
Maxims
Self-Study Questions
Maxim 1: The breath
Maxim 2: The assisted breath
Maxim 3: Assistance by volume or pressure control
Maxim 4: Breaths classified by trigger and cycle events
Maxim 5: Patient or machine initiation of trigger and cycle events
Maxim 6: Spontaneous and mandatory breaths
Maxim 7: Breath sequences
Maxim 8: Ventilatory patterns
Maxim 9: Targeting schemes
Maxim 10: Mode classification
Menu

## Vents
COMPARISON OF VENTILATORS BY MODE CLASSIFICATIONS AND MANUFACTURER'S MODE NAMES
Tag | Mode Classification | Covidien PB 840 or PB 980 | Dräger Evita Infinity V500 | Hamilton G5 | Maquet Servo-i or Servo-u | Carescape R860
VC-CMVs | Volume control CMV with set-point targeting | Assist/Control Volume Control | Volume Control Continuous Mandatory Ventilation (not patient triggered) | Synchronized Controlled Mandatory Ventilation | Volume Control (with square or ramp flow waveform) | Assist/Control Volume Control
Volume Control Assist Control (patient triggered)
VC-CMVd | Volume control IMV with dual/adaptive and adaptive targeting | Mode not offered | Volume Assist Control with Pressure Limitation | Mode not offered | Volume Control (with flow assist waveform) | Assist/Control Volume Control with Pressure Limit
Volume Control Continuoues Mandatory Ventilation with Pressure Limitation
VC-IMVda,s | Volume control IMV with dual/adaptive and set-point targeting | Mode not offered | Volume Control Mandatory Minute Volume with Pressure Limited Ventilation | Mode not offered | Mode not offered | Mode not offered
VC-IMVa,s | Volume control IMV with adaptive and set-point targeting | Mode not offered | Volume Control Mandatory Minute Volume Ventilation | Mode not offered | Mode not offered | Mode not offered
VC-IMVa,sr | Volume control IMV with adaptive and set-point/servo targeting | Mode not offered | Volume Control Mandatory Minute Volume Ventilation with Automatic Tube Compensation | Mode not offered | Mode not offered | Mode not offered
VC-IMVda,a | Volume control IMV with dual/adaptive and adaptive targeting | Mode not offered | Mode not offered | Mode not offered | Automode (with flow assist waveform) - Volume Control to Volume Support | Mode not offered
VC-IMVsa,a | Volume control IMV with set-point/adaptive and adaptive targeting | Mode not offered | Mode not offered | Mode not offered | Automode - Volume Control (with square or descending ramp flow waveform) to Volume Support | Mode not offered
VC-IMVd,d | Volume control IMV with dual targeting | Mode not offered | Mode not offered | Mode not offered | Volume Control (with flow assist waveform) | Mode not offered
VC-IMVd,s | Volume control IMV with dual and set-point targeting | Mode not offered | Volume Control SIMV with Pressure Limitation | Mode not offered | SIMV (Volume Control with flow assist waveform) | SIMV Volume Control with Pressure Limit
VC-IMVd,sr | Volume control IMV with dual and set-point/servo targeting | Mode not offered | Volume Control SIMV with Pressure Limitation and Automatic Tube Compensation | Mode not offered | Mode not offered | SIMV Volume Control with Pressure Limit and Tube Compensation
VC-IMVda,sr | Volume control IMV with dual/adaptive and set-point/servo targeting | Mode not offered | Volume Control Mandatory Minute Volume with Pressure Limitation and Automatic Tube Compensation | Mode not offered | Mode not offered | Mode not offered
VC-IMVs,s | Volume control IMV with set-point targeting | SIMV Volume Control with PS | Volume Control SIMV | SIMV | SIMV Volume Control (with square or ramp flow waveform) | SIMV Volume Control
Volume Control Continuous Mandatory Ventilation (not patient triggered)
VC-IMVs,sr | Volume control IMV with set-point and set-point/servo targeting | SIMV Volume Control with Tube Compensation | Volume Control SIMV with Automatic Tube Compensation | SIMV with Tube Resistance Compensation | Mode not offered | SIMV Volume Control with Tube Compensation
PC-CMVa | Pressure control CMV with adaptive targeting | Assist/Control Volume Control Plus | Volume Control Continuous Mandatory Ventilation with AutoFlow/Volume Guarantee | Adaptive Pressure Ventilation Continuous Mandatory Ventilation | Pressure Regulated Volume Control | Assist/Control Pressure Regulated Volume Control
Volume Control/Assist Control with AutoFlow/Volume Guarantee
PC-CMVas | Pressure control CMV with adaptive targeting/set-point targeting | Mode not offered | Mode not offered | Mode not offered | Mode not offered | Pressure Regulated Volume Control with Pressure Minimum
PC-CMVar | Pressure control CMV with adaptive/servo targeting | Mode not offered | Volume Control Continuous Mandatory Ventilation with AutoFlow/Volume Guarantee and Automatic Tube Compensation | Adaptive Pressure Ventilation Continuous Mandatory Ventilation with Tube Resistance Compensation | Mode not offered | Assist/Control Pressure Regulated Volume Control with Tube Compensation
Volume Control/Assist Control with AutoFlow/Volume Guarantee and Automatic Tube Compensation
PC-CMVs | Pressure control CMV with set-point targeting | Assist/Control Pressure Control | Pressure Control/Assist Control (patient triggered) | Pressure Controlled Continuous Mandatory Ventilation | Pressure Control | Assist Control Pressure Control
Noninvasive Pressure Control
PC-CMVsr | Pressure control CMV with set-point/servo targeting | Mode not offered | Pressure Control Assist Control with Automatic Tube Compensation | Pressure Controlled Continuous Mandatory Ventilation with Tube Resistance Compensation | Mode not offered | Assist Control Pressure Control with Tube Compensation
PC-IMVa,a | Pressure control IMV with adaptive targeting | Mode not offered | Mode not offered | Mode not offered | Automode (Pressure Regulated Volume Control to Volume Support) | Mode not offered
PC-IMVa,r | Pressure control IMV with adative and servo targeting | SIMV Volume Control Plus with Tube Compensation | Mode not offered | Mode not offered | Mode not offered | Mode not offered
PC-IMVas,s | Pressure control IMV with adative/set-point and set-point targeting | Mode not offered | Mode not offered | Mode not offered | Mode not offered | SIMV Pressure Regulated Volume Control with Pressure Minimum
PC-IMVa,s | Pressure control IMV with adative and set-point targeting | SIMV Volume Control Plus with PS | Volume Control SIMV with AutoFlow/Volume Guarantee | Adaptive Pressure Ventilation SIMV | SIMV Pressure Regulated Volume Control | Bilevel Airway Pressure Ventilation Volume Guarantee
SIMV Pressure Regulated Volume Control
PC-IMVas,sr | Pressure control IMV with adaptive/set-point, and set-point/servo targeting | Mode not offered | Mode not offered | Mode not offered | Mode not offered | Mode not offered
PC-IMVar,sr | Pressure control IMV with adaptive/servo and set-point/servo targeting | Mode not offered | Volume Control SIMV with AutoFlow/Volume Guarantee and Automtic Tube Compensation | Adaptive Pressure Ventilation SIMV with Tube Resistance Compensation | Mode not offered | Bilevel Airway Pressure Ventilation Volume Guarantee with Tube Compensation
SIMV Pressure Regulated Volume Control with Tube Compensation
Volume Control Mandatory Minute Volume with AutoFlow/Volume Guarantee and Automatic Tube Compensation
PC-IMVoi,oi | Pressure control IMV with optimal and intelligent targeting | Mode not offered | Mode not offered | Adaptive Support Ventilation | Mode not offered | Mode not offered
IntelliVent-Adaptive Support Ventilation
PC-IMVoir,oir | Pressure control IMV with optimal, intelligent, and servo targeting | Mode not offered | Mode not offered | Adaptive Support Ventilation with Tube Resistance Compensation | Mode not offered | Mode not offered
IntelliVent-Adaptive Support Ventilation with Tube Resistance Compensation
PC-IMVs,r | Pressure control IMV with set-point and servo targeting | SIMV-Pressure Control with Tube Compensation | Mode not offered | Mode not offered | Mode not offered | SIMV-Pressure Control with Tube Compensation
|  | Bilevel with Tube Compensation |  |  |  | Bilevel with Tube Compensation
PC-IMVsa,s | Pressure control IMV with set-point/adaptive and set-point targeting | SIMV-Pressure Control with PS | Pressure Control Continuous Mandatory Ventilation (not patient triggered) | Pressure SIMV | SIMV (Pressure Control) | SIMV Pressure Control
|  | Bilevel with PS | Pressure Control Airway Pressure Release Ventilation |  | Bi-Vent | Airway Pressure Release Ventilation
|  |  |  | Noninvasive Ventilation - Spontaneous Timed | Automode (Pressure Control to PS) | Bilevel Airway Pressure Ventilation
Nasal CPAP - PS
Airway Pressure Release Ventilation
|  |  | Pressure Control Pressure Support Ventilation | DuoPAP
PC-IMVsr,sr | Pressure control CSV with set-point/servo targeting | Mode not offered | Pressure Control Continuous Mandatory Ventilation with Automatic Tube Compensation | Pressure SIMV with Tube Resistance Compensation | Mode not offered | SIMV Pressure Control with Tube Compensation
|  |  | Pressure Control SIMV with Autmatic Tube Compensation | Airway Pressure Release Ventilation with Tube Resistance Compensation |  | Bilevel Airway Pressure Ventilation with Tube Compensation
Pressure Control Airway Pressure Release Ventilation with Automatic Tube Compensation
Pressure Control Pressure Support Ventilation with Automatic Tube Compensation
PC-CSVa | Pressure control CSV with adaptive targeting | Spont Volume Support | Spontaneous CPAP/Volume Support | Mode not offered | Volume Support | Volume Support
PC-CSVar | Pressure control CSV with adaptive/servo targeting | Mode not offered | Spontaneous CPAP/Volume Support with Automatic Tube Compensation | Mode not offered | Mode not offered | Volume Support with Tube Compensation
PC-CSVb | Pressure control CSV with bio-variable targeting | Mode not offered | Spontaneous CPAP/Variable Pressure Support | Mode not offered | Mode not offered | Mode not offered
PC-CSVbr | Pressure control CSV with bio-variable/servo targeting | Mode not offered | Spontaneous CPAP/Variable Pressure Support with Automatic Tube Compensation | Mode not offered | Mode not offered | Mode not offered
PC-CSVi | Pressure control CSV with intelligent targeting | Mode not offered | SmartCare | Mode not offered | Mode not offered | Mode not offered
PC-CSVr | Pressure control CSV with servo targeting | Spont Tube Compensation | Spontaneous Proportional Pressure Support | Mode not offered | Neurally Adjusted Ventilatory Assist (NAVA) | Mode not offered
|  | Spont Proportional Assist |  | Mode not offered | Noninvasive NAVA
PC-CSVs | Pressure control CSV with set-point | Spont PS | Spontaneous CPAP/Pressure Support | Spontaneous | Spontaneous/CPAP | Continuous Positive Airway Pressure/Pressure Support
|  |  | Spontaneous CPAP |  | Pressure Support | Noninvasive Ventilation
|  |  |  | Noninvasive Ventilation | Nasal CPAP | Spontaneous Breathing Trial
PC-CSVsr | Pressure control CSV with set-point/servo targeting | Mode not offered | Spontaneous CPAP/Pressure Support with Automatic Tube Compensation | Spontaneous with Tube Resistance Compensation | Mode not offered | Continuous Positive Airway Pressure/Pressure Support with Tube Compensation
Spontaneous Breathing Trial with Tube Compensation
CONTROL VARIABLES |  | MODE NAME CONVENTIONS
VC = volume control |  | A/C = assist/control
PC = pressure control |  | SIMV = synchronized intermittent mandatory ventilation
TC = time control |  | NAVA = Neurally Adjusted Ventilatory Assist
BREATH SEQUENCES |  | PAV = Proportional Assist Ventilation
CMV = continuous mandatory ventilation |  | CPAP = continuous positive airway pressure
IMV = IMV |  | PS = Pressure Support
CSV = continuous spontaneous ventilation
TARGETING SCHEMES
s = set-point
d = dual
b = bio-variable
r = servo
a = adaptive
o = optimal
i = intelligent
Menu

## ModeMap
|  |  | pressure | CMV | N/A | N/A
|  |  | volume | IMV | set-point | set-point
|  |  | time | CSV | set-point/servo | set-point/servo
[5 empty] | dual | servo
[5 empty] | dual/set-point | dual
[5 empty] | servo | adaptive
[5 empty] | bio-variable | optimal
[5 empty] | bio-variable/servo | intelligent
[5 empty] | adaptive | optimal/intelligent
[5 empty] | adaptive/servo | optimal/intelligent/servo
[5 empty] | optimal | dual/set-point
[5 empty] | intelligent | dual//adaptive
optimal/intelligent
optimal/intelligent/servo
dual/adaptive
set-point/adaptive
dual/set-point
unhide rows to see lookup tables
Ventilator Mode Map V10.0 © 2012 by Mandu Press Ltd
MODE CLASSIFICATION
[5 empty] | Primary | Secondary
|  |  | Control | Breath | Targeting | Targeting
MANUFACTURER | MODEL | MODE NAME (as specified in operator's manual) | Variable | Sequence | Scheme | Scheme | TAG
Airon | mini | IMV + CPAP | pressure | IMV | dual | set-point | PC-IMVd,s
Airon | mini | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Airon | model A and B | Volume Limited IMV + CPAP | volume | IMV | set-point | set-point | VC-IMVs,s
Airon | model A and B | Pressure Limited IMV + CPAP | volume | IMV | dual | set-point | VC-IMVd,s
Bio-Med | Crossvent-4+ | A/C | volume | CMV | set-point | N/A | VC-CMVs
Bio-Med | Crossvent-4+ | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Bio-Med | Crossvent-4+ | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Bunnell | Life Pulse | High frequency jet ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | 3100B | High frequency oscillatory ventilation | time | IMV | set-point | set-point | TC-IMVs,s
CareFusion | Avea | Volume A/C | volume | CMV | set-point | N/A | VC-CMVs
CareFusion | Avea | Volume SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
CareFusion | Avea | Volume SIMV with Artificial Airway Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
CareFusion | Avea | Volume A/C with Demand Flow | volume | IMV | dual | dual | VC-IMVd,d
CareFusion | Avea | Volume SIMV with Demand Flow | volume | IMV | dual | set-point | VC-IMVd,s
CareFusion | Avea | Volume SIMV with Demand Flow and Artificial Airway Compensation | volume | IMV | dual | set-point/servo | VC-IMVd,sr
CareFusion | Avea | Pressure A/C | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | Avea | Time Cycled Pressure Limited A/C | pressure | CMV | dual | N/A | PC-CMVd
CareFusion | Avea | Pressure A/C with Machine Volume | pressure | CMV | dual | N/A | PC-CMVd
CareFusion | Avea | Pressure A/C with Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Avea | Time Cycled Pressure Limited A/C with Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Avea | Volume A/C with Vsync | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Avea | Pressure Regulated Volume Control A/C | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Avea | Pressure A/C with Flow Cycle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | Pressure A/C with Flow Cycle and Artificial Ariway Compensation | pressure | IMV | set-point | set-point/servo | PC-IMVs,sr
CareFusion | Avea | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | Pressure SIMV with Artificial Ariway Compensation | pressure | IMV | set-point | set-point/servo | PC-IMVs,sr
CareFusion | Avea | CPAP/Pressure Support Ventilation with Volume Limit | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | CPAP/Pressure Support Ventilation with Volume Limit and Artificial Airway Compensation | pressure | IMV | set-point | set-point/servo | PC-IMVs,sr
CareFusion | Avea | Infant Nasal IMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | Infant Nasal IMV with Artificial Airway Compensation | pressure | IMV | set-point | set-point/servo | PC-IMVs,sr
CareFusion | Avea | Airway Pressure Release Ventilation/Biphasic | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | Time Cycled Pressure Limited A/C with Flow Cycle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Avea | Time Cycled Pressure Limited SIMV with Artificial Airway Compensation | pressure | IMV | set-point | set-point/servo | PC-IMVs,sr
CareFusion | Avea | Time Cycled Pressure Limited SIMV | pressure | IMV | dual | set-point | PC-IMVd,s
CareFusion | Avea | Pressure SIMV with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Pressure SIMV with Volume Guarantee and Artificial Airway Compensation | pressure | IMV | adaptive | set-point/servo | PC-IMVa,sr
CareFusion | Avea | Time Cycled Pressure Limited A/C with Flow Cycle and Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Time Cycled Pressure Limited SIMV with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Time Cycled Pressure Limited SIMV with Volume Guarantee and Artificial Airway Compensation | pressure | IMV | adaptive | set-point/servo | PC-IMVa,sr
CareFusion | Avea | Volume A/C with Vsync and Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Volume SIMV with Vsync | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Volume SIMV with Vsync and Artificial Airway Compensation | pressure | IMV | adaptive | set-point/servo | PC-IMVa,sr
CareFusion | Avea | Pressure Regulated Volume Control A/C with Flow Cycle | pressure | IMV | adaptive | adaptive | PC-IMVa,a
CareFusion | Avea | Pressure Regulated Volume Control SIMV with Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Pressure Regulated Volume Control SIMV | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Avea | Pressure Regulated Volume Control SIMV with Artificial Ariway Compensation | pressure | IMV | adaptive | set-point/servo | PC-IMVa,sr
CareFusion | Avea | Time Cycled Pressure Limited A/C with Flow Cycle and Volume Guarantee and Artificial Airway Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
CareFusion | Avea | Volume A/C with Vsync and Flow Cycle and Artificial Airway Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
CareFusion | Avea | Pressure Regulated Volume Control A/C with Flow Cycle and Artificial Airway Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
CareFusion | Avea | Pressure Regulated Volume Control SIMV with Flow Cycle and Artificial Airway Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
CareFusion | Avea | Time Cycled Pressure Limited A/C with Flow Cycle and Artificial Airway Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
CareFusion | Avea | CPAP/Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | Avea | CPAP/Pressure Support Ventilation and Artificial Airway Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
CareFusion | Bird Mark 7 | Not Named | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Bird Mark 7 | Not Named | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | EnVe | Volume A/C | volume | CMV | set-point | N/A | VC-CMVs
CareFusion | EnVe | Volume SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
CareFusion | EnVe | Pressure A/C | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | EnVe | Pressure Regulated Volume Control A/C | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | EnVe | Noninvasive Positive Pressure Ventilation Pressure | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | EnVe | Pressure A/C with Flow Cycle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | EnVe | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | EnVe | Noninvasive Positive Pressure Ventilation Pressure with Flow Cylcle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | EnVe | Pressure Regulated Volume Control A/C with Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | EnVe | Pressure Regulated Volume Control SIMV | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | EnVe | CPAP/PSV | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | EnVe | Noninvasive Positive Pressure Ventilation CPAP/PSV | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | EnVe | CPAP + Volume Targeted Pressure Support Ventilation | pressure | CSV | adaptive | N/A | PC-CSVa
CareFusion | Infant Flow SiPAP | BiPhasic | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Infant Flow SiPAP | BiPhasic tr | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Infant Flow SiPAP | Nasal CPAP | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | LTV 1200 | Volume A/C | volume | CMV | set-point | N/A | VC-CMVs
CareFusion | LTV 1200 | Volume SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
CareFusion | LTV 1200 | Pressure A/C | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | LTV 1200 | Pressure A/C with Flow Termination | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | LTV 1200 | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | LTV 1200 | CPAP/PSV | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | LTV 1200 | Noninvasive Positive Pressure Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | ReVel | A/C + Volume | volume | CMV | set-point | N/A | VC-CMVs
CareFusion | ReVel | SIMV + Volume | volume | IMV | set-point | set-point | VC-IMVs,s
CareFusion | ReVel | A/C + Pressure | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | ReVel | A/C + Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | ReVel | Noninvasive Positive Pressure Ventilation Pressure | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | ReVel | Noninvasive Positive Pressure Ventilation Pressure with Flow Cylcle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | ReVel | A/C + Pressure with Flow Cycle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | ReVel | SIMV + Pressure | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | ReVel | A/C + PRVC with Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | ReVel | SIMV + PRVC | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | ReVel | CPAP + PS | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | ReVel | Noninvasive Positive Pressure Ventilation CPAP/Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | ReVel | CPAP + Volume Targeted Pressure Support | pressure | CSV | adaptive | N/A | PC-CSVa
CareFusion | Vela | Volume A/C | volume | CMV | set-point | N/A | VC-CMVs
CareFusion | Vela | Volume SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
CareFusion | Vela | Pressure A/C | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | Vela | Noninvasive Positive Pressure Ventilation A/C | pressure | CMV | set-point | N/A | PC-CMVs
CareFusion | Vela | Pressure A/C with Assured Volume | pressure | CMV | dual | N/A | PC-CMVd
CareFusion | Vela | Volume A/C with Vsync | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Vela | Pressure Regulated Volume Control A/C | pressure | CMV | adaptive | N/A | PC-CMVa
CareFusion | Vela | Pressure A/C with Flow Cycle | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Vela | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Vela | Noninvasive Positive Pressure Ventilation SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Vela | Airway Pressure Release Ventilation/Biphasic | pressure | IMV | set-point | set-point | PC-IMVs,s
CareFusion | Vela | Volume A/C with Vsync and Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Vela | Volume SIMV with Vsync | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Vela | Pressure Regulated Volume Control A/C with Flow Cycle | pressure | IMV | adaptive | adaptive | PC-IMVa,a
CareFusion | Vela | Pressure Regulated Volume Control SIMV with Flow Cycle | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Vela | Pressure Regulated Volume Control SIMV | pressure | IMV | adaptive | set-point | PC-IMVa,s
CareFusion | Vela | CPAP/Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
CareFusion | Vela | Noninvasive Positive Pressure Ventilation CPAP PS | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | Newport e360T | Volume Control A/C Mandatory Ventilation | volume | CMV | set-point | N/A | VC-CMVs
Covidien | Newport e360T | Volme Control SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | Newport e360T | Pressure Control A/C Mandatory Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | Newport e360T | Biphasic Pressure Release A/C Mandatory Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | Newport e360T | Volume Target Pressure Control A/C Mandatory Ventilation | pressure | CMV | adaptive | N/A | PC-CMVa
Covidien | Newport e360T | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | Newport e360T | Biphasic Pressure Release SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | Newport e360T | Volume Target Pressure Control SIMV | pressure | IMV | adaptive | adaptive | PC-IMVa,a
Covidien | Newport e360T | Pressure Control Spont | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | Newport e360T | Volume Control Spont | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | Newport e360T | Volume Target Pressure Control Spont | pressure | CSV | adaptive | N/A | PC-CSVa
Covidien | Newport HT70 Plus | A/C Mandatory Ventilation Volume Control | volume | CMV | set-point | N/A | VC-CMVs
Covidien | Newport HT70 Plus | SIMV Volume Control | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | Newport HT70 Plus | A/C Mandatory Ventilation Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | Newport HT70 Plus | SIMV Pressure Control | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | Newport HT70 Plus | Spont | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | PB 540 | Volume A/C | volume | CMV | set-point | N/A | VC-CMVs
Covidien | PB 540 | Volume SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | PB 540 | Pressure A/C | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | PB 540 | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 540 | Pressure Support Ventilation/CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | PB 7200 | Continuous Mandatory Ventilation | volume | CMV | set-point | N/A | VC-CMVs
Covidien | PB 7200 | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | PB 7200 | Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | PB 7200 | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 7200 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | PB 840 | A/C Volume Control | volume | CMV | set-point | N/A | VC-CMVs
Covidien | PB 840 | SIMV Volume Control with Pressure Support | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | PB 840 | SIMV Volume Control with Tube Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
Covidien | PB 840 | A/C Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | PB 840 | A/C Volume Control Plus | pressure | CMV | adaptive | N/A | PC-CMVa
Covidien | PB 840 | SIMV-Pressure Control with Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 840 | SIMV-Pressure Control with Tube Compensation | pressure | IMV | set-point | servo | PC-IMVs,r
Covidien | PB 840 | Bilevel with Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 840 | Bilevel with Tube Compensation | pressure | IMV | set-point | servo | PC-IMVs,r
Covidien | PB 840 | SIMV Volume Control Plus with Pressure Support | pressure | IMV | adaptive | set-point | PC-IMVa,s
Covidien | PB 840 | SIMV Volume Control Plus with Tube Compensation | pressure | IMV | adaptive | servo | PC-IMVa,r
Covidien | PB 840 | Spont Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | PB 840 | Spont Tube Compensation | pressure | CSV | servo | N/A | PC-CSVr
Covidien | PB 840 | Spont Proportional Assist | pressure | CSV | servo | N/A | PC-CSVr
Covidien | PB 840 | Spont Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Covidien | PB 980 | A/C Volume Control | volume | CMV | set-point | N/A | VC-CMVs
Covidien | PB 980 | SIMV Volume Control with Pressure Support | volume | IMV | set-point | set-point | VC-IMVs,s
Covidien | PB 980 | SIMV Volume Control with Tube Compensation | volume | IMV | set-point | servo | VC-IMVs,r
Covidien | PB 980 | A/C Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Covidien | PB 980 | A/C Volume Control Plus | pressure | CMV | adaptive | N/A | PC-CMVa
Covidien | PB 980 | SIMV-Pressure Control with Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 980 | SIMV-Pressure Control with Tube Compensation | pressure | IMV | set-point | servo | PC-IMVs,r
Covidien | PB 980 | Bilevel with Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Covidien | PB 980 | Bilevel with Tube Compensation | pressure | IMV | set-point | servo | PC-IMVs,r
Covidien | PB 980 | SIMV Volume Control Plus with Pressure Support | pressure | IMV | adaptive | set-point | PC-IMVa,s
Covidien | PB 980 | SIMV Volume Control Plus with Tube Compensation | pressure | IMV | adaptive | servo | PC-IMVa,r
Covidien | PB 980 | Spont Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Covidien | PB 980 | Spont Tube Compensation | pressure | CSV | servo | N/A | PC-CSVr
Covidien | PB 980 | Spont Proportional Assist | pressure | CSV | servo | N/A | PC-CSVr
Covidien | PB 980 | Spont Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Dräger | Babylog VN500 | Pressure Control A/C | pressure | CMV | set-point | N/A | PC-CMVs
Dräger | Babylog VN500 | Pressure Control A/C with Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Babylog VN500 | Pressure Control A/C with Volume Guarantee and Automatic Tube Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Dräger | Babylog VN500 | Pressure Control A/C with Automatic Tube Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
Dräger | Babylog VN500 | Pressure Control Continuous Mandatory Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Babylog VN500 | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Babylog VN500 | Pressure Control Pressure Support Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Babylog VN500 | Pressure Control Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Babylog VN500 | Pressure Control Mandatory Minute Volume Ventilation with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Babylog VN500 | Pressure Control Continuous Mandatory Ventilation with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Babylog VN500 | Pressure Control SIMV with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Babylog VN500 | Pressure Control Pressure Support Ventilation with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Babylog VN500 | Pressure Control Continuous Mandatory Ventilation with Volume Guarantee and Automatic Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Babylog VN500 | Pressure Control SIMV with Volume Guarantee and Automatic Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Babylog VN500 | Pressure Control Pressure Support Ventilation with Volume Guarantee and Automatic Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Babylog VN500 | Pressure Control Continuous Mandatory Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Babylog VN500 | Pressure Control SIMV with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Babylog VN500 | Pressure Control Pressure Support Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Babylog VN500 | Pressure Control Airway Pressure Release Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Babylog VN500 | Spontaneous CPAP/Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Dräger | Babylog VN500 | Spontaneous Proportional Pressure Support | pressure | CSV | servo | N/A | PC-CSVr
Dräger | Babylog VN500 | Spontaneous Proportional Pressure Support with Automatic Tube Compensation | pressure | CSV | servo | N/A | PC-CSVr
Dräger | Babylog VN500 | Spontaneous CPAP/Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Dräger | Babylog VN500 | Spontaneous CPAP/Volume Support wth Automatic Tube Compensation | pressure | CSV | adaptive/servo | N/A | PC-CSVar
Dräger | Babylog VN500 | Spontaneous CPAP/Pressure Support with Automatic Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
Dräger | Carina | Pressure Control A/C | pressure | CMV | set-point | N/A | PC-CMVs
Dräger | Carina | Volume Control A/C with AutoFlow | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Carina | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Carina | Volume Control SIMV with AutoFlow | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Carina | SPN-CPAP/PS | pressure | CSV | set-point | N/A | PC-CSVs
Dräger | Carina | SPN-CPAP/PS with Volume Guarantee | pressure | CSV | adaptive | N/A | PC-CSVa
Dräger | Evita Infinity V500 | Volume Control Continuous Mandatory Ventilation (not patient triggered) | volume | CMV | set-point | N/A | VC-CMVs
Dräger | Evita Infinity V500 | Volume Control Assist Control (patient triggered) | volume | CMV | set-point | N/A | VC-CMVs
Dräger | Evita Infinity V500 | Volume Control Continuous Mandatory Ventilation with Pressure Limitation | volume | CMV | dual | N/A | VC-CMVd
Dräger | Evita Infinity V500 | Volume Control A/C with Pressure Limitation | volume | CMV | dual | N/A | VC-CMVd
Dräger | Evita Infinity V500 | Volume Control SIMV with AutoFlow | volume | IMV | set-point | set-point | VC-IMVs,s
Dräger | Evita Infinity V500 | Volume Control SIMV with Automatic Tube Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
Dräger | Evita Infinity V500 | Volume Control SIMV with Pressure Limitation | volume | IMV | dual | set-point | VC-IMVd,s
Dräger | Evita Infinity V500 | Volume Control SIMV with Pressure Limitation and Automatic Tube Compensation | volume | IMV | dual | set-point/servo | VC-IMVd,sr
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume Ventilation | volume | IMV | adaptive | set-point | VC-IMVa,s
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume Ventilation with Automatic Tube Compensation | volume | IMV | adaptive | set-point/servo | VC-IMVa,sr
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume with AutoFlow/Volume Guarantee | pressure | IMV | adaptive | set-point | VC-IMVa,s
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume with Pressure Limitation | volume | IMV | dual/adaptive | set-point | VC-IMVda,s
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume with Pressure Limitation and Automatic Tube Compensation | volume | IMV | dual/adaptive | set-point/servo | VC-IMVda,sr
Dräger | Evita Infinity V500 | Pressure Control A/C (patient triggered) | pressure | CMV | set-point | N/A | PC-CMVs
Dräger | Evita Infinity V500 | Volume Control Continuous Mandatory Ventilation with AutoFlow/Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Evita Infinity V500 | Volume Control A/C with AutoFlow/Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Evita Infinity V500 | Volume Control Continuous Mandatory Ventilation with AutoFlow/Volume Guarantee and Automatic Tube Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Dräger | Evita Infinity V500 | Volume ControlA/C with AutoFlow/Volume Guarantee and Automatic Tube Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Dräger | Evita Infinity V500 | Pressure Control A/C with Automatic Tube Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
Dräger | Evita Infinity V500 | Pressure Control Continuous Mandatory Ventilation (not patient triggered) | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita Infinity V500 | Pressure Control SIMV Plus | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita Infinity V500 | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita Infinity V500 | Pressure Control Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita Infinity V500 | Pressure Control Pressure Support Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita Infinity V500 | Volume Control SIMV with AutoFlow | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Evita Infinity V500 | Volume Control Mandatory Minute Volume with AutoFlow/Volume Guarantee and Automatic Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Evita Infinity V500 | Volume Control SIMV with AutoFlow and Automtic Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Evita Infinity V500 | Pressure Control Continuous Mandatory Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita Infinity V500 | Pressure Control SIMV with Autmatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita Infinity V500 | Pressure Control Airway Pressure Release Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita Infinity V500 | Pressure Control Pressure Support Ventilation with Automatic Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita Infinity V500 | Spontaneous CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Dräger | Evita Infinity V500 | Spontaneous CPAP/Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Dräger | Evita Infinity V500 | Spontaneous Proportional Pressure Support | pressure | CSV | servo | N/A | PC-CSVr
Dräger | Evita Infinity V500 | Spontaneous CPAP/Variable Pressure Support | pressure | CSV | bio-variable | N/A | PC-CSVb
Dräger | Evita Infinity V500 | Spontaneous CPAP/Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Dräger | Evita Infinity V500 | SmartCare | pressure | CSV | intelligent | N/A | PC-CSVi
Dräger | Evita Infinity V500 | Spontaneous CPAP/Volume Support with Automatic Tube Compensation | pressure | CSV | adaptive/servo | N/A | PC-CSVar
Dräger | Evita Infinity V500 | Spontaneous CPAP/Variable Pressure Support with Automatic Tube Compensation | pressure | CSV | bio-variable/servo | N/A | PC-CSVbr
Dräger | Evita Infinity V500 | Spontaneous CPAP/Pressure Support with Automatic Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
Dräger | Evita XL | Continuous Mandatory Ventilation | volume | CMV | set-point | N/A | VC-CMVs
Dräger | Evita XL | Continuous Mandatory Ventilation with Pressure Limited Ventilation | volume | CMV | dual | N/A | VC-CMVd
Dräger | Evita XL | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Dräger | Evita XL | SIMV with Automatic Tube Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
Dräger | Evita XL | SIMV with Pressure Limited Ventilation | volume | IMV | dual | set-point | VC-IMVd,s
Dräger | Evita XL | SIMV with Pressure Limited Ventilation and Automatic Tube Compensation | volume | IMV | dual | set-point/servo | VC-IMVd,sr
Dräger | Evita XL | Mandatory Minute Volume Ventilation | volume | IMV | adaptive | set-point | VC-IMVa,s
Dräger | Evita XL | Mandatory Minute Volume Ventilation with Automatic Tube Compensation | volume | IMV | adaptive | set-point/servo | VC-IMVa,sr
Dräger | Evita XL | Mandatory Minute Volume with Pressure Limited Ventilation | volume | IMV | dual/adaptive | set-point | VC-IMVda,s
Dräger | Evita XL | Mandatory Minute Volume with Pressure Limited Ventilationand Automatic Tube Compensation | volume | IMV | dual/adaptive | set-point/servo | VC-IMVda,sr
Dräger | Evita XL | Pressure Controlled Ventilation Plus Assisted | pressure | CMV | set-point | N/A | PC-CMVs
Dräger | Evita XL | Continuous Mandatory Ventilation with AutoFlow | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Evita XL | Continuous Mandatory Ventilation with AutoFlow and Tube Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Dräger | Evita XL | Pressure Controlled Ventilation Plus/Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita XL | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Evita XL | Mandatory Minute Volume with AutoFlow | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Evita XL | SIMV with AutoFlow | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Evita XL | Mandatory Minute Volume with AutoFlow and Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Evita XL | SIMV with AutoFlow and Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Dräger | Evita XL | Pressure Controlled Ventilation Plus/Pressure Support and Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita XL | Airway Pressure Release Ventilation with Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Dräger | Evita XL | CPAP/Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Dräger | Evita XL | SmartCare | pressure | CSV | intelligent | N/A | PC-CSVi
Dräger | Evita XL | CPAP/Pressure Support with Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
Dräger | Oxylog 3000 plus | Volume Control A/C Volume ControlCMV | volume | CMV | set-point | N/A | VC-CMVs
Dräger | Oxylog 3000 plus | Volume Control SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Dräger | Oxylog 3000 plus | Volume Control A/C with AutoFlow | pressure | CMV | adaptive | N/A | PC-CMVa
Dräger | Oxylog 3000 plus | Pressure Control IMV Plus | pressure | IMV | set-point | set-point | PC-IMVs,s
Dräger | Oxylog 3000 plus | Volume Control SIMV with AutoFlow | pressure | IMV | adaptive | set-point | PC-IMVa,s
Dräger | Oxylog 3000 plus | SpnCPAP/PS | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | Engström Carestation | Volume Controlled Ventilation | volume | CMV | set-point | N/A | VC-CMVs
GE Healthcare | Engström Carestation | Volume Controlled Ventilation with Pressure Limit | volume | CMV | dual | N/A | VC-CMVd
GE Healthcare | Engström Carestation | SIMV Volume Controlled | volume | IMV | set-point | set-point | VC-IMVs,s
GE Healthcare | Engström Carestation | SIMV Volume Controlled with Airway Resistance Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
GE Healthcare | Engström Carestation | SIMV Volume Controlled with Pressure Limit | volume | IMV | dual | set-point | VC-IMVd,s
GE Healthcare | Engström Carestation | SIMV Volume Controlled with Pressure Limit and Airway Resistance Compensation | volume | IMV | dual | set-point/servo | VC-IMVd,sr
GE Healthcare | Engström Carestation | Pressure Controlled Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
GE Healthcare | Engström Carestation | Pressure Controlled Ventilation-Volume Guaranteed | pressure | CMV | adaptive | N/A | PC-CMVa
GE Healthcare | Engström Carestation | Pressure Controlled Ventilation-Volume Guaranteed with Airway Resistance Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
GE Healthcare | Engström Carestation | Pressure Controlled Ventilation with Airway Resistance Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
GE Healthcare | Engström Carestation | SIMV Pressure Controlled | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | Engström Carestation | Bilevel Airway Pressure Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | Engström Carestation | Bilevel Airway Pressure Ventilation-Volume Guaranteed | pressure | IMV | adaptive | set-point | PC-IMVa,s
GE Healthcare | Engström Carestation | SIMV Pressure Controlled Volume Guaranteed | pressure | IMV | adaptive | set-point | PC-IMVa,s
GE Healthcare | Engström Carestation | Bilevel Airway Pressure Ventilation-Volume Guaranteed with Airway Resistance Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
GE Healthcare | Engström Carestation | SIMV Pressure Controlled Volume Guaranteed with Airway Resistance Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
GE Healthcare | Engström Carestation | SIMV Pressure Controlled with Airway Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
GE Healthcare | Engström Carestation | Bilevel Airway Pressure Ventilation with Airway Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
GE Healthcare | Engström Carestation | CPAP/Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | Engström Carestation | CPAP/Pressure Support with Airway Resistance Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
GE Healthcare | Engström Carestation | Volume Guarantee Pressure Support | pressure | CSV | adaptive | N/A | PC-CSVa
GE Healthcare | Engström Carestation | Volume Guarantee Pressure Support with Airway Resistance Compensation | pressure | CSV | adaptive/servo | N/A | PC-CSVar
GE Healthcare | Engström Carestation | CPAP/Pressure Support with Airway Resistance Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
GE Healthcare | CARESCAPE R860 | A/C Volume Control | volume | CMV | set-point | N/A | VC-CMVs
GE Healthcare | CARESCAPE R860 | A/C Volume Control with Pressure Limit | volume | CMV | dual | N/A | VC-CMVd
GE Healthcare | CARESCAPE R860 | SIMV Volume Control | volume | IMV | set-point | set-point | VC-IMVs,s
GE Healthcare | CARESCAPE R860 | SIMV Volume Control with Tube Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
GE Healthcare | CARESCAPE R860 | SIMV Volume Control with Pressure Limit | volume | IMV | dual | set-point | VC-IMVd,s
GE Healthcare | CARESCAPE R860 | SIMV Volume Control with Pressure Limit and Tube Compensation | volume | IMV | dual | set-point/servo | VC-IMVd,sr
GE Healthcare | CARESCAPE R860 | A/C Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
GE Healthcare | CARESCAPE R860 | A/C Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
GE Healthcare | CARESCAPE R860 | A/C Pressure Regulated Volume Control with Tube Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
GE Healthcare | CARESCAPE R860 | A/C Pressure Control with Tube Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
GE Healthcare | CARESCAPE R860 | SIMV Pressure Control | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | CARESCAPE R860 | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | CARESCAPE R860 | Bilevel Airway Pressure Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | CARESCAPE R860 | Bilevel Airway Pressure Ventilation Volume Guaranteed | pressure | IMV | adaptive | set-point | PC-IMVa,s
GE Healthcare | CARESCAPE R860 | SIMV Pressure Regulated Volume Control | pressure | IMV | adaptive | set-point | PC-IMVa,s
GE Healthcare | CARESCAPE R860 | Bilevel Airway Pressure Ventilation Volume Guaranteed with Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
GE Healthcare | CARESCAPE R860 | SIMV Pressure Regulated Volume Control with Tube Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
GE Healthcare | CARESCAPE R860 | SIMV Pressure Control with Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
GE Healthcare | CARESCAPE R860 | Bilevel Airway Pressure Ventilation with Tube Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
GE Healthcare | CARESCAPE R860 | CPAP/Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | CARESCAPE R860 | CPAP/Pressure Support with Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
GE Healthcare | CARESCAPE R860 | Noninvasive Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | CARESCAPE R860 | Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
GE Healthcare | CARESCAPE R860 | Volume Support with Tube Compensation | pressure | CSV | adaptive/servo | N/A | PC-CSVar
GE Healthcare | CARESCAPE R860 | CPAP/Pressure Support with Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
GE Healthcare | CARESCAPE R860 | Spontaneous Breathing Trial | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | CARESCAPE R860 | Spontaneous Breathing Trial with Tube Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
GE Healthcare | iVent101 | A/C Volume Control | volume | CMV | set-point | N/A | VC-CMVs
GE Healthcare | iVent101 | A/C Volume Control with Adaptive Peak Flow | volume | CMV | adaptive | N/A | VC-CMVa
GE Healthcare | iVent101 | SIMV Volume Control | volume | IMV | set-point | set-point | VC-IMVs,s
GE Healthcare | iVent101 | SIMV Volume Control with Adaptive Peak Flow | volume | IMV | adaptive | set-point | VC-IMVa,s
GE Healthcare | iVent101 | A/C Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
GE Healthcare | iVent101 | A/C Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
GE Healthcare | iVent101 | SIMV Pressure Control | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | iVent101 | Adaptive Bi-Level | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | iVent101 | SIMV Pressure Regulated Volume Control | pressure | IMV | adaptive | set-point | PC-IMVa,s
GE Healthcare | iVent101 | CPAP/Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
GE Healthcare | iVent201 | A/C Volume Control | volume | CMV | set-point | N/A | VC-CMVs
GE Healthcare | iVent201 | A/C Volume Control with Adative Flow and I-Time | volume | CMV | adaptive | N/A | VC-CMVa
GE Healthcare | iVent201 | SIMV Volume Control | volume | IMV | set-point | set-point | VC-IMVs,s
GE Healthcare | iVent201 | SIMV Volume Control with Adaptive Flow and I-Time | volume | IMV | adaptive | set-point | VC-IMVa,s
GE Healthcare | iVent201 | A/C Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
GE Healthcare | iVent201 | SIMV Pressure Control | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | iVent201 | Adaptive Bi-Level | pressure | IMV | set-point | set-point | PC-IMVs,s
GE Healthcare | iVent201 | CPAP/PSV | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | C3 | Pressure Controlled Ventilation Plus | pressure | CMV | set-point | N/A | PC-CMVs
Hamilton | C3 | SIMV Plus | pressure | CMV | adaptive | N/A | PC-CMVa
Hamilton | C3 | Pressure Controlled SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | C3 | Noninvasive Ventilation - Spontaneous Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | C3 | Nasal CPAP - Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | C3 | Duo Positive Airway Pressure | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | C3 | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | C3 | SIMV Plus | pressure | IMV | adaptive | set-point | PC-IMVa,s
Hamilton | C3 | Adaptive Support Ventilation | pressure | IMV | optimal/intelligent | optimal/intelligent | PC-IMVoi,oi
Hamilton | C3 | Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | C3 | Noninvasive Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | G5 | SIMV | volume | CMV | set-point | N/A | VC-CMVs
Hamilton | G5 | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Hamilton | G5 | SIMV with Tube Resistance Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
Hamilton | G5 | Pressure Controlled Continuous Mandatory Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Hamilton | G5 | Adaptive Pressure Ventilation Continuous Mandatory Ventilation | pressure | CMV | adaptive | N/A | PC-CMVa
Hamilton | G5 | Adaptive Pressure Ventilation Continuous Mandatory Ventilation with Tube Resistance Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Hamilton | G5 | Pressure Controlled Continuous Mandatory Ventilation with Tube Resistance Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
Hamilton | G5 | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | G5 | Noninvasive Ventilation - Spontaneous Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | G5 | Nasal CPAP - Pressure Support | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | G5 | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | G5 | DuoPAP | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | G5 | Adaptive Pressure Ventilation SIMV | pressure | IMV | adaptive | set-point | PC-IMVa,s
Hamilton | G5 | Adaptive Pressure Ventilation SIMV with Tube Resistance Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Hamilton | G5 | Adaptive Support Ventilation | pressure | IMV | optimal/intelligent | optimal/intelligent | PC-IMVoi,oi
Hamilton | G5 | IntelliVent-Adaptive Support Ventilation | pressure | IMV | optimal/intelligent | optimal/intelligent | PC-IMVoi,oi
Hamilton | G5 | Adaptive Support Ventilation with Tube Resistance Compensation | pressure | IMV | optimal/intelligent/servo | optimal/intelligent/servo | PC-IMVoir,oir
Hamilton | G5 | IntelliVent-Adaptive Support Ventilation with Tube Resistance Compensation | pressure | IMV | optimal/intelligent/servo | optimal/intelligent/servo | PC-IMVoir,oir
Hamilton | G5 | Pressure SIMV with Tube Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Hamilton | G5 | Airway Pressure Release Ventilation with Tube Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Hamilton | G5 | Spontaneous with Tube Resistance Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
Hamilton | G5 | Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | G5 | Noninvasive Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | Galileo | Synchronized Controlled Mandatory Ventilation A/C | volume | CMV | set-point | N/A | VC-CMVs
Hamilton | Galileo | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Hamilton | Galileo | SIMV with Tube Resistance Compensation | volume | IMV | set-point | set-point/servo | VC-IMVs,sr
Hamilton | Galileo | Pressure Controlled Continuous Mandatory Ventilation A/C | pressure | CMV | set-point | N/A | PC-CMVs
Hamilton | Galileo | Adaptive Pressure Ventilation Continuous Mandatory Ventilation | pressure | CMV | adaptive | N/A | PC-CMVa
Hamilton | Galileo | Adaptive Pressure Ventilation Continuous Mandatory Ventilation with Tube Resistance Compensation | pressure | CMV | adaptive/servo | N/A | PC-CMVar
Hamilton | Galileo | Pressure Controlled Continuous Mandatory Ventilation A/C with Tube Resistance Compensation | pressure | CMV | set-point/servo | N/A | PC-CMVsr
Hamilton | Galileo | Pressure SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | Galileo | Duo Positive Airway Pressure | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | Galileo | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | Galileo | Adaptive Pressure Ventilation SIMV | pressure | IMV | adaptive | set-point | PC-IMVa,s
Hamilton | Galileo | Adaptive Pressure Ventilation SIMV with Tube Resistance Compensation | pressure | IMV | adaptive/servo | set-point/servo | PC-IMVar,sr
Hamilton | Galileo | Adaptive Support Ventilation | pressure | IMV | optimal/intelligent | optimal/intelligent | PC-IMVoi,oi
Hamilton | Galileo | Adaptive Support Ventilation with Tube Resistance Compensation | pressure | IMV | optimal/intelligent/servo | optimal/intelligent/servo | PC-IMVoir,oir
Hamilton | Galileo | Pressure SIMV with Tube Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Hamilton | Galileo | Duo Positive Airway Pressure with Tube Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Hamilton | Galileo | Airway Pressure Release Ventilation with Tube Resistance Compensation | pressure | IMV | set-point/servo | set-point/servo | PC-IMVsr,sr
Hamilton | Galileo | Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | Galileo | Noninvasive Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | Galileo | Spontaneous with Tube Resistance Compensation | pressure | CSV | set-point/servo | N/A | PC-CSVsr
Hamilton | T1 | Pressure Controlled Ventilation Plus | pressure | CMV | set-point | N/A | PC-CMVs
Hamilton | T1 | Synchronized Controlled Mandatory Ventilation Plus | pressure | CMV | adaptive | N/A | PC-CMVa
Hamilton | T1 | Pressure Controlled SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | T1 | Noninvasive Ventilation - Spontaneous Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | T1 | Duo Positive Airway Pressure | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | T1 | Airway Pressure Release Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Hamilton | T1 | SIMV Plus | pressure | IMV | adaptive | set-point | PC-IMVa,s
Hamilton | T1 | Adaptive Support Ventilation | pressure | IMV | optimal/intelligent | optimal/intelligent | PC-IMVoi,oi
Hamilton | T1 | Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Hamilton | T1 | Noninvasive Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Impact Instrumentation | Uni-Vent Eagle II | A/C Volume Target | volume | CMV | set-point | N/A | VC-CMVs
Impact Instrumentation | Uni-Vent Eagle II | SIMV Volume Target | volume | IMV | set-point | set-point | VC-IMVs,s
Impact Instrumentation | Uni-Vent Eagle II | A/C Pressure Target | pressure | CMV | set-point | N/A | PC-CMVs
Impact Instrumentation | Uni-Vent Eagle II | SIMV Pressure Target | pressure | IMV | set-point | set-point | PC-IMVs,s
Impact Instrumentation | Uni-Vent Eagle II | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | Servo 300 | Volume Control | volume | CMV | set-point | N/A | VC-CMVs
Maquet | Servo 300 | SIMV (Volume Control) | volume | IMV | set-point | set-point | VC-IMVs,s
Maquet | Servo 300 | Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Maquet | Servo 300 | Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
Maquet | Servo 300 | SIMV (Pressure Control) | pressure | IMV | set-point | set-point | PC-IMVs,s
Maquet | Servo 300 | Pressure Support/CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | Servo 300 | Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Maquet | SERVO-I | Volume Control (with square or ramp flow) | volume | CMV | set-point | N/A | VC-CMVs
Maquet | SERVO-I | Volume Control with Flow Adaptation | volume | IMV | dual | dual | VC-IMVd,d
Maquet | SERVO-I | SIMV (Volume Control with square or ramp flow) | volume | IMV | set-point | set-point | VC-IMVs,s
Maquet | SERVO-I | SIMV (Volume Control with Flow Adaptation) | volume | IMV | dual | set-point | VC-IMVd,ds
Maquet | SERVO-I | Automode (Volume Control with square or ramp flow to Volume Support) | volume | IMV | set-point/adaptive | adaptive | VC-IMVsa,a
Maquet | SERVO-I | Automode (Volume Control with Flow Adaptation to Volume Support) | volume | IMV | dual/adaptive | dual//adaptive | VC-IMVda,da
Maquet | SERVO-I | Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Maquet | SERVO-I | Noninvasive Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Maquet | SERVO-I | Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
Maquet | SERVO-I | SIMV (Pressure Control) | pressure | IMV | set-point | set-point | PC-IMVs,s
Maquet | SERVO-I | Bi-Vent | pressure | IMV | set-point | set-point | PC-IMVs,s
Maquet | SERVO-I | Automode (Pressure Control to Pressure Support) | pressure | IMV | set-point/adaptive | set-point | PC-IMVsa,s
Maquet | SERVO-I | SIMV Pressure Regulated Volume Control | pressure | IMV | adaptive | set-point | PC-IMVa,s
Maquet | SERVO-I | Automode (Pressure Regulated Volume Control to Volume Support) | pressure | IMV | adaptive | adaptive | PC-IMVa,a
Maquet | SERVO-I | Spontaneous/CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-I | Nasal CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-I | Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-I | Noninvasive Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-I | Neurally Adjusted Ventilatory Assist | pressure | CSV | servo | N/A | PC-CSVr
Maquet | SERVO-I | Noninvasive Neurally Adjusted Ventilatory Assist | pressure | CSV | servo | N/A | PC-CSVr
Maquet | SERVO-I | Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Maquet | SERVO-U | Volume Control (with square or ramp flow) | volume | CMV | set-point | N/A | VC-CMVs
Maquet | SERVO-U | Volume Control with Flow Adaptation | volume | IMV | dual | dual | VC-IMVd,d
Maquet | SERVO-U | SIMV (Volume Control with ramp or square flow) | volume | IMV | set-point | set-point | VC-IMVs,s
Maquet | SERVO-U | SIMV (Volume Control with Flow Adaptation) | volume | IMV | dual | set-point | VC-IMVd,ds
Maquet | SERVO-U | Automode (Volume Control with square or ramp flow to Volume Support) | volume | IMV | set-point/adaptive | adaptive | VC-IMVsa,a
Maquet | SERVO-U | Automode (Volume Control with Flow Adaptation to Volume Support) | volume | IMV | dual/adaptive | dual//adaptive | VC-IMVda,da
Maquet | SERVO-U | Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Maquet | SERVO-U | Pressure Regulated Volume Control | pressure | CMV | adaptive | N/A | PC-CMVa
Maquet | SERVO-U | Noninvasive Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Maquet | SERVO-U | SIMV (Pressure Control) | pressure | IMV | set-point | set-point | PC-IMVs,s
Maquet | SERVO-U | Bi-Vent | pressure | IMV | set-point | set-point | PC-IMVs,s
Maquet | SERVO-U | Automode (Pressure Control to Pressure Support) | pressure | IMV | set-point/adaptive | set-point | PC-IMVsa,s
Maquet | SERVO-U | SIMV Pressure Regulated Volume Control | pressure | IMV | adaptive | set-point | PC-IMVa,s
Maquet | SERVO-U | Automode (Pressure Regulated Volume Control to Volume Support) | pressure | IMV | adaptive | adaptive | PC-IMVa,a
Maquet | SERVO-U | Spontaneous/CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-U | Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-U | Noninvasive Pressure Support | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-U | Noninvasive CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Maquet | SERVO-U | Neurally Adjusted Ventilatory Assist | pressure | CSV | servo | N/A | PC-CSVr
Maquet | SERVO-U | Noninvasive Neurally Adjusted Ventilatory Assist | pressure | CSV | servo | N/A | PC-CSVr
Maquet | SERVO-U | Volume Support | pressure | CSV | adaptive | N/A | PC-CSVa
Oricare | V8800 | A/C | volume | CMV | set-point | N/A | VC-CMVs
Oricare | V8800 | SIMV (Volume Control Ventilation) + Pressure Support Ventilation | volume | IMV | set-point | set-point | VC-IMVs,s
Oricare | V8800 | Pressure Controlled A/C | pressure | CMV | set-point | N/A | PC-CMVs
Oricare | V8800 | Noninvasive Ventilation Pressure Control Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Oricare | V8800 | Pressure Regulation Volume Controlled | pressure | CMV | adaptive | N/A | PC-CMVa
Oricare | V8800 | SIMV (Pressure Control Ventilation) + Pressure Support Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Oricare | V8800 | BIVENT/Airway Pressure Release Ventilatino + Pressure Support Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Oricare | V8800 | SIMV (Pressure Regulated Volume Control) + Pressure Support Ventilation | pressure | IMV | adaptive | set-point | PC-IMVa,s
Oricare | V8800 | Spontaneous Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Oricare | V8800 | Noninvasive Ventilation/CPAP + Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
Percussionaire | IVP | Intrapulmonary Percussive Ventilation | time | IMV | set-point | set-point | TC-IMVs,s
Percussionaire | VDR4 | Volumetric Diffusive Respiration | time | IMV | set-point | set-point | TC-IMVs,s
Philips Respironics | Trilogy 202 | Pressure Control | pressure | CMV | set-point | N/A | PC-CMVs
Philips Respironics | Trilogy 202 | Pressure Control with Average Volume Assured Pressure Support | pressure | CMV | adaptive | N/A | PC-CMVa
Philips Respironics | Trilogy 202 | SIMV | volume | IMV | set-point | set-point | VC-IMVs,s
Philips Respironics | Trilogy 202 | Pressure Control SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | Trilogy 202 | Spontaneous/Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | Trilogy 202 | Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | Trilogy 202 | Timed with Average Volume Assured Pressure Support | pressure | IMV | adaptive | set-point | PC-IMVa,s
Philips Respironics | Trilogy 202 | Spontaneous/Timed with Average Volume Assured Pressure Support | pressure | IMV | adaptive | set-point | PC-IMVa,s
Philips Respironics | Trilogy 202 | Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | Trilogy 202 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | Trilogy 202 | Spontaneous with Average Volume Assured Pressure Support | pressure | CSV | adaptive | N/A | PC-CSVa
Philips Respironics | Esprit | Volume Control Ventilation A/C | volume | CMV | set-point | N/A | VC-CMVs
Philips Respironics | Esprit | Volume Control Ventilation A/C with Flow-Trak | volume | IMV | dual | dual | VC-IMVd,d
Philips Respironics | Esprit | Volume Control IMV | volume | IMV | set-point | set-point | VC-IMVs,s
Philips Respironics | Esprit | Volume Control IMV with Flow-Trak | volume | IMV | dual | set-point | VC-IMVd,ds
Philips Respironics | Esprit | Pressure Control A/C | pressure | CMV | set-point | set-point | PC-CMVs
Philips Respironics | Esprit | Pressure Control IMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | Esprit | Non-invasive Positive Pressure Ventilation Spontaneous/Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | Esprit | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | Esprit | Non-invasive Positive Pressure Ventilation Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | V200 | Volume Control Ventilation A/C | volume | CMV | set-point | N/A | VC-CMVs
Philips Respironics | V200 | Volume Control Ventilation A/C with Flow-Trak | volume | IMV | dual | dual | VC-IMVd,d
Philips Respironics | V200 | Volume Control IMV | volume | IMV | set-point | set-point | VC-IMVs,s
Philips Respironics | V200 | Volume Control IMV with Flow-Trak | volume | IMV | dual | set-point | VC-IMVd,ds
Philips Respironics | V200 | Pressure Control A/C | pressure | CMV | set-point | N/A | PC-CMVs
Philips Respironics | V200 | Pressure Control IMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | V200 | Non-invasive Positive Pressure Ventilation Spontaneous/Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | V200 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | V200 | Non-invasive Positive Pressure Ventilation Spontaneous | pressure | CSV | set-point | N/A | PC-CSVs
Philips Respironics | V60 | Pressure Controlled Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Philips Respironics | V60 | Spontaneous/Timed | pressure | IMV | set-point | set-point | PC-IMVs,s
Philips Respironics | V60 | Proportional Pressure Support | pressure | IMV | set-point | servo | PC-IMVs,r
Philips Respironics | V60 | Average Volume-Assured Pressure Support | pressure | IMV | adaptive | adaptive | PC-IMVa,a
Philips Respironics | V60 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
SLE Limited | SLE5000 | Continuous Mandatory Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
SLE Limited | SLE5000 | Continuous Mandatory Ventilation with Targeted Tidal Volume | pressure | CMV | adaptive | N/A | PC-CMVa
SLE Limited | SLE5000 | Patient Triggered Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
SLE Limited | SLE5000 | Patient Triggered Ventilation with Targeted Tidal Volume | pressure | CMV | adaptive | N/A | PC-CMVa
SLE Limited | SLE5000 | SIMV with Pressure Support Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
SLE Limited | SLE5000 | SIMV Ventilation with Targeted Tidal Volume | pressure | IMV | adaptive | set-point | PC-IMVa,s
SLE Limited | SLE5000 | High Frequency Oscillation | time | IMV | set-point | set-point | TC-IMVs,s
SLE Limited | SLE5000 | High Frequency Oscillation with Continuous Mandatory Ventilation | time | IMV | set-point | set-point | TC-IMVs,s
SLE Limited | SLE5000 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
SLE Limited | SLE5000 | Pressure Support Ventilation | pressure | CSV | set-point | N/A | PC-CSVs
SLE Limited | SLE5000 | Pressure Support Ventilation with Targeted Tidal Volume | pressure | CSV | adaptive | N/A | PC-CSVa
Smiths | babyPAC 100 | Controlled Mandatory Ventilation + PEEP | pressure | CMV | set-point | N/A | PC-CMVs
Smiths | babyPAC 100 | Controlled Mandatory Ventilation + Active PEEP | pressure | IMV | set-point | set-point | PC-IMVs,s
Smiths | babyPAC 100 | IMV + CPAP | pressure | IMV | set-point | set-point | PC-IMVs,s
Smiths | babyPAC 100 | CPAP | pressure | CSV | set-point | N/A | PC-CSVs
Heinen + Löwenstein | leoniplus | Intermittent Positive Pressure Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Heinen + Löwenstein | leoniplus | Synchronized Intermittent Positive Pressure Ventilation | pressure | CMV | set-point | N/A | PC-CMVs
Heinen + Löwenstein | leoniplus | Synchronized Intermittent Positive Pressure Ventilation with Volume Guarantee | pressure | CMV | adaptive | N/A | PC-CMVa
Heinen + Löwenstein | leoniplus | Nasal Intermittent Positive Pressure Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Heinen + Löwenstein | leoniplus | IMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Heinen + Löwenstein | leoniplus | SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Heinen + Löwenstein | leoniplus | Pressure Support Ventilation - Synchronized Intermittent Positive Pressure Ventilation | pressure | IMV | set-point | set-point | PC-IMVs,s
Heinen + Löwenstein | leoniplus | Pressure Support Ventilation - SIMV | pressure | IMV | set-point | set-point | PC-IMVs,s
Heinen + Löwenstein | leoniplus | High Frequency Oscillation | time | IMV | set-point | set-point | TC-IMVs,s
Heinen + Löwenstein | leoniplus | SIMV with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Heinen + Löwenstein | leoniplus | Pressure Support - Synchronized Intermittent Positive Pressure Ventilation with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Heinen + Löwenstein | leoniplus | Pressure Support - SIMV with Volume Guarantee | pressure | IMV | adaptive | set-point | PC-IMVa,s
Heinen + Löwenstein | leoniplus | High Frequency Oscillation | pressure | IMV | adaptive | set-point | PC-IMVa,s
Heinen + Löwenstein | leoniplus | CPAP | pressure | CSV | N/A | set-point | PC-CSVs
Heinen + Löwenstein | leoniplus | Nasal CPAP | pressure | CSV | N/A | set-point | PC-CSVs
Carl Reiner | TwinStream | Bi Level Ventilation | time | IMV | set-point | set-point | PC-IMVs,s
Carl Reiner | TwinStream | Pulsatile Bi Level Ventilation | time | IMV | set-point | set-point | TC-IMVs,s
Carl Reiner | TwinStream | Pulsatile CPAP | time | IMV | set-point | set-point | PC-IMVs,s
Carl Reiner | TwinStream | CPAP | pressure | CSV | N/A | set-point | PC-CSVs
Stephan | Stephanie | A/C Flow Controlled | volume | CMV | set-point | N/A | VC-CMVs
Stephan | Stephanie | A/C Flow Controlled with Pressure Limit | volume | CMV | dual | N/A | VC-CMVd
Stephan | Stephanie | Controlled Mandatory Ventilation Flow Controlled | volume | IMV | set-point | set-point | VC-IMVs,s
Stephan | Stephanie | Controlled Mandatory Ventilation Flow Controlled with Pressure Limit | volume | IMV | dual | set-point | VC-IMVd,s
Stephan | Stephanie | SIMV Flow Controlled | volume | IMV | set-point | set-point | VC-IMVs,s
Stephan | Stephanie | SIMV Flow Controlled with Pressure Limit | volume | IMV | dual | set-point | VC-IMVd,s
Stephan | Stephanie | A/C Pressure Controlled | pressure | CMV | set-point | N/A | PC-CMVs
Stephan | Stephanie | A/C Pressure Controlled with Volume Limitation | pressure | CMV | adaptive | N/A | PC-CMVa
Stephan | Stephanie | SIMV Pressure Cotnrolled | pressure | IMV | set-point | set-point | PC-IMVs,s
Stephan | Stephanie | SIMV Pressure Cotnrolled with Volume Limitation | pressure | IMV | adaptive | set-point | PC-IMVa,s
Stephan | Stephanie | High Frequency Oscillation | time | IMV | set-point | set-point | TC-IMCs,s
Stephan | Stephanie | Pressure Support Ventilation | pressure | CSV | N/A | set-point | PC-CSVs
Stephan | Stephanie | Proportional Assist Ventilation | pressure | CSV | N/A | servo | PC-CSVr
Stephan | Stephanie | CPAP | pressure | CSV | N/A | set-point | PC-CSVs
Stephan | Sophie | A/C Pressure Controlled | pressure | CMV | set-point | N/A | PC-CMVs
Stephan | Sophie | A/C Pressure Controlled with Volume Limitation | pressure | CMV | adaptive | N/A | PC-CMVa
Stephan | Sophie | A/C Pressure Controlled with Inspiratory Time Termination | pressure | IMV | set-point | set-point | PC-IMVs,s
Stephan | Sophie | SIMV Pressure Cotnrolled | pressure | IMV | set-point | set-point | PC-IMVs,s
Stephan | Sophie | SIMV Pressure Cotnrolled with Volume Limitation | pressure | IMV | adaptive | set-point | PC-IMVa,s
Stephan | Sophie | High Frequency Oscillation | time | IMV | set-point | set-point | TC-IMCs,s
Stephan | Sophie | Pressure Support Ventilation | pressure | CSV | N/A | set-point | PC-CSVs
Stephan | Sophie | CPAP | pressure | CSV | N/A | set-point | PC-CSVs
CONTROL VARIABLES
VC = volume control
PC = pressure control
TC = time control
BREATH SEQUENCES
CMV = continuous mandatory ventilation
IMV = intermittent mandatory ventilation
CSV = continuous spontaneous ventilation
TARGETING SCHEMES
s = set-point
d = dual
b = bio-variable
r = servo
a = adaptive
o = optimal
i = intelligent
Menu

## Maxims
Ten Fundamental Maxims for Ventilator Technology
Reference
Chatburn RL, Khatib ME, Mireles-Cabodevila E. A Taxonomy for Mechanical Ventilation: 10 Fundamental Maxims. Respir Care 2014;59(11):1747–1763.
https://app.box.com/s/sbnajwyhdo53yf6wqcv50orxypumb16j
1. A breath is one cycle of positive flow (inspiration) and negative flow (expiration) defined in terms of the flow-time curve.
Inspiratory time is defined as the period from the start of positive flow to the start of negative flow. Expiratory time is defined as the period from the start of expiratory flow to the start of inspiratory flow. The flow-time curve is the basis for many variables related to ventilator settings.
2. A breath is assisted if the ventilator does work on the patient.
An assisted breath is one for which the ventilator does some portion of the work of breathing. For constant flow inflation, work is defined as inspiratory pressure multiplied by tidal volume. Therefore, an assisted breath is identified as a breath for which airway pressure (displayed on the ventilator) rises above baseline during inspiration. An unassisted breath is one for which the ventilator simply provides the inspiratory flow demanded by the patient and pressure stays constant throughout the breath.
3. A ventilator assists breathing using either pressure control or volume control based on the equation of motion for the respiratory system.
Providing assistance means doing work on the patient, which is accomplished by controlling either pressure or volume. A simple mathematical model describing this fact is known as the equation of motion for the passive respiratory system: Pressure = (Elastance × Volume) + (Resistance × Flow)
Volume control means that both volume and flow are preset prior to inspiration. In other words, the right hand side of the equation of motion remains constant while pressure changes with changes in elastance and resistance.
Pressure control means that inspiratory pressure is preset as either a constant value or it is proportional to the patient’s inspiratory effort. In other words, the left hand side of the equation of motion remains constant while volume and flow change with changes in elastance and resistance.
Time control means that neither pressure nor volume/flow are preset; the only preset targets are frequency and inspiratory time
4. Breaths are classified by the criteria that trigger (start) and cycle (stop) inspiration.
The start of inspiration is called the trigger event. The end of inspiration is called the cycle event.
5. Trigger and cycle events can be initiated by the patient or the machine.
A spontaneous breath is a breath for which the patient both triggers and cycles the breath. A spontaneous breath may occur during a mandatory breath (eg Airway Pressure Release Ventilation). A spontaneous breath may be assisted or unassisted.
A mandatory breath is a breath for which the machine triggers and/or cycles the breath. A mandatory breath can occur during a spontaneous breath (eg, High Frequency Jet Ventilation). A mandatory breath is, by definition, assisted.
7. There are 3 breath sequences: Continuous mandatory ventilation , Intermittent Mandatory Ventilation, and Continuous Spontaneous Ventilation.
A breath sequence is a particular pattern of spontaneous and/or mandatory breaths. The 3 possible breath sequences are: continuous mandatory ventilation, (CMV, spontaneous breaths are not allowed between mandatory breaths), intermittent mandatory ventilation (IMV, spontaneous breaths may occur between mandatory breaths), and continuous spontaneous ventilation (CSV, all breaths are spontaneous).
8. There are 5 basic ventilatory patterns; VC-CMV, VC-IMV, PC-CMV, PC-IMV, PC-CSV
The combination VC-CSV is not possible because volume control implies machine cycling and machine cycling makes every breath mandatory, not spontaneous. TC-IMV is rare but possible with modes like High Frequency Oscillatory Ventilation and Intrapulmonary Percussive Ventilation
9. Within each ventilatory pattern there are several variations that can be distinguished by their targeting scheme(s).
A targeting scheme is a description of how the ventilator achieves preset targets. A target is a predetermined goal of ventilator output. Examples of within-breath targets include inspiratory flow or pressure and rise time, tidal volume and constant of proportionality between inspiratory pressure and patient effort. Examples of between-breath targets and targeting schemes include average tidal volume, percent minute ventilation and combined PCO2, volume, and frequency values describing a “zone of comfort”. The targeting scheme (or combination of targeting schemes) is what distinguishes one ventilatory pattern from another. There are 7 basic targeting schemes that comprise the wide variety seen in different modes of ventilation:
Set-point (s): A targeting scheme for which the operator sets all the parameters of the pressure waveform (pressure control modes) or volume and flow waveforms (volume control modes).
Dual (d): A targeting scheme that allows the ventilator to switch between volume control and pressure control during a single inspiration.
Bio-variable (b): A targeting scheme that allows the ventilator to automatically set the inspiratory pressure or tidal volume randomly to mimic the variability observed during normal breathing.
Servo (r): A targeting scheme for which inspiratory pressure is proportional to inspiratory effort.
Adaptive (a): A targeting scheme that allows the ventilator to automatically set one target (eg, pressure within a breath) to achieve another target (eg, average tidal volume over several breaths).
Optimal (o): A targeting scheme that automatically adjusts the targets of the ventilatory pattern to either minimize or maximize some overall performance characteristic (eg, minimize the work rate done by the ventilatory pattern).
Intelligent (i): A targeting scheme that uses artificial intelligence programs such as fuzzy logic, rule based expert systems, and artificial neural networks.
10. A mode of ventilation is classified according to its control variable, breath sequence, and targeting scheme(s
The preceding 9 maxims create a theoretical foundation for a taxonomy of mechanical ventilation. The taxonomy is based on these theoretical constructs and has 4 hierarchical levels:
Control Variable (Pressure or Volume, for the primary breath)
Breath Sequence (CMV, IMV, or CSV)
Primary Breath Targeting Scheme (for CMV or CSV)
Secondary Breath Targeting Scheme (for IMV)
The “primary breath” is either the only breath there is (mandatory for CMV and spontaneous for CSV) or it is the mandatory breath in IMV. The targeting schemes can be represented by single, lower case letters: set-point = s, dual = d, servo = r, bio-variable = b, adaptive = a, optimal = o, intelligent = i.
Return

## Q1
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 1: The breath
1. Define the following terms (see Glossary from Main Menu):
Breath
Inspiratory Time
Inspiratory Hold
Inspiratory Hold Time
Expiratory Time
Total Cycle Time (Ventilatory Period)
I:E Ratio
Duty Cycle
Tidal Volume
Minute Ventilation
Mean Airway Pressure
Answers to the following questions can be determined by experimentation with the simulator
2. If minute ventilation remains constant and frequency increases, what happens to tidal volume?
Tidal volume decreases
3. If frequency decreases, what happens to total cycle time?
Total cycle time increases
4. If inspiratory time remains constant and frequency increases, what happens to expiratory time?
Expiratory time decreases
5. If the I:E is 1:3, what is the duty cycle?
Duty cycle = 1/4 = 25%
6. If I:E changes from 1:2 to 1:3, does I:E increase or decrease?
I:E decreases (1/3 < 1/2)
7. If the I:E increases and frequency is constant, what happens to the mean airway pressure?
Mean airway pressure increases
8. If inspiratory hold is activated, what happens to inspiratory time?
Inspiratory time increases
9. If I:E increases but frequency does not change, what happens to expiratory time?
Expiratory time decreases
10. IF I:E stays constant, what happens to mean airway pressure as frequency increases?
Mean airway pressure remains constant
Return

## Q2
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 2: The assisted breath
1. Define the following terms (see Glossary from Main Menu):
Assisted Breath
Loaded Breath
Work of Breathing
Transrespiratory Pressure
Transpulmonary Pressure
Loaded Breath
Answers to the following questions can be determined by experimentation with the simulator
2. If the peak inspiratory pressure is the same as the PEEP, what is the ventilator work?
This is the same as CPAP and the ventilator work is zero, the breath is unassisted
3. In volume control, what happens to ventilator work and total work when the patient makes an inspiratory effort (Pmus > 0)?
Ventilator work decreases as patient work increases because volume and flow remain constant; work shifts from ventilator to patient (work shifting index > 0%), but total work remains constant
4. In pressure control, what happens to ventilator work when the patient makes an inspiratory effort (Pmus > 0)?
Ventilator work decreases, patient work increases, but unlike volume control, the total work also increases
5. For a patient that is intubated on a conventional ventilator, is transrespiratory pressure positive or negative during inspiration?
Positive: For example Paw = 10, Pes = -5, PTP = Paw - Pes = (10 -(-5) = 10 + 5 = 15
6. For a patient in an Iron Lung, is transrespiratory pressure positive or negative during inspiration?
Positive
7. When tidal volume increases, what happens to driving pressure?
Driving pressure (aka, tidal pressure) increases
8. If ventilatory assistance increases, what happens to tidal volume?
Increased ventilatory assistance means increased ventialtor work delivered to patient, so tidal volume increases
9. If ventilatory assistance increases, what happens to peak expiratory flow?
If tidal volume increases, peak expiratory flow increases
10. If ventilatory assistance decreases, what happens to mean inspiratory pressure?
Mean inspiratory pressure decreases
Return

## Q3
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 3: Assistance by volume or pressure control
1. Define the following terms (see Glossary from Main Menu):
Equation of Motion
Volume Control
Pressure Control
Time Control
Answers to the following questions can be determined by experimentation with the simulator
2. When switching from square flow waveform to ramp flow waveform, with the same tidal volume and peak flow, what happens to inspiratory time?
Inspiratory time doubles
3. In volume control, what happens to tidal volume and flow when compliance or resistance change?
Tidal volume and flow are not affected by lung model changes
4. In volume control what happens to inspiratory pressure when compliance or resistance change?
Inspiratory pressure increases as compliance decreases or resistance increases
5. In volume control what happens to inspiratory volume and flow when inspiratory effort increases (Pmus > 0)?
Tidal volume and flow are not affected by effort model changes
6. In volume control what happens to inspiratory pressure when inspiratory effort increases (Pmus > 0)?
Inspiratory pressure decreases as Pmus increases
7. In pressure control, what happens to tidal volume and flow when compliance or resistance change?
Tidal volume and flow increase as compliance increases or resistance decreases
8. In pressure control what happens to inspiratory pressure when compliance or resistance change?
Inspiratory pressure increases is not affected by lung model changes
9. In pressure control what happens to inspiratory volume and flow when inspiratory effort increases (Pmus > 0)?
Tidal volume and flow increase as Pmus increases
10. In pressure control what happens to inspiratory pressure when inspiratory effort increases (Pmus > 0)?
Inspiratory pressure is not affected by effort model changes
11. True or False: In volume control with both square and ramp flow waveform, the initial pressure rise (above PEEP) is equal to resistance times flow
TRUE
12. True or False. In volume control with both square and ramp flow waveforms, the end inspiratory pressure (above PEEP) is equal to elastance times tidal volume (called tidal pressure or driving pressure)
FALSE
End inspiratory pressure (above PEEP) is equal to driving pressure (tidal pressure) only for ramp flow (assuming end inspiratory flow is zero)
13 True or False. In volume control with ramp flow waveforms, the end inspiratory pressure (above atmospheric pressure) is equal to plateau pressure
TRUE
Return

## Q4
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 4: Breaths classified by trigger and cycle events
1. Define the following terms (see Glossary from Main Menu):
Trigger
Trigger Variable
Cycle
Cycle Variable
2. True or False: The cycle variable determines the inspiratory time.
TRUE
3. True or False: Once the inspiratory time is set, the trigger variable, by controlling the trigger frequency, determines the expiratory time
TRUE
4. In PC-CSV (eg, Pressure Support), cycling occurs when inspiratory flow decreases to a preset percentage of the peak inspiratory flow
TRUE
5. In PC-CSV (eg, Pressure Support), what happens to the tidal volume if the cycle threshold is increased?
Tidal volume decreases because the average inspiratory time decreases
6. True or False: Triggering is NOT affected by the sensitivity setting
FALSE
7. What happens to trigger sensitivity if the trigger threshold is increased?
Sensitivity decreases as the trigger threshold increases (ie, more effort is required to start inspiration)
8. What is the main cycle variable in volume control?
Tidal volume
9. What is the main cycle variable in pressure control?
Inspiratory time
10. How do trigger variables differ between volume control and pressure control?
They are the same
Return

## Q5
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 5: Patient or machine initiation of trigger and cycle events
1. Define the following terms (see Glossary from Main Menu):
Patient Triggering
Patient Cycling
Machine Triggering
Machine Cycling
Trigger Window
Synchronization Window
Pressure Triggering
Pressure Cycling
Volume Triggering
Volume Cycling
Flow Triggering
Flow Cycling
2. What variables might the ventilator monitor to allow patient triggering?
Pressure, volume, or flow change at the airway opening; chest impedance, electrical signal from the diaphragm
3. What variables might the ventilator monitor to allow machine triggering?
Time, patient trigger rate, minute ventilation
4. What variables might the ventilator monitor to allow patient cycling
Inspiratory pressure or flow, electrical signal from the diaphragm
5. What variables might the ventilator monitor to allow machine cycling
Inspiratory time or tidal volume
6. True of False: Flow cycling during Pressure Support is an example of machine cyclling
FALSE because the time for inspiratory flow to decay from the peak value to the cycle threshold is dependent on respiratory system mechanics (ie, inspiratory time is random)
7. Pressure cycling during volume control is an exmample of patient cycling
TRUE because inspiraoty pressure during volume control depends on respiratory system mechanics (however this is an alarm condition)
8. True or False: The trigger window is usually longer than the synchronization window
TRUE because the trigger window is nearly the whole expiratory time, but the synchronization window is only a small portion of the expiratory time just before the scheduled machine trigger time
9. True or False: The trigger window makes Assist/Control possible whereas the synchronization window makes SIMV possible
TRUE: if trigger and cycle variables are pressure or flow, then triggering and cycling can be achieved by the passive respiratory system mechanics alone (eg, Vortran ventilator)
10. True or False: Patient triggering and cycling can occur even if the patient is paralyzed
TRUE: if trigger and cycle variables are pressure or flow, then triggering and cycling can be achieved by the passive respiratory system mechanics alone (eg, Vortran ventilator)
Return

## Q6
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 6: Spontaneous and mandatory breaths
1. Define the following terms (see Glossary from Main Menu):
Spontaneous breath
Mandatory breath
2. True or False: If inspiration is patient triggered it is considered to be spontaneous
FALSE, the definition of a spontaneous breath is one for which inspiration is both triggered and cycleD by the patient
3. A breath for which inspiration is patient triggered and patient cycled is classified as spontaneous
TRUE
4. A brath for which inspiration is patient triggered and machine cycled is classified as mandatory
TRUE
5. A breath for which inspiration is machine triggered and patient cycled is classified as mandatory
TRUE
6. A breath for which inspiration is machine triggered and machine cycled is classified as mandatory
TRUE
7. True or False: A spontaneous breath may occur durring a mandatory breath
TRUE, a pressure control breath with a long inspiratory time may allow one or more spontaneous breaths during the preset inspiratory time (eg, APRV), if there is an active exhalation valve
8. True or False: A mandatory breath may occur during a spontaneous breath
TRUE, this is what happens during various forms of high frequency ventilation
9. True or False: Spontaneous breaths generally promote patient-ventilator synchrony better than mandatory breaths
TRUE, the purpose of spontaneous breaths is to allow the patient to retain substantial control of inspiratory time and ventilatory frequency
10. True or False: One advantage of mandatory breaths is that they may sustain oxygenation better than spontaneous breaths
TRUE, for patients with severe ARDS, mandatory breaths can maintain a longer inspiratory time and hence a higher mean inspiratory pressure and mean airway pressure
11. For mandatory PC breaths, what happens to tidal volume if the pressure rise time is increased?
Tidal volume decreases because inspiratory time is fixed and mean inspiratory pressure decreases; tidal volume is proportional to mean inspiratory pressure and inspiratory time
12. For spontaneous PC breaths, what happens to tidal volume if the pressure rise time is increased?
Tidal volume increases because inspiratory time increases much more than mean inspiratory pressure decreases; tidal volume is proportional to mean inspiratory pressure and inspiratory time
Return

## Q7
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 7: Breath sequences
1. Define the following terms (see Glossary from Main Menu):
Continuous mandatory ventilation (CMV)
Intermittent mandatory ventilation (IMV)
Continuous spontaneous ventilation (CSV)
2. True or False: During CMV, the set frequency is the lowest frequency of mandatory breaths that can occur
TRUE
3. True or False: During IMV, the set frequency is the lowest frequency of mandatory breaths that can occur
FALSE, it is the highest frequency of mandatory breaths that can occur
4. Describe the three types of IMV
Type 1: the set frequency of mandatory breaths will always occur
Type 2: mandatory breaths are suppressed if the spontaneous breath frequency is higher than the set frequency
Type 3: mandatory breaths are suppressed if the spontaneous breath frequency causes the total minute ventilation to exceed the set minute ventilation target
5. True or False: Some ventilators use an active exhalation valve to allow spontaneous breaths during a pressure control mandatory breath, so this is considered a form of IMV
FALSE, the definition of IMV is that spontaneous breaths may occur between mandatory breaths
6. True or False: Decreasing the set frequency during CMV will decrease ventialtory support
This is only true if the patient triggered breath frequency does not increase (eg, a paralyzed patient)
7. True or False: Decreasing the set frequency during IMV will decrease ventilatory support
This is only true if the spontaneous breaths are not assisted (eg, with Pressure Support) to a tidal volume close to mandatory breaths. Otherwise, the patient may simply substitute assisted spontaneous breaths for the missing mandatory breaths
8. True of False: CSV always results in better patient-ventilator synchrony than CMV or IMV
False, Type 2 and type 3 IMV allow mandatory breaths to "get out of the way" if the patient can sustain spontaneous triggering, so synchrony is potentially the same with the added benefit of safety in the case of apnea
9. What is a disadvantage of CMV in terms of autoPEEP?
As the frequency of patient triggered breaths increases, expiratory time decreases and the risk of autoPEEP increases
10. What is the primary disadvantage of CSV?
If the patient becomes apneic, there will be an alarm and ventilation reverts to a backup mode which may not have been set properly
Return

## Q8
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 8: Ventilatory patterns
1. Define the following terms (see Glossary from Main Menu):
VC-CMV
VC-IMV
PC-CMV
PC-IMV
PC-CSV
TC-IMV
2. When you see the term "assist/control" in the literature what does it mean?
In the adult literature it generally means VC-CMV, but in the pediatric literature it means PC-CMV
3. The mode called Mandatory Minute Volume on Dräger ventilators is what type of ventilatory pattern?
VC-IMV
4. Which mode would you expect to provide better patient synchrony, VC-CMV or PC-CMV?
While both modes may create trigger and cycle dyssynchrony, but PC-CMV is better because the patient can determine inspiratory flow and tidal volume
5. Which ventilatory pattern best serves the goal of safety?
VC-CMV; minute ventilation is essentially independent of lung mechanics
6. Which ventilatory pattern best serves the goal of comfort?
PC-CSV; the patient can adjust inspiratory flow, tidal volume, inspiratory time and ventilatory frequency
7. Which ventialtory pattern is often called "time-cycled, pressure-limited" in the pediatric literature?
PC-IMV
8. The mode called Pressure Support has which ventilatory pattern?
PC-CSV
9. Give examples of modes that are classified as TC-IMV?
High frequency oscillatory ventilation and Intrapulmonary Percussive Ventilation
10. Why is VC-CSV not a valid ventilatory pattern?
VC implies that all breaths are volume cycled, meaning that they are machine cycled
A breath that is machine cycled can never be a spontaneous breath (by definition)
Therefore, VC and CSV are muturally exclusive terms
Return

## Q9
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 9: Targeting schemes
1. Define the following terms (see Glossary from Main Menu):
Target
Targeting scheme
Set-point targeting
Dual targeting
Bio-variable targeting
Servo targeting
Adatpive targeting
Optimal targeting
Intelligent targeting
2. Give examples of within-breath targets
VC: tidal volume, inspiratory flow; PC: inspiratory pressure, inspiratory time
3. Give examples of between-breath targets
Average tidal volume (eg Pressure Regulated Volume Control), minute ventilation (eg, Mandatory Minute Volume Ventilation)
4. What is the difference between a targeting scheme and a feedback control scheme
They are the same thing in the context of mechanical ventilation
5. Give examples of mode names that use set-point targeting
Volume Assist/Control or SIMV, Pressure Assist/Control or SIMV, Pressure Support, CPAP
6. What is the purpose of dual-targeting in a volume control mode?
It swithces the breath to pressure control if it senses that the patient is not getting enough inspiratory flow
7. Give examples of names of modes that use servo targeting
Proportional Assist Ventilation (PAV), Neurally Adjusted Ventilatory Assist (NAVA), Automatic Tube Compensation
8. Give examples of names of modes that use adaptive targeting
Volume Control Plus, Volume Gurantee, Pressure Regulated Volume Support, Volume Support
9. Give an example of mode that uses optimal targeting
Adaptive Support Ventilation
10. Give an example of a mode that uses intelligent targeting
IntelliVent ASV, SmartCare PS
Return

## Q10
Self-Study Questions
To reveal answers, select columns C and E and right click and select Unhide
To conceal answers, select column D and then right click and select Hide
Maxim 10: Mode classification
1. Define the following terms (see Glossary from Main Menu):
Mode of ventilation
Taxonomy
Mode variety
2. Give the complete mode classification as an abbreviation (tag) for Volume Assist/Control, Pressure Regulated Volume Control, and Pressure Support
A/C = VC-CMVs, PRVC = PC-CMVa, PS = PC-CSVs
3. What is the difference between Volume Control on the Medtronic PB 840 ventilator and the Maquet Servo-i ventialtor?
PB 840 is VC-CMVs; Servo-i is VC-CMVd
4. Give examples of different varieties of PC-CSVr
Both Proportional Assist Ventilation (PAV) and Neurally Adjusted Ventilatory Assist (NAVA) are PC-CSVr but they have very different trigger and cycle variables and different targeting scheme software
5. How would you classify Tube Compensation?
It is another variety of PC-CSVr
5. What is the classification of the mode called Volume Assist?
PC-CSVa
6. On the Dräger V500 ventilator, how does activating AutoFlow in Volume A/C affect the mode classification?
It changes from VC-CMVs to PC-CMVa
7. How is AutoMode on the Maquet Servo-i classified?
There are 4 ways to configure AutoMode (note that adaptive targeting is used both to set the mandatory breath frequency and to set the inspiratory pressure, depending on the mode)
-Volume Control (square or ramp flow) to Volume Support: VC-IMVsa,a
-Volume Control (with Flow Adaptation) to Volume Support: VC-IMVda,da
-Pressure Regulated Volume Control to Volume Support: PC-IMVa,a
-Pressure Control to Pressure Support: PC-IMVsa,s
8. What is the primary advantage and primary disadvantage of PC-CSVa?
Primary advantage, with normal inspiratory effort it serves the goals of comfort (patient can modify flow and volume) and safety (tidal volume regulation to average set value)
Primary disadvantage: with high inspiratory effort tidal volume dosage can become excessive while ventilatory support decreases to very low levels because mode cannot distinguish between increasing compliance and increasing effort
9. What is the advantage of Mandatory Minute Volume and Adaptive Support Ventilation over conventional IMV?
With MMV, spontaneous breaths can suppress mandatory breaths and thus improve patient-ventilator synchrony
10. Why are there so many different modes of ventilation?
They all differ in the extent to which they serve the 3 goals of mechanical ventilation, Safety, Comfort, and Liberation
There is no compelling data to suggest that any mode is superior to any other mode for all patients, hence engineers create modes that at least theoretically better serve the 3 goals
11. During normal, spontaneous, unassisted breathing (ie, no ventilator) what mode of ventilation does the brain impose? Does any ventilator simulate this?
Neural stimulation of the inspiratory muscles is translated into increasing traspulmonary pressure, which generates inspiratory flow
Hence, the control variable is pressure
Inspiration is, by definition, both patient triggered and patient cycled
Hence the ventilatory pattern is PC-CSV
According to John West, MD (Respiratory Physiology the Essentials, 8th edition), the medullary respiratory center of the brain is responsible for the basic rhythm of ventilation, even when all known afferent stimuli have been abolished
Hence, inspiration would seem to be both time triggered and time cycle, although unlike a mechanical system the timing varies somewhat randomly around mean values
Given the above, we would have to conclude that the brain uses a bio-variable targeting scheme
The only mechanical ventilator mode similar to this is Dräger's Variable Pressure Support
Return